## 1999-11-17
RECALL_CLASSIFICATION_DATE | PRODUCT_TYPE | CLASSIFICATION | RECALL_NUMBER | PRODUCT | RECALLING_FIRM | MANUFACTURER | RECALL_INITIALIZATION_DATE | REASON | VOLUME | DISTRIBUTION
1999-11-17 | Devices | II | z-130-0 | Spline Ha Coated Cylinder Dental Implants Endosseous Implants Part 1852 | Manufacturer | Sulzer Calcitek Inc | May 3 1999 | 13 Mm And 15 Mm Implants Were Interchanged Each Size In Trays Intended Labeled For The Other | 1 010 Implants Were Distributed | Nationwide And Internationl
1999-11-17 | Devices | II | z-135/139-0 | A Percutaneous Transluminal Angioplasty Pta Catheter Catalog Numbers Pta Os 253 303 And 403 B Z Med Dilation Catheter Catalog Pdz S C Z Med Ii Dilatation Catheter Catalog Numbers Pds 630 643 And 655 D Pta Os Balloon Catheter Catalog Numbers Vsb 030 And 040 E Ghost Ii Dilation Catheter Catalog Numbers Gii 109 113 121 129 130 137 138 141 | Manufacturer | Numed Inc | January 29 1999 | Non Sterility | 302 Catheters Were Distributed | Massachusetts United Kingdom South Africa
1999-11-17 | Devices | II | z-154/159-0 | Immulite Anti Thyroglobulin Antibody And Immulite Anti Thyroid Peroxidase Antibody For In Vitro Diagnostic Use A Immulite Anti Thyroglobulin Antibody 50 Tests B Immulite Anti Thyroglobulin Antibody 100 Tests C Immulite Anti Thyroglobulin Antibody 500 Tests D Immulite Anti Thyroid Peroxidase Antibody 50 Tests E Immulite Anti Thyroid Peroxidase Antibody 100 Tests F Immulite Anti Thyroid Peroxidase Antibody 500 Tests | Diagnostic Products Corporation | Euro Dpc Ltd | July 15 1999 | The Sample Diluent Is Non Homogeneous And May Produce Inaccurate Test Results | 1 269 Kits Were Distributed | Nationwide And United Kingdom
1999-11-17 | Devices | II | z-160/169-0 | Port A Cath Implantable Access Systems In Kits And Trays | Manufacturer | Sims Deltec Inc | October 18 1999 | Fracture Of Device Catheter Can Result In Leaks | 17 485 Implantable Access Systems Were Distributed | Nationwide And International
1999-11-17 | Devices | II | z-170/171-0 | Revision J Software In The Nellcor Puritan Bennett 740 Ventilator System And 760 Ventilator System | Mallinckrodt | Nellcor Puritan Bennett Ireland Ltd | September 13 1999 | Software Can Cause Under Or Over Delivery Of Oxygen | 223 Units | Nationwide And International
1999-11-17 | Devices | II | z-172-0 | Medi Vac Non Conductive Suction Tube With Maxi Grip Connectors A Sterile Non Conductive Tubing Used With A Vacuum Powered Suction Apparatus | Allegiance Healthcare Corporation | Allegiance Healthcare Corporation Mexicali Baja Cal Mexico | October 28 1999 | The Outer Package Seal May Not Hold Thus The Sterility Of The Product May Be Compromised | 8 340 Units Were Distributed | Nationwide And International
1999-11-17 | Devices | II | z-176/189-0 | Alm Prismalix Prx Ceiling Mounted Surgical Lighting System Used To Provide Visible Illumination Of Surgical Patients And Other Patient Examination Procedures A Prx 6401 Series B Prx 8401 Series C Prx 6601 Series D Prx 8601 Series E Prx 4401 Series F Prx 6441 Series G Prx 8441 Series H Prx 4441 Series I Prx 6001 Series J Prx 8001 K Prx 4001 Series L Prx 4401 Series M Prx 4441 Series N Prx 6601 | Alm Surgical Equipment | Alm Sa Parc De Limere Ardon Orleans Cedex | September 17 1999 | The Main Arm Of The Device May Fall From The Shaft Because Of Missing Shaft Retaining Clips | 491 Devices Were Distributed | Nationwide Guam Canada
1999-11-17 | Devices | II | z-190/198-0 | Sri Brand Of Surgical Procedure Packs An Assembly Of Finished Medical Devices And Other Items Used During The Preparation And Or Performance Of Surgical Medical Procedures A Pack No | Manufacturer | Sterile Recoveries Inc | September 15 1999 | The Devices Were Labeled And Distributed As Sterile Devices Without Assurance That They Were Sterile | 410 Units Were Distributed | California
1999-11-17 | Devices | II | z-199-0 | Tyshak Brand Peripheral Balloon Dilation Catheter Model 610532 Rated Burst Pressure 3 | Manufacturer | Numed Inc | August 26 1999 | The Firm Inadvertently Attached A Tyshak Ii Balloon To A Tyshak I Catheter | 7 Catheters Were Distributed | Pennsylvania
1999-11-17 | Devices | II | z-200-0 | Oxy Peep Oxygen Diluter Kit Model P 63000 Dilutes Oxygen Being Administered To A Patient | Manufacturer | Pegasus Research Corporation Santa Ana California | October 11 1999 | The Exhalation Valve In The Mask May Stick | 516 Units Were Distributed | Nationwide
1999-11-17 | Devices | II | z-201/204-0 | Blood Collection Safety Systems As Follows A Safe Point M D Multi Draw Needle Cover System 21 Gauge X 1 1 2 Needle Packaged 150 Or 200 Pieces Per Box Or Individually Packaged | North American Medical Products | North American Medical Products Inc | September 23 1999 | Lack Of Assurance Of Sterility | 28 900 Catheters Bulk Lots | Nationwide
1999-11-17 | Devices | II | z-205-0 | Proclear Compatibles Daily Wear Contact Lens | Manufacturer | Biocompatibles Eyecare Inc | November 3 1999 | Leaking Blister Seal Of Contact Lens Packaging Compromised Sterility | 119 028 Lenses Were Distributed In The United States | Nationwide And International