## 1999-8-25
RECALL_CLASSIFICATION_DATE | PRODUCT_TYPE | CLASSIFICATION | RECALL_NUMBER | PRODUCT | RECALLING_FIRM | MANUFACTURER | RECALL_INITIALIZATION_DATE | REASON | VOLUME | DISTRIBUTION
1999-8-25 | Drugs | II | d-310-9 | Naproxen Tablets Usp 500 Mg Rx In 500 Tablet Bottles | Manufacturer | Geneva Pharmaceuticals Inc | June 16 1999 | Metal Wire Particle Contamination | 3 440 Bottles Were Distributed | Nationwide
1999-8-25 | Drugs | II | d-311/403-9 | 1 Antiliriumæ Injection Physostigmine Salicylate | Manufacturer | Steris Laboratories Phoenix Arizona | February 19 1999 | Lack Of Sufficient Data From An Adequately Validated Laboratory Test Method To Assure Products Meet All Regulatory Requirments Including But Not Limited To Methods Validation And Potency | Number Of Vials Or Units Distributed 1 230 412 Vials 2 35 012 Vials 3 389 970 15 Ml Bottles And 34 337 30 Ml Bottles 4 159 381 Vials 5 66 529 Bottles 6 131 795 Vials 7 78 126 Vials 8 32 374 Vials 9 13 450 Vials 10 29 990 Vials 11 1 244 548 10 Ml Vials And 1 088 261 30 Ml Vials 12 64 210 1 Ml And 363 968 5 Ml Vials 13 1 118 536 Vials 14 97 821 Vials 15 31 960 Vials 16 386 874 Vials 17 10 970 Vials 18 713 274 Vials 19 70 527 Vials 20 569 200 Vials 21 973 200 1 Ml Vials And 102 529 10 Ml Vials 22 336 186 Vials 23 48 926 Vials 24 87 824 Vials 25 13 068 Vials 26 210 813 Vials 27 127 826 Vials 28 107 441 Vials 29 30 449 494 Vials 31 10 018 Vials 32 17 058 Vials 33 78 638 Vials 34 163 090 Ampuls 35 139 070 Ampuls 36 96 192 Vials 37 45 361 Vials 38 136 607 Vials 39 82 786 Vials 40 28 794 Vials 41 338 345 Vials 42 173 074 Vials 43 311 675 Vials 44 65 512 1 Ml And 4 901 10 Ml Vials 45 1 591 Vials 46 98 145 Vials 47 17 775 Vials 48 8 643 Vials 49 147 894 Vials 50 111 133 Vials 51 33 620 Vials 52 484 679 5 Ml And 202 836 10 Ml Vials 53 683 697 Vials 54 94 269 Vials 55 47 769 Vials 56 2 811 407 Bottles 57 5 730 538 Bottles 58 48 995 Vials 59 119 418 Vials 60 6 727 Vials 61 1 347 045 Ampuls 62 139 521 Vials 63 1 685 494 Vials 64 43 929 Vials 65 35 018 Bottles 66 28 874 Bottles 67 54 170 Vials 68 16 984 10 Ml And 83 115 30 Ml Vials 69 18 057 Bottles 70 67 234 5 Ml 30 933 10 Ml And 28 277 15 Ml Bottles 71 210 270 Vials 72 534 974 Vials 73 78 350 10 Ml And 161 085 30 Ml Vials 74 40 581 Vials 75 167 238 5 Ml And 68 219 10 Ml Bottles 76 141 019 Bottles 77 188 280 0 3 Ml And 15 057 5 Ml Vials 78 12 619 Vials 79 115 140 Vials 80 745 750 Vials 81 14 773 Vials 82 409 785 Vials 83 36 590 Vials 84 229 098 Vials 85 34 072 Vials 86 280 424 Bottles 87 338 524 Vials 88 892 567 1 Ml And 1 568 078 5 Ml Vials 89 1 065 376 Vials 90 142 586 Bottles 91 45 948 Vials 92 24 721 Vials 93 20 740 Vials | Nationwide And International
1999-8-25 | Drugs | III | d-309-9 | Allegra Dæ Extended Release Tablets Fexofenadine Hcl 60 Mg And Pseudoephedrine Hcl 120 Mg In Physician 6 Ís Samples 30 And 100 Tablet Bottles Blister Strips Rx For The Relief Of Symptoms Associated With Seasonal Allergic Rhinitis | Manufacturer | Hoechst Marion Roussel Inc | May 17 1999 | Dissolution Rate Failure | The Firm Has Distributed The Following Amounts Of Product 100 Tablet Blister Strip Boxes 5 326 Boxes Between 3 5 98 12 22 98 30 Tablet Bottles Physician Samples 62 514 Bottles Between 2 13 98 11 3 98 Physician Blister Samples 129 264 Boxes Containing 8 3 Day Starter Boxes Between 4 3 98 5 27 98 100 Tablet Bottles 90 676 Bottles Between 4 27 98 6 22 98 | Nationwide
1999-8-25 | Devices | II | z-1116/1119-9 | Angiocomm Workstation Previously Statview 1 | Manufacturer | Quinton Instrument Company Bothell Washington | June 28 1999 | Under Certain Specific Circumstances Sequences Of A Single Plane Study May Overlap Sequences Of Bi Plane Study When Viewing Both On The Angiocomm Single Sequence Review Station Or The Single Sequence Network Interface Station Nis | 11 Review Stations And 7 Sequence Nis Stations Were Distributed | Alabama New York Texas West Virginia Washington State
1999-8-25 | Devices | II | z-1146-9 | Light Handle Covers Single Use Cover For Or Light Handles To Maintain Sterility During Or Procedures Sold In Single And Multi Packs And As A Component In Surgical Trays | Deroyal Surgical | Atlanta Thermoplastics Tucker Georgia | June 25 1999 | Surgical Device Cover On Handle May Come Off And Compromise Sterility During Surgery | 631 000 Covers Were Distributed | Nationwide
1999-8-25 | Devices | II | z-1147/1153-9 | Hoyer Power Lifter Actuators A 12 Volt Dc Electromechanical Device Marketed As A Component Of Several Patient Lifting Devices Part Of A Retrofit Kit For Unpowered Patient Lifting Devices Or As A Stand Alone Component A Model No | Sunrise Medical Home Healthcare Group | Scanac of Sweden | May 14 1999 | A Short Circuit In The System Will Cause The Batteries To Discharge Resulting In The Generation Of Heat Which May Cause Fire Fumes Smoke Or Loss Of Power | 1 803 Units Representing The Actuators Were Shipped The Actuators Were Situated In Finished Products | Nationwide Puerto Rico Thailand
1999-8-25 | Devices | II | z-1209-9 | Bennett Contour 2 Stereotactic Ready Mammography Machine An X Ray Device | Manufacturer | Trex Medical Corporation Copaigue New York | August 29 1997 | The Subtray Mandrels Pins Used To Attach The Mammography Bucky Tray Or The Optional Cytoguide To The Vertical C Arm Of The Mammography Machine Can Overtime Become Loose Or Bend And Break | 154 Units Were Distributed | Nationwide And International
1999-8-25 | Devices | II | z-1210-9 | Apligraft Graftskin Indicated For Use With Standard Therapeutic Compression For The Treatment Of Non Infected Partial And Full Thickness Skin Ulcers Due To Venous Insufficiency Of Greater Than 1 Month Duration Which Has Not Adequately Responded To Conventional Therapy | Manufacturer | Organigenesis Inc | July 1 1999 | Burkholderia Cepacia Contamination | 16 Units Were Distributed For Commercial Use | Alabama Arizona Florida Massachusetts Minnesota Washington State Texas
1999-8-25 | Devices | II | z-1211-9 | Iti Extender System Octa Abutment Also Referred To As The Bone Graft Octa Abutment H 2 | The Straumann Company | Institut Sraumann Ag Waldenbrug Switzerland | July 7 1999 | Potential For Abutment To Fracture Due To Error On The Tolerance Value For The Depth Of The Hole On The Modified Engineering Drawing | 27 Units Were Distributed | Connecticut Iowa Florida Minnesota
1999-8-25 | Devices | II | z-1213/1214-9/ code lot numbers: a) m80233000 through m86559000, m88059000 & m88080000; b) m85013000 through m86544000, m88053000, m88054000, m88055000, m88058000, m88073000, m88074000, m88049000, & m88293000 | Shiley Adult Tracheostomy Tubes Intended For Use In Providing Tracheal Access For Airway Management A Catalog Nos | Manufacturer | Mallinckrodt Inc | July 6 1999 | The Lug Pin Which Hold The Soft Swivel Flange To The Cannula May Break Causing Displacement Of The Flange | 263 246 Tubes Were Distributed Firm Estimated That Little Product Remained On Market At Time Of Recall Initiation | Nationwide And International
1999-8-25 | Devices | II | z-1216-9 | Synchron Systems Digoxin Reagent Catalog No | Beckman Coulter | Beckman Coulter Carlsbad California | June 4 1999 | The Devices May Produce False Low Digoxin Results | 2 554 Units Of Lot M 810065 1917 Units Of Lot M 904060 And 678 Units Of M 901265 Were Distributed | Nationwide And International
1999-8-25 | Devices | II | z-1217-9 | Heart Valve Harvested From A Cadaveric Donor | Manufacturer | New York Organ Donor Network Nyodn New York New York | March 13 1998 | The Specimen Collected From Donor Was Found To Be Repeatedly Reactive For Hbsag | 1 Unit Was Distributed | Georgia
1999-8-25 | Devices | II | z-1218-9 | Sorin Blood Cardioplegia Console Model Bcc 02 Part R Dp 48 0001 102 Designed For Use During Cardiopulmonary Bypass Procedures To Control The Introduction Of Blood And Crystaloid Delivery In The Management Of Cardioplegia Delivery | Sorin Biomedical | Seamed Corporation Redmond Washington | March 15 1999 | The Device May Deliver The Blood Cardioplegia Solution At Uncontrolled High Speed Overspeed Pump Condition | 80 Units Were Distributed | Florida Georgia Illinois Massachusetts Maine Michigan Missouri Nebraska Pennsylvania Rhode Island Texas Washington State Canada Italy Japan
1999-8-25 | Devices | III | z-1182/1201-9 | Omega Plus Compression Hip Screw System With Varied Product Codes A Product Code 33630304 B Product Code 33630305 C Product Code 33630354 D Product Code 33630355 E Product Code 33630405 F Product Code 33630454 G Product Code 33630455 H Product Code 33631002 I Product Code 33631002 J Product Code 33631004 K Product Code 33631004 L Product Code 33631005 M Product Code 33631008 N Product Code 33631102 O Product Code 33631103 P Product Code 33631106 Q Product Code 33631108 R Product Code 33631204 S Product Code 33631306 T Product Code 33631406 | Howmedica Osteonics Corporation | Pfizer Hospital Products Group Howmedics Inc | January 20 1999 | The Inner Dimension Of The Sideplate Had Dimensions That Were Too Small Resulting In The Sideplate Not Sliding Fully Into The Lag Screw | 1 218 Were Distributed | Korea Ireland And Canada
1999-8-25 | Devices | III | z-1202-9 | Series 7000 Total Knee Modular Tibial Tray A Component Of The Total Knee System Implanted Into A Patient Along With Other Knee System Components As Part Of A Knee Replacement | Manufacturer | Howmedica Osteonics Corporation Allendale New Jersey | January 20 1999 | The Tibial Tray Labeled As Catalog No | 6 Devices Were Distributed | Florida Massachusetts Michigan North Carolina New York England
1999-8-25 | Devices | III | z-1203-9 | Cryovalve Allograft Heart Valve | Manufacturer | Cryolife Inc | June 11 1999 | Additional Information For The Organ Procurement Agency Concerning This Donor Revealed That The Donor Possessed Genital Condylomata Genital Warts | 1 Unit | Colorado
1999-8-25 | Devices | III | z-1204/1208-9 | Human Beta 2 Microglobulin Kits Products Are In Vitro Diagnostic Kits And An Antiserum A Product Code No | The Binding Site | The Binding Site Ltd | May 19 1999 | We Are Classifying The Firm's Action As A Voluntary Recall | 87 Kits And 75 Vials Were Distributed | Alabama California Colorado Florida Oklahoma Massachusetts New Jersey New York
1999-8-25 | Devices | III | z- 1212-9 | Architect Free T 3 Reagent Kit A Chemiluminescent Microparticle Immunoassay Cmia For The Quantitative Determination Of Free Triiodothyronine Free T 3 In Human Serum And Plasma | Manufacturer | Abbott Laboratories Abbott Park Illinois | August 4 1999 | Low Recovery When Used With Multi Constituent Controls | 247 Kits Were Distributed | Alabama Florida Illinois Kansas Maryland Michigan New Hampshire New York Ohio Pennsylvania Texas California Germany Japan And Canada
1999-8-25 | Devices | III | z-1215-9 | Kavo Intrasept 905 Intrasurg 500 Dental Operative Unit An Ac Powered Drilling Device Intended To Supply Power To And Serve As A Base For Other Dental Devices Such As A Dental Handpiece During Dental And Maxillofacial Surgery | The Straumann Co | Kavo Dental Gmbh Biberach Riss Germany | June 30 1999 | Software Defect May Cause The Unit To Shutdown Prior To Reaching The Preset Torque Limit | 7 Units Were Distributed | Arizona California Delaware Massachusetts Montana Texas
1999-8-25 | Devices | III | z-1219-9 | Apligraf Graftskin Indicated For Use With Standard Therapeutic Compression For The Treatment Of Non Infected Partial And Full Thickness Skin Ulcers Due To Venous Insufficiency Of Greater Than 1 Month Duration Which Has Not Responded To Conventioanl Therapy | Manufacturer | Organogenesis Inc | July 21 1999 | The Ph Was Found To Be Out Of Specification Range | 58 Units Were Distributed | Alabama Arizona California Georgia Illinois Indiana New York Massachusetts Maryland Maine Michigan Missouri North Carolina Ohio Pennsylvania South Carolina Texas