## 1995-01-25
RECALL_CLASSIFICATION_DATE | PRODUCT_TYPE | CLASSIFICATION | RECALL_NUMBER | PRODUCT | RECALLING_FIRM | MANUFACTURER | RECALL_INITIALIZATION_DATE | REASON | VOLUME | DISTRIBUTION
1995-01-25 | Drugs | II | d-050-5 | Alupent Metaproterenol Sulfate Usp Inhalation Complete With Mouthpiece 200 Metered Doses Rx Used In The Treatment Of Asthma | Boehringer Ingelheim Pharmaceuticals | 3 M Pharmaceuticals Inc | January 9 1995 | Product Does Not Meet Particle Size Specification | 75 034 Units Of Lot 930181 A 71 566 Units Of Lot 930183 A 59 048 Units Of Lot 930184 B 7 845 Units Of Lot 930184 C 12 128 Units Of Lot 930732 B Were Distributed | Nationwide
1995-01-25 | Drugs | III | d-046-5 | Anacin Coated Analgesic Tablets Otc In 2 And 12 Count Tins | Manufacturer | Whitehall Robins Hammonton New Jersey | April 4 1994 | Product Does Not Meet Content Uniformity Specification For Caffeine Ingredien | 113 400 Tablets Were Distributed | Nationwide
1995-01-25 | Drugs | III | d-047-5 | Momentum Muscular Backache Formula Caplets Otc Analgesic In Bottles Of 24 | Manufacturer | Whitehall Robins Division Of American Home Products Corporation Hammonton New Jersey | April 4 1994 | Product Does Not Meet Stability Specification For Disintegratio | 212 688 Tablets Were Distributed Firm Estimates None Remains On Market | Nationwide
1995-01-25 | Drugs | III | d-048-5 | Dristan 12 Hour Nasal Spray In 1 2 Fluid Ounce Bottles Otc Nasal Spray | Whitehall Robins | Whitehall Robins Inc | November 2 1993 | Warning Statement On Immediate Container Label Did Not Appear On Unit Carto | 36 496 Dozen Bottles Were Distributed Firm Estimates None Remains On Market | Nationwide
1995-01-25 | Drugs | III |  | Preparation H Hemorrhoidal Suppositories Otc Rectal Vasoconstrictor Suppository In 12 24 36 And 48 Count Cartons | Manufacturer | Whitehall Robins Division Of American Home Products Corporation Hammonton New Jersey | April 4 1994 | Preservative Ingredient Is Marginally Subpoten | Firm Estimates None Remains On The Market | Nationwide
1995-01-25 | Drugs | III | d-051-5 | Amantadine Hydrochloride Capsules Usp 100 Mg Rx Oral Drug For The Prevention Or Chemoprophylaxis Of And The Treatment Of Respiratory Tract Illness | Udl Laboratories | Chase Laboratories Newark New Jersey | January 12 1995 | Product Does Not Meet Usp Content Uniformity Specification | 5 262 Cartons Of 100 Unit Dose Capsules Were Distributed Firm Estimated That 50 Of The Product Remained On Market At Time Of Recall Initiation | Nationwide
1995-01-25 | Drugs | III | d-052-5 | Regular Strength Acetaminophen Tablets 325 Mg Packaged In 100 And 1000 Tablet Bottles Then Repackaged Under The Goldline And Kerr Labels | Manufacturer | Granutec Inc | December 21 1994 | Presence Of Metallic Particles In The Produc | 6 2 Million Tablets Were Distributed | Florida North Carolina Ohio
1995-01-25 | Devices | II | z-254-5 | Siemens Conventional Mevatron Linear Accelerator System Used To Provide Teletherapy X Ray Beams For The Treatment Of Cancer | Manufacturer | Siemens Medical Systems Inc | June 14 1994 | A Design Defect In The Software Allows The Entry And Acceptance Of Erroneous Treatment Parameters When Using The Mevatron Linac Syste | Approximately 800 Units | Nationwide And International
1995-01-25 | Devices | II | z-258-5 | Rms 2000 Radiation Management System | Manufacturer | Varian Associates Inc | February 19 1993 | A Bug In The System's Novelle Netware Ver | 50 Units | Nationwide And International
1995-01-25 | Devices | II | z-259/263-5 | Clinac Medical Linear Accelerators A Model Cl 12 8 B Model Cl 12 10 C Model Cl 18 20 D Model Cl 1800 E Model Cl 2100 C | Manufacturer | Varian Associates Inc | April 21 1992 | A Carousel Misalignment Caused Several Incidences Of Random Transverse Beam Asymmetr | Approximately 400 Units | Nationwide And International
1995-01-25 | Devices | II | z-252/253-5 | Medrad Disposable Syringes Fluid Delivery Products Packaged As A Ct Tri Paks B 200 Ml Fasturn Syringes | Medrad | Medrad Inc | March 22 1994 | The Sterility Of The Device Has Been Compromised By A Loss Of Package Integrity | 96 Packaged Devices | California Colorado Florida Kentucky Tennessee Utah Virginia
1995-01-25 | Devices | II | z-299-5 | Image Receptor Support Device Irsd Model 4512 590 12681 Used To Provide Breast Cancer Screening And Or Diagnosis Through Radiography Of The Breast | Manufacturer Fda Approved The Firm's Corrective Action Plan December 21 | Lorad Corporation Danbury Connecticut | December 21 1994 | The Fourth Position Of The Automatic Exposure Control Aec Detector In The Irsd Was Insufficiently Leaded And Transmission Limit Exceeded | 24 Units | Nationwide
1995-01-25 | Devices | II | z-301-5 | S Product Auto Suture Premium Ceea Disposable Stapler | Manufacturer | United States Surgical Corporation Norwalk Connecticut | October 7 1994 | The Low Profile Anvil Head Can Come Off During Firin | 11 010 Units | Nationwide And International
1995-01-25 | Devices | III | z-270/274-5 | Allergan Resolve Gp Daily Cleanser For Cleaning Rigid Gas Permeable And Hard Contact Lenses A Resolve Gp Daily Cleaner In 30 Ml Bottles Part 515 B Resolve Gp Daily Cleaner In 5 Ml Bottles C Wet N Soak Kit Part 6635 Contains 5 Ml Bottle D Wet N Soak Kit Part 180 Kt Contains 5 Ml Bottle E Wet N Soak Kit Part 854 Contains 5 Ml Bottle | Allergan | Allergan America Hormigueros Puerto Rico | November 9 1994 | The Resolve Gp Daily Cleaner Was Misformulated To Contain Less Than The Required Amount Of One Of Its Surfactant Ingredients And Does Not Meet Its Physical Appearance Specification | Approximately 293 960 Bottles Were Distributed Firm Estimated That 50 000 Bottles Remained At Time Of Recall Initiation 6 | Nationwide And Russia
1995-01-25 | Devices | III | z-275/277-5 | Elisa Serology Test Kits A Herpes Simplex Virus 1 Igm Elisa Serology Test Kit Catalog 2305450 For The Detection Of Igm Antibodies To Herpes Simplex Virus 1 Antigen In Human Serum B Chlamydia Trachomatis Igg Elisa Serology Test Kits Catalog 2306200 For The Detection Of Igg Antibodies To Chlamydia Trachomatis Antigen In Human Serum C Mumps Igg Elisa Serology Test Kits Catalog 2305900 For The Detection Of Igg Antibodies To Mumps Antigen In Human Serum For Vitro Diagnostic Use | Manufacturer | Clark Laboratories Inc | September 19 1994 | The Absorbance Values For Controls And Specimens Are Lower Than Expecte | A 608 Units B 98 Units C 214 Units Were Distributed | Maryland Florida Greece Turkey Chile Mexico Israel Australia Portugal Spain
1995-01-25 | Devices | III | z-278-5 | Spottest Voges Proskauer Reagent A In Vitro Diagnostic Kit Used For Determining The Ability Of Bacteria To Produce Acetylmethyl Carbinol As An End Product Of Glucose Metabolism | Repacker | Difco Laboratories Mausten Wisconsin Repacker | October 20 1994 | The Label On The Product Incorrectly States The Device Is Spottest Voges Proskauer Reagent | 88 Boxes | Kentucky Texas New Jersey Georgia Florida Illinois Michigan Ireland Finland Greece Taiwan
1995-01-25 | Devices | III | z-279/280-5 | Qbc Centrifuge Power Supply 120 220 Volt Supplied With The Qbc Centrifuge A Qbc Centrifuge Catalog 424740 B Qbc Centrifuge Power Supply Supplied As A Replacement Part Catalog 42474004 | Manufacturer | Becton Dickinson Primary Care Diagnostics Sparks Maryland | October 10 1994 | A Defective Component In The Power Supply May Produce A Visual Flash And Smok | 1 034 Units Were Distributed | Nationwide
1995-01-25 | Devices | III | z-281/282-5 | Precise Hcg Test Kit Used In The Detection Of Human Chorionic Gonadotropin Hcg In Urine Specimens For The Early Detection Of Pregnancy A Catalog 496220 200 And 40 Test Kits B Catalog 496240 200 And 40 Test Kits | Manufacturer | Becton Dickinson Advanced Diagnostics Sparks Maryland | September 26 1994 | False Positive Results In Excess Of Expected Specificity Levels Of 9 | 2 971 Units Were Distributed | Nationwide
1995-01-25 | Devices | III | z-309-5 | Peptostreptococuss Anaerobius Atcc 27337 Bacti Disk An In Vitro Diagnostic Product Recommended For Use In Laboratory Quality Assurance Procedures | Manufacturer | Remel Limited Partnership Lenexa Kansas | July 28 1994 | An Additional Colony Type Peptostreptococcus Spp Was Present Upon Rehydratio | Firm Estimates None Remains On The Market | Nationwide Australia Canada
1995-01-25 | Devices | III | z-310-5 | Loefflers Medium Slant Tubes An In Vitro Diagnostic Used In The Cultivation Of Corynebacterium Diphtheriae | Manufacturer | Remel Limited Partnership Lenexa Kansas | August 1 1994 | The Medium Is A Broth Instead Of A Slan | 140 Tubes Were Distributed Firm Estimates None Remains On The Market | California Texas Missouri Oklahoma
1995-01-25 | Devices | III | z-311-5 | Edms Anaerocult C Mini Set Mini Sachets And Incubation Bags An In Vitro Diagnostic Product Packed 25 Sachets And Incubation Bags Per Box | Remel Limited Partnership | Em Diagnostic Systems Gibbstown New Jersey | May 9 1994 | The Actual Expiration Date Was 6 30 94 And Was Correct On The Outside Box But The Individual Product Was Relabeled With An Expiration Date Of 11 30 9 | 50 Units Were Distributed | Maryland Massachusetts Delaware Michigan Virginia North Carolina California Kansas
1995-01-25 | Devices | III | z-315/316-5 | Fibrinogen Reagents And Veronal Buffer Used For Quantitative Determination Of Fibrinogen In Human Serum For Coagulation Studies A Dade Owren's Veronal Buffer Catalog B 4234 25 B Dade Data F Fibrinogen Determination Reagents Catalog B 4233 15 | Baxter Diagnostic | Baxter Diagnostic Of Puerto Rico Inc | June 19 1994 | When Either Of The Two Lots Of Buffer Were Used As A Diluent In The Fibrinogen Determination The Control Values Were Out Of Range When Read Against A Calibration Curve Constructed With Another Lot Number Of Buffe | A 64 629 Vials B 5 908 Packages Were Distributed | Nationwide And International