## 1995-03-29
RECALL_CLASSIFICATION_DATE | PRODUCT_TYPE | CLASSIFICATION | RECALL_NUMBER | PRODUCT | RECALLING_FIRM | MANUFACTURER | RECALL_INITIALIZATION_DATE | REASON | VOLUME | DISTRIBUTION
1995-03-29 | Devices | II | z-455-5 | Monoject Hypodermic Needles Single Use Epoxy Needle Attachment Polypropylene Hub Luer Lock Hub 19 Ga | Sherwood Medical Company | Sherwood Medical Company Deland Florida | December 21 1994 | The Needles May Contain A Coating Of Silicone Oil In The Interior Of The Needle Hu | 642 600 Needles Of Lot 237593 And 276 900 Needles Of Lot 238057 Were Distributed | Nationwide And Canada
1995-03-29 | Devices | II | z-516-5 | Chemspor Biological And Chemical Indicator A Bi For Monitoring Steam Sterilization Processes In Wet Environments | American Sterilizer Company | Amsco International Inc | February 22 1995 | The Device Is Labeled For 48 Hour Incubation And 510 K Clearance Is For 7 Day | 9 988 Boxes Of Product Were Distributed | Nationwide And International
1995-03-29 | Devices | II | z-517-5 | Proof Flash Self Contained Biological And Chemical Indicator A Bi For Monitoring Steam 270 Xf Steam Gravity Sterilization Cycles | American Sterilizer Company | Amsco International Inc | February 22 1995 | The Product Is Labeled For A 24 Hour Or 48 Hour Incubation Perio | 25 462 Cases 50 Vials Per Box Were Distributed | Nationwide And International
1995-03-29 | Devices | II | z-518-5 | Saxon Gold Ultra Spermicidal Lubricated Latex Condoms With Nonoxynol 9 Packed In Plastic Film | Manufacturer | Safetex Division Of Aladan Corporation Colonial Heights Virginia | September 21 1994 | Condoms Were Found To Have A Lower Concentration Of The Spermicide Nonoxynol 9 Than Was Represented In The Device's Labeling And Premarket Notificatio | 20 Gross Were Distributed Firm Estimates None Remains On The Market | New York Kansas
1995-03-29 | Devices | II | z-540/541-5 | Mrinnervu Disposable Endorectal Coils For Mri Imaging For The Colon And Cervix A Cervix Coil Model Bcr 15 B Colorectal Coil Model Pcc 15 | Manufacturer | Medrad Inc | January 31 1995 | The Packaged Devices Had Broken Shaft | Approximately A 812 Units B 915 Units Were Distributed Firm Estimated That 334 Rectal Units And 252 Cervical Units Remained On Market At Time Of Recall Initiation | Nationwide And International
1995-03-29 | Devices | II | z-542/543-5 | Coveen Urisheath Uriliner Male External Catheters A Pediatric Geriatric Catheter Product 5120 H 25 Mm B Large Catheter Product 5135 H 35 Mm | Coloplast | Coloplast A S Humlebaek Denmark | January 31 1995 | The Labeling Incorrectly Identifies The Product As Non Latex Produc | 3 500 Units Of Lot 86811 And 200 Units Of Lot 86764 Were Distributed | Nationwide
1995-03-29 | Devices | II | z-545/547-5 | Adult Child And Infant Cpr Resuscitator Bags Single Patient Use Disposable With Related Accessories Mask Peep Valve Peep Adapter And Manometer Port | Mercury Medical | Plaxtron Industrial M Sdn | August 29 1994 | The Devices Have Excess Glue Around The Intake Valve On The Interior Of The Ba | 23 610 Units Were Distributed | Nationwide And Puerto Rico
1995-03-29 | Devices | II | z-548/549-5 | Transflator Ii Balloon Inflation Syringes A Transflator Balloon Inflation Syringe Catalog 05 0802 B Transflator Ii Balloon Syringe With Ptca Accessory Kit Catalog 05 0804 | Manufacturer | Surgical Technologies Inc | February 2 1995 | Devices Have Broken Transmitting Led Solder Joints Which May Result In Loss Of Transmission And The Signal Lost Read Out On The Monitor | 1 732 Units Were Distributed Firm Estimates None Remains On The Market | Nationwide Belgium Canada Italy The Netherlands United Kingdom
1995-03-29 | Devices | II | z-550-5 | Medical Research Laboratory Dpd Defibrillator Sold Under Datascope Part 0998 00 0004 Used In Conjunction With Datascope Passport Patient Monitoring System | Datascope Corporation | Medical Research Laboratories Buffalo Grove Illinois | March 2 1995 | There Is A Potential Of The Devices When Used In The Synchronous Mode With The Datascope Passport Monitor For The Purpose Of Cardioversion To Discharge As Late As Approximately 160 Ms After The Peak Of The R Wave Of The Electrocardiogra | 656 Units Were Distributed | Nationwide And International
1995-03-29 | Devices | II | z-558-5 | Encore Orthopedic Foundation Brand Knee Tibial Insert 13 Mm Size 4 | Manufacturer | Encore Orthopedics Inc | October 7 1994 | The Label On The Interior Pouch May Incorrectly Identify The Insert Thickness As 15 Mm However The Actual Thickness Is 13 Mm Which Is Correctly Identified On The Exterior Box Labe | 15 Inserts Were Distributed Firm Estimates None Remain On Market | New York Oklahoma Pennsylvania South Carolina Utah Texas Washington State
1995-03-29 | Devices | II | z-559-5 | Slit Rotation Knob On The Retroskop Cs An Illumination Module Microscope Attachment | Manufacturer | Carl Zeiss Inc | February 6 1995 | The Knob Originally Designed To Be Removable So It Could Be Sterilized Has The Potential To Loosen And Fall Off From The Retroskop After Being Sterilize | 80 Attachments Were Distributed | Nationwide Canada
1995-03-29 | Devices | II | z-565-5 | D 4600 Air Exchange Line Sterile Single Patient Use An Accessory For Ophthalmic Use | Manufacturer | Storz Instrument Company St | November 28 1994 | Due To A Packaging Deficiency The Pouches Were Improperly Sealed Therefore Compromising Sterilit | 2 213 Boxes 10 Pouches Per Box Were Distributed | Nationwide And International
1995-03-29 | Devices | III | z-442/443-5 | MMI Barium Enema Kits For Use In X Ray Diagnosis Of The Lower Gastrointestinal Tract A Catalog 354 00 Assembled Empty Barium Enema Kit B Catalog 381 00 Assembled Filled Barium Enema Kit | MMI | Ridge Products Lakewood New Jersey | May 20 1994 | Devices May Have A Fracture In The Seal At The Base Of The Bag Which Could Result In Leakag | A 242 Kits B 48 Kits Were Distributed | New York Florida Michigan
1995-03-29 | Devices | III | z-502/505-5 | Mdt Mckesson Desert Air Plus Dental Air Compressors A Model 301862 Single Head 115 V B Model 301863 Single Head 230 V C Model 301864 Twin Head 230 V D Model 301865 Triple Head 230 V | Manufacturer | Mdt Diagnostic Company North Charleston South Carolina | June 10 1994 | The Nylon Coalescing Filter Bowl May Break Apart Due To Vibrational Fatigu | A 124 Units B 65 Units C 166 Units D 18 Units Were Distributed | Nationwide And Canada
1995-03-29 | Devices | III | z-506/509-5 | Mdt Mckesson Desert Air Plus Dental Air Compressors A Model 301862 Single Head 115 V B Model 301863 Single Head 230 V C Model 301864 Twin Head 230 V D Model 301865 Triple Head 230 V | Manufacturer | Mdt Diagnostic Company North Charleston South Carolina | December 6 1993 | The Clear Styrene Desiccant Enclosure May Explod | A 123 Units B 65 Units C 165 Units D 16 Units Were Distributed | Nationwide And Canada
1995-03-29 | Devices | III | z-551/552-5 | Bausch Lomb Gold Medalist Toric Visibility Tinted Contact Lenses Labeled With The Following Parameters A Blister Lot Number Sy 009541 Bc 8 | Bausch Lomb | Bausch Lomb Inc | January 24 1995 | An Unidentified Number Of Lenses Are Labeled With The Incorrect Base Curve Sphere Cylinder And Axis Parameter | A 66 Lenses B 71 Lenses Were Distributed | Alaska Kansas Minnesota Virginia California Florida Georgia Illinois Iowa Missouri New Jersey New York Ohio Oregon Wisconsin Canada Australia Ireland
1995-03-29 | Devices | III | z-560-5 | Dade Alcohol And Ammonia Control Level I Normal And Level Ii High | Manufacturer | Baxter Diagnostics Inc | August 19 1994 | The Labeling Was Misleading As To The Ranges Listed In The Package Inser | 29 022 Units Were Distributed | Nationwide And International
1995-03-29 | Devices | III |  | Dade Immunoassay Controls Tri Level Level I Level Ii And Level Iii A Catalog B 5700 5 B Catalog B 5700 6 C Catalog B 5700 7 D Catalog B 5700 8 | Manufacturer | Baxter Diagnostics Inc | October 20 1994 | The Labeling Is Misleading As The Results For The Free T 4 Clinical Assays 2 Step Cal 535 A And Direct Cal 535 S Were Reversed On The Clinical Assay Sheets Provided With The Test Kit | 35 Kits Were Distributed | Texas Michigan California Wisconsin Florida New York
1995-03-29 | Devices | III |  | Frozen Scallops 95 710 318 |  |  | N A | Of The Presence Therein Of Decomposed Shrim |  | N/A