## 1996-02-21
RECALL_CLASSIFICATION_DATE | PRODUCT_TYPE | CLASSIFICATION | RECALL_NUMBER | PRODUCT | RECALLING_FIRM | MANUFACTURER | RECALL_INITIALIZATION_DATE | REASON | VOLUME | DISTRIBUTION
1996-02-21 | Drugs | II | d-086-6 | Dilantin Extended Phenytoin Sodium Capsules Usp Kapseals 100 Mg Antiepileptic Drug In Bottles Of 100 | Parke Davis | Warner Lambert Company Fajardo Puerto Rico | December 12 1995 | Product May Not Meet Dissolution Specifications Through Expiration Dat | 39 846 Bottles Were Distributed | Nationwide
1996-02-21 | Drugs | III | d-085-6 | Panama Jack Brand Dark Tanning Lotion Waterproof 4 Spf In 3 Fluid Ounce Plastic Bottles | Manufacturer | Creative Cosmetics Rockledge Florida | April 3 1995 | Product Does Not Meet Manufacturer's Microbial Specification | 7 344 Units Were Distributed | Nationwide
1996-02-21 | Devices | II | z-326/329-6 | Multinex A Breath By Breath Gas Analyzer That Uses Microprocessor Technology To Calculate And Display End Tidal Co 2 Inspired Co 2 Inspired And Expired N 2 O Inspired 02 Inspired And Expired Anesthetic Agents Respiration Rate Sa 2 And Pulse Rate A Multinex Plus Model Number 0998 00 0071 13 14 15 16 B Multinex Plus W Id Model 0998 00 0071 17 18 19 20 C Multinex Model Number 0998 00 0071 01 02 03 04 D Multinex Id Model 0998 00 0071 09 10 11 12 | Manufacturer | Datascope Corp | August 31 1995 | The 4610 Gas Analyzer Board And The 4640 Gas Identification Board Have Been Found To Have Defects Which May Cause A Random And Intermittent Failure Message Along With Either A Temporary Loss Of Gas Monitoring Or A Disabling Of The Anesthesia Identification Functio | 1 248 Units Were Distributed | Nationwide And International
1996-02-21 | Devices | II | z-341-6 | Monogram Instrument Tibial Baseplate Impactor Extractor Used To Install And Remove The Tibial Baseplate | Manufacturer | Howmedica Inc | March 14 1995 | The Product's Threaded Shaft Was Breaking During Impactio | 408 Units Were Distributed | Nationwide And International
1996-02-21 | Devices | II | z- 344/345-6 | A Alta Tibial Rod Driver Locking Bolt Bolt Only Used For The Insertion Of The Alta Tibial Rod B Alta Tibial Rod Driver And Bolt Driver Bolt Assembly Meant To Absorb The Impact Of The Blows From The Mallet During Insertion Of The Tibial Rod | Manufacturer | Howmedica Inc | July 29 1994 | The Locking Bolt Was Breaking Near The Threaded End Smooth Shaft Interface During Surgical Procedure | 709 Units Were Distributed Firm Estimates None Remains On The Market | Nationwide And International
1996-02-21 | Devices | II | z-374/376-6 | Barrier Vu Thru Cardiovascular Pack A Sterile Barrier Vu Thru Cardiovascular Pack Containing In Part A C V Split Sheet Product Reorder Number 127 B Sterile Barrier Vu Thru Cardiovascular Pack Ii Containing In Part A C V Split Sheet Product Reorder Number 0173 C Bulk Non Sterile Barrier Vu Thru C V Split Sheet Product Reorder Number 9573 | Johnson Johnson Medical J J | Johnson Johnson Medical Inc | December 8 1995 | The Fabric Of The Clipping Tabs Are Subject To Tearing Presenting Potential Harm To Patien | 569 Cases 2 770 Units Were Distributed | Nationwide And International
1996-02-21 | Devices | II | z-379/380-6 | A Gesco Brand 3 | Medchem Products | Gesco International Inc | May 19 1995 | The Catheter Was Manufactured Without A Radiopaque Stripe Thus Precluding Radiographic Confirmatio | A 900 Units B 700 Units Were Distributed Firm Estimates None Remains On Market | Nationwide
1996-02-21 | Devices | II | z-408-6 | Rite Angle Disposable Prophy Angle With Soft Cup Used In Performing Dental Prophylaxis | Dentsply International | Team Technologies Inc | April 26 1995 | In Some Cases The Device Has Demonstrated A Propensity For Stalling Stopping Falling Apart Vibrating Freezing Up And Over Heating During Usag | 112 320 Were Distributed | Nationwide
1996-02-21 | Devices | II | z-409-6 | Power Sonic Brand Battery Packs Used In Uro View Or Uroview Model 2000 X Ray Imaging Systems | Manufacturer | Oec Medical Systems Inc | December 14 1995 | The Replacement Battery Pack Could Have Been Damaged During Shipment Which Could Possibly Lead To Premature Failure And In The Worst Case A Change Of A Battery Fir | 8 Units | Florida Illinois Kentucky North Carolina New York Ohio Virginia Washington D
1996-02-21 | Devices | II | z-411-6 | Dextrolyte Peritoneal Dialysis 5 Prong Cycler Set | National Medical Care | National Medical Care Reynosa Reynosa Mexico | May 26 1995 | The Connector Lines Are Misassemble | 2 990 Sets Were Distributed | Nationwide
1996-02-21 | Devices | II | z-412-6 | Baxter Arterial Bloodline Used For Dialysis | National Medical Care | National Medical Care Reynosa Reynosa Mexico | June 2 1995 | The Arterial Bloodlines Were Packaged In Cartons Labeled As Venous Bloodline | 725 Cases Were Distributed | Nationwide
1996-02-21 | Devices | II | z-429/430-6 | A Stryker Universal Cement Restrictor With Inserter Part 206 701 Intended To Be Implanted During Total Joint Or Joint Revision Arthroplasty B Stryker Sterile Disposable Cement Sculps Part 206 716 Intended To Be Used In Joint Or Joint Revision Arthroplasty To Remove Excess Bone Cement From Around The Prosthesis | Stryker Instruments | Biomedical Devices Irvine California | August 24 1995 | The Sterility Of The Product Cannot Be Guaranteed Because The Packaging May Be Compromise | A 60 Boxes 6 Units Per Box B 32 Boxes 6 Units Per Box Were Distributed | Nationwide
1996-02-21 | Devices | II | z-433-6 | Belmont Acuray 071 A Dental Periapical X Ray System | Belmont Equipment Corporation | Takara Medical Company Kobe Japan | September 20 1995 | The Horizontal Arm Of The Device May Separate From Its Wall Mount Bracket And Fal | 3 851 Units Were Distributed | Nationwide And Canada
1996-02-21 | Devices | II | z-435/440-6 | Radiographic Intensifying Screens Intended To Expose Medical Radiographic Film A Model Trimax T 2 B Model Trimax T 6 C Model Trimax T 12 D Model Trimax T 4 E Model Trimax T 8 F Model Trimax T 6 12 | 3 M Medical Imaging Systems Division | 3 M Italia Spa Viale Della Liberta Italy | November 3 1995 | The Intensifying Screens Will Chemically React With Radiographic Film And Become Heavily Discolore | 39 150 Screens Were Distributed | Nationwide And International
1996-02-21 | Devices | II | z-443-6 | Prospect Tm Entamoeba Histolytic Microplate Assay Kit Used For Detection Of Entamoeba Histolytica Specific Antigens | Manufacturer | Alekon Biomedical Inc | March 23 1995 | The Test Could Provide A False Positive Result For Entamoeba Histolytic In Pediatric Sample | 281 Kits Were Distributed | Nationwide And International
1996-02-21 | Devices | III | z-385/386-6 | Ati Eo Gas Disposable Test Pack With Ati Test Biological Indicator A Disposable Test Pack With A Self Contained Biological Indicator Used To Monitor Ethylene Oxide Sterilization Cycles A Catalog 00234 00234 Bx B Catalog 23404 23404 Bx | Pymah Corporation | Pymah Corporation North Hollywood California | March 27 1995 | Devices Were Found Not To Meet Labeling Claims Of An 18 Month Shelf Lif | 677 25 Cases Firm Estimated That 200 Cases 64 Packages Per Case Remained On Market At Time Of Recall Initiation | Nationwide Canada Germany Spain Ireland Poland Turkey
1996-02-21 | Devices | III | z-410-6 | Ektachem Clinical Chemistry Slides Digoxin Dgxn For In Vitro Diagnostic Use To Quantitatively Measure The Drug Digoxin In Serum And Plasma | Manufacturer | Johnson Johnson Clinical Diagnostics Inc | July 17 1995 | The Digoxin Slides May Exhibit A Positive Bias Overtime On The First Slide From Each Cartridg | 14 904 Cartridges Were Distributed | Nationwide And International
1996-02-21 | Devices | III | z-434-6 | Caulk Friction Grip Carbide Super Burs Used For Cutting Teeth | Manufacturer | Dentsply International Caulk Division York Pennsylvania | March 28 1995 | The Device Is Subject To Dulling Of The Cutting Edges Or The Bur Tip And Possible Breakage In Use At The Weld Locatio | 537 000 Were Distributed | Nationwide And International
1996-02-21 | Devices | III | z-444-6 | Ektachem Clinical Chemistry Slides For Cholesterol Chol For In Vitro Diagnostic Use To Quantitatively Measure Cholesterol In Serum Plasma | Johnson Johnson Clinical Diagnostics | Johnson Johnson Clinical Diagnostics Inc | November 29 1993 | There Was A Drift In Cholesterol Results Within The Same Day Caused By Slide Temperature And Humidit | 1 336 000 Cartridges Were Distributed | Nationwide And International
1996-02-21 | Devices | III | z-445-6 | Kodak Ektachem Clinical Chemistry Slides Trig For In Vitro Diagnostic Use For The Quantitative Measurement Of Triglycerides Concentration In Serum Or Plasma | Johnson Johnson Clinical Diagnostics | Johnson Johnson Clinical Diagnostics Formerly Eastman Kodak Company Clinical Diagnostics Division Rochester New York | May 18 1995 | Triglyceride Slides On Analyzer Stability Did Not Support A Four Week Expiration Clai | Approximately 760 000 Cartridges Were Distributed 8 | Nationwide And International
1996-02-21 | Devices | III | z-446-6 | Opus Total B Hcg Test Modules For In Vitro Diagnostic Use For The Quantitative Measurement Of Intact And Free Beta Subunits Of Human Chorionic Gonadotropin Hcg In Serum Plasma And Urine | Manufacturer | Behring Diagnostics Inc | January 30 1996 | Insufficient Conjugate In The Wells May Cause False Negatives Or Lower Than Expected Result | 409 Boxes 50 Units Per Box Were Distributed | Nationwide And International
1996-02-21 | Devices | III | z-447-6 | Ilab Il Test Tm Albumin Reagent Intended For The Quantitative In Vitro Diagnostic Determination Of Serum Albumin Using The Ilab Chemistry Systems | Instrumentation Laboratory Company | Instrumentation Laboratory Company Orangeburg New York Responsible Firm | January 14 1994 | The Albumin Reagent Had Mold Growth And Reagent Deterioratio | 2 Kits Were Distributed | Canada
1996-02-21 | Devices | III | z-448-6 | Monarch Il Test Albumin Reagent For The Quantitative In Vitro Diagnostic Determination Of Serum Albumin Using The Il Monarch R Chemistry Systems | Instrumentation Laboratory Company | Instrumentation Laboratory Company Ilc Orangeburg New York Responsible Firm | January 19 1994 | The Albumin Reagent Had Mold Growth And Reagent Deterioratio | 704 Kits Were Distributed | Florida Oregon Texas Kansas Louisiana Wisconsin Missouri Pennsylvania Ohio Kentucky Michigan Arkansas