## 1996-05-01
RECALL_CLASSIFICATION_DATE | PRODUCT_TYPE | CLASSIFICATION | RECALL_NUMBER | PRODUCT | RECALLING_FIRM | MANUFACTURER | RECALL_INITIALIZATION_DATE | REASON | VOLUME | DISTRIBUTION
1996-05-01 | Drugs | II | d-127/128-6 | Cephradine Capsules Usp Rx Antibiotic A 250 Mg In Bottles Of 500 Under The Rugby Label B 500 Mg In Bottles Of 100 Under The Biocraft Label | Biocraft Laboratories | Biocraft Laboratories Inc | October 6 1995 | Potency Not Assured Through Expiration Dat | A 504 Bottles B 5 467 Bottles Were Distributed Firm Estimated That Less Than 5 Percent Of Product Remained On Market At Time Of Recall Initiation 4 | Georgia
1996-05-01 | Drugs | II | d-129-6 | Alupent Metaproterenol Sulfate Usp Inhalation Aerosol Refill 200 Metered Doses Rx Bronchodilator | Boehringer Ingelheim Pharmaceuticals | 3 M Pharmaceuticals Northridge California | April 9 1996 | A Single Unit Of Atrovent Aerosol Was Found Within Lot 951051 | 8 413 Canisters Were Distributed | Nationwide
1996-05-01 | Drugs | III | d-130-6 | Niferex Forte Elixir Rx In 4 Ounce Bottles Indicated For Iron Deficiency Anemia | Manufacturer | Central Pharmaceuticals Inc | December 29 1995 | Subpotency Of Vitamin B 12 Ingredien | Firm Estimates None Remains On The Market | Nationwide And Haiti
1996-05-01 | Devices | II | z-542/545-6 | Extended Stay Stretchers And Warming Stretchers 5 A Stryker Medical Thermaltek Series Emergency Care Warming Stretcher Model 1030 B Stryker Medical Thermaltek Series Pacu Warming Stretcher Model 1530 C Stryker Patient Handling Synergy Series Extended Stay Stretcher Model 1050 D Stryker Patient Handling Synergy Series Extended Stay Stretcher Model 1550 | Stryker Patient Handling | Belden Wire Cable Company Cord Products Division Carmel Indiana Power Cords Stryker Medical Division Of Stryker Corporation Kalamazoo Michigan Stretchers | December 4 1995 | The Outer Casings Of The Power Cords On These Electrical Powered Stretchers Are Cracking Because The Insulation Used On The Outer Jacket Was Not Compatible With The Insulation Used On The Inner Conductor | 949 Stretchers Were Distributed | Nationwide Canada Colombia The Netherlands Singapore
1996-05-01 | Devices | II | z-585-586-6 | Mevasim And Mevasim S Radiation Therapy Simulator Systems A Medical Linear Accelerator Used For The Treatment Of Superficial Shallow And Deep Seated Tumors | Siemens Medical Systems Oncology Care Systems | Mecaserto Radiotherapie France | June 14 1995 | The Mevasim And Mevasim S Film Grid Assemblies Have The Potential To Fall From The Cassette Holder When The Gantry Is Rotated To The Lateral Position And The Mevasim S Field Collimating Blades Have The Potential Of Slowly Drifting When The Gantry Is In The Lateral Positio | 279 Units Were Distributed 6 | Nationwide And International
1996-05-01 | Devices | II | z-587-6 | Vericord Vs 3 Radiotherapy Verification And Record System Used With Model Sl 75 5 Linear Accelerators Used To Treat Superficial Shallow And Deep Seated Tumors | Philips Medical Systems | Philips Medical Systems Crawley West Sussex England | December 18 1995 | Software Version | 16 Units Were Distributed | Florida New Jersey Ohio Pennsylvania Canada
1996-05-01 | Devices | II | z-597/600-6 | Genesis Ii Resuscitators A Genesis Ii Adult Child Automatic Time Volume Cycled Resuscitator With Masks Part 155020 Ot B Genesis Ii Adult Only Automatic Time Volume Cycled Resuscitator With Mask Part No | Nellcor Puritan Bennett | O Two Systems International Inc | January 8 1996 | The Flapper Valve Of The Actuating Diaphragm Of The Demand Valve Had Been Partially Sucked Into The Diaphragm Insert Causing The Automatic Shutoff Valve To Function Thereby Shutting Off All Flow Through Both Manual And Automatic Circuit | 168 Units Were Distributed | Missouri Iowa Massachusetts Alabama California Mexico Peru Ireland United Kingdom
1996-05-01 | Devices | II | z-608/609-6 | Natural Knee Tibial Insert Used In Conjunction With Intermedics Natural Knee System To Replace A Diseased Or Otherwise Dysfunctional Knee Joint A Natural Knee Size A Congruent Tibial Insert W Screw Left 11 Mm Catalog No | Manufacturer | Intermedics Orthopedics Inc | March 18 1996 | A Package Labeled As An 11 Mm Left Congruent Insert Actually Contained A 22 Mm Right Ultra Congruent Inser | A 7 Units B 9 Units Were Distributed | California Colorado Florida Michigan Mississippi Ohio Pennsylvania Tennessee Texas Wisconsin
1996-05-01 | Devices | II | z-612/626-6 | Ercp Endoscopic Retrograde Cholangiopancreatography Cannulas Diagnostic Devices A Product No | Manufacturer | Mill Rose Laboratories Inc | November 8 1994 | The Depth Markings On The Cannula Tip Slough Off Into The Patient During Invasive Surgical Procedure | 2 394 Cannulas Were Distributed | Nationwide
1996-05-01 | Devices | II | z-629-6 | Starplex 60 Ml Jr | Laboratory Product Sales | Starplex Scientific Etobicoke Ontario Canada | July 20 1995 | Some Of The Lots In Distribution Had Not Been Irradiated With The Correct Dos | Undetermined | New York
1996-05-01 | Devices | II | z-630-6 | Abiomed Bvs 5000 Bi Ventricular Support System Solenoid Valve In Bvs Console 9 Catalog 0005 0000 A Mechanical Circulatory Support System For Use In Patients Suffering From Postcardiotomy Ventricular Dysfunction | Manufacturer | Abiomed Inc | March 22 1996 | The Solenoid Valves Used In The Bvs Console Are Failing In The Closed Position Causing Low Pressure Low Flow Alarms To Activat | 139 Valve Assemblies 2 Valves In Each Assembly Are In Commerce | Nationwide And International
1996-05-01 | Devices | II | z-631-6 | Laser Treatment Systems Used For Hair Removal Recall Z 631 6 | Manufacturer Fda Approved The Firm's Corrective Action Plan April 16 | Lorad Corporation Danbury Connecticut | April 16 1996 | The Devices Failed To Comply With 21 Cfr 104 | 9 Units | Nationwide
1996-05-01 | Devices | II | z-634-6 | Corflo 300 Enteral Feeding Pump A Volumetric Rotary Peristaltic Intermittent Pump Designed For Both Adult And Pediatric Enteral Feeding | Corpak Medsystems | Corpak Medsystems Wheeling Illinois Pump Illini Technology Inc | March 27 1996 | Some Of The Devices May Have Defective Ic Chips Which May Cause The Pump To Malfunctio | 181 Pumps Were Distributed With The Firm Estimating That All Of The Units Remain On The Market | Nationwide Canada
1996-05-01 | Devices | II | z-635-6 | Mitochondria Eia Test Kit For In Vitro Diagnostic Use Catalog No | Manufacturer | Clark Laboratories Inc | March 20 1996 | The Product Insert Packaged With The Kits Stated An Incorrect Length Of Time To Run The Tes | 64 Kits Were Distributed | California Maryland Pennsylvania Tennessee International
1996-05-01 | Devices | II | z-636/637-6 | Rolls Brand Manual Wheelchairs A Model 2000 Series B Model 9000 Series | Manufacturer | Invacare Corporation North Ridgeville Ohio | October 19 1995 | The Front Steel Caster Forks Which Support The Chair Wheels Have The Potential For Buckling Under Stress Which Could Cause The Chair To Tip Over And The User To Potentially Fall Out Of The Wheelchai | 118 Chairs Were Distributed | Nationwide
1996-05-01 | Devices | II | z-647/695-6 | Various Custom Procedure Trays Containing Barrier Vu Thru Cardiovascular Split Sheet Labeled With A Unique Work Order Number Used For Various Procedures Ranging From General Surgery To Open Heart Surgery | Manufacturer | Maxxim Medical Doing Business As Sterile Design Clearwater Florida | December 17 1995 | The Fabric Of The Clipping Tabs Of The Split Sheet Are Subject To Tearing Which May Result In Unsecured Line | 803 Trays Were Distributed | California Pennsylvania Florida Michigan Texas Oklahoma Virginia Illinois Missouri New York Washington State Oregon Indiana Maryland New Jersey Illinois
1996-05-01 | Devices | III | z-627-6 | Drs 4 Dialyzer Reprocessing System Eprom Esio Software Version Sxd 329 | Seratronics | Fresenius Usa Inc | May 11 1995 | A Flaw In The Software May Cause The Machine To Lock Up During Self Test Upon Power U | 39 Units Were Distributed 13 | Nationwide
1996-05-01 | Devices | III | z-628-6 | Dulbecco's Phosphate Buffered Saline Dpbs In 500 Ml And 1000 Ml Bottles Sterile For Cell Culture In Vitro Diagnostic Use | Manufacturer | Jrh Biosciences Lenexa Kansas | March 26 1996 | The Above Lot Of Material May Have Units Which Do Not Contain Dulbecco's Phosphate Buffered Saline But Contain Wate | 306 500 Ml And 4 1000 Ml Bottles Were Distributed | California Virginia Connecticut New Jersey Massachusetts Maryland Pennsylvania Texas
1996-05-01 | Devices | III | z-633-6 | Daily Wear Polymacon Soft Hydrophilic Contact Lenses | Manufacturer | Osi Corporation San Francisco California | November 22 1995 | The Lenses Are Labeled As Polymacon 38 Water When They May In Fact Be Ocufilcon D 55 Wate | 2 159 Units Were Distributed Firm Estimated That 1 245 Units Remained In Commerce As Of December 1995 | Nationwide
1996-05-01 | Devices | III | z-638/639-6 | Ectra Retrograde Knives Packaged As Follows A Ectra Ii Disposable Procedure Kit Sterile Part No | Manufacturer | Smith Nephew Endoscopy Andover Massachusetts | March 20 1996 | The Tips Are Subject To Breakage During Us | Firm Estimated That 4 300 Units Remained In Distribution At Time Of Recall Initiation | Nationwide And International