## 1996-10-22
RECALL_CLASSIFICATION_DATE | PRODUCT_TYPE | CLASSIFICATION | RECALL_NUMBER | PRODUCT | RECALLING_FIRM | MANUFACTURER | RECALL_INITIALIZATION_DATE | REASON | VOLUME | DISTRIBUTION
1996-10-22 | Drugs | I | d-254-6 | All Sizes And Strengths Of Adrenal Cortex Injection A Rx Injectable Drug Labeled As Distributed By Hallmark Labs Chicago Il | Phyne Pharmaceuticals | Primary Ontor Beauty Products Inc | October 1 1996 | Mycobacterium Abscessus Contamination Of Unopened Vial Of Hallmark Labs Labeled Product Distributedd By Phyne Pharmaceuticals And Epidemiological Data Collected By Cdc Correlating The Use Of This Product With Mycobacterium Abscesses At The Site Of Injectio | Undetermined | Nationwide
1996-10-22 | Devices | II | z-001-7 | Gemini Pc 1 Volumetric Infusion Pump Controller European 220 Volt Model V 8 | Manufacturer | Imed Corporation San Diego California | June 18 1996 | The Pumps Contain A Faulty Toroid Inductor Which May Heat Up Adjacent Wires Causing An Electrical Shor | 645 Units Were Distributed | International
1996-10-22 | Devices | II | z-007/014-7 | Max I Probe Endodontic Irrigation Syringe A Catalog Maxe 125 3 Cc Syringe 25 G Probe 1 2 Naoc 1 Solution B Catalog Maxe 625 3 Cc Syringe 25 G Probe 4 6 Naoc 1 Solution C Catalog Maxe 628 3 Cc Syringe 28 G Probe 4 6 Naoc 1 Solution D Catalog Maxe 630 3 Cc Syringe 30 G Probe 4 6 Naoc 1 Solution E Catalog Maxe 650 Trial Kit Of 2 Trays 1 Tray Of 3 Cc Syringes And 1 Tray Of 10 Cc Syringes With Each Tray Containing 1 25 G 28 G And 2 30 G Probes With 4 6 Naoc 1 Solution F Catalog Maxe 825 3 Cc Syringe 25 G Probe 2 3 Naoc 1 Solution 5 G Catalog Maxe 828 3 Cc Syringe 28 G Probe 2 3 Naoc 1 Solution H Catalog Maxe 830 3 Cc Syringe 30 G Probe 2 3 Naoc 1 Solution | Manufacturer | Mpl Technologies Inc | September 9 1996 | The External Finger Grip Flanges May Separate From The Barrel Of The Syringe Rendering The Syringe Inoperabl | 12 455 Syringes Were Distributed | Nationwide And Belgium
1996-10-22 | Devices | II | z-017-7 | Lifesaver Portable Oxygen Units Catalog No | Manufacturer | Hudson Respiratory Care Inc | June 14 1996 | Device Was Distributedd With A Regulator With A Flow Range Of 1 8 Lpm Instead Of The Correct Regulator With A Flow Range Of 2 15 Lp | 42 Units Were Distributed | California Texas Illinois Pennsylvania Utah Missouri Massachusetts Canada Honduras
1996-10-22 | Devices | II | z-018-7 | The Thairapy Vest Airway Clearance System Used To Provide Airway Clearance Therapy | Manufacturer | American Biosystems Inc | February 8 1996 | A Shorting In An Output Transistor Could Allow The Device Blower To Operate At Pressures Higher Than Specification Resulting In Vest Pressures Constricting The Thorax And Impairing Inspiratio | 173 Devices Were Distributed | Nationwide
1996-10-22 | Devices | II | z-019-7 | Low Porosity Woven Vascular Prosthesis Catalog Olp 1206 Indicated For Replacement Or Bypass Procedures In Aneurysmal And Occlusive Disease Of The Abdominal Aorta Visceral Arteries And Proximal Peripheral Arteries Exclusive Of The Coronary Arteries | Manufacturer | Intervascular Inc | August 21 1996 | One Thread Is Missing In The Involved Woven Vascular Prothesis Which May Potentially Increase Water Permeability Along The Line Of The Pulled Threa | 65 Unimplanted Devices Were In Commerce At Time Of Recall Initiation | Tennessee New York Louisiana Illinois Alabama Florida California Pennsylvania Arizona Georgia Texas New Jersey Ohio Japan
1996-10-22 | Devices | II | z-020-024-7 | Input Ps Catheter Introducer Sheath With Hydro Pel Coating Used To Facilitate Placing A Catheter Through The Skin Into A Vein Or Artery A 5 F 1000511 5 F 11 Cm Input 2 Freeway 050011 5 F 11 Cm Input Introducer Intl | Usci Manufacturing Facility | C R Bard Irl | August 13 1996 | Introducer Hub Is Separating At Or Near The Weld Site Of The Upper And Lower Hub Body Part | Approximately 980 539 Units Were Distributed | Nationwide And International
1996-10-22 | Devices | II | z-027/028-7 | Microvision C Radiographic Film Used For Mammography | Sterling Diagnostic Imaging | Sterling Diagnostic Imaging Inc | August 30 1996 | The Devices Have Demonstrated Uncharacteristic Aging That Is Below Working Limits For Speed For This Product Which May Result In The Mis Diagnosis In Mammography Reading | Firm Estimated That 19 388 Boxes Of Film Remained On Market At Time Of Recall Initiation | Nationwide Mexico Brazil Latin America
1996-10-22 | Devices | II | z-1216/1218-6 | A Mobilaire Iii Oxygen Concentrator Model No | Manufacturer | Invacare Corporation Sanford Florida | March 8 1996 | Kinked Oxygen Transfer Hoses In The Units Caused By Defective Exterior Foam Insulation May Affect The Flow Of Oxygen And Therefore Affect The Concentration Of Oxygen The Patient Receive | 193 Units Were Distributed | Nationwide