## 1996-11-20
RECALL_CLASSIFICATION_DATE | PRODUCT_TYPE | CLASSIFICATION | RECALL_NUMBER | PRODUCT | RECALLING_FIRM | MANUFACTURER | RECALL_INITIALIZATION_DATE | REASON | VOLUME | DISTRIBUTION
1996-11-20 | Drugs | III | d-035-7 | Pseudoephedrine Hcl Adrenergic Bulk Pharmaceutical Non Rx Under The Following Labels China Meheco Zhuhai Import Export Company Buckton Scott Limited And H | H Reisman Corporation | China Meheco Zhuhai Import Export Company China Bulk Manufacturer | May 24 1995 | Pseudoephedrine Hcl Was Mistakenly Mixed With Ephedrine Hc | 145 Kg Were Distributed Firm Estimated That 10 Percent Of Product Remained On Market At Time Of Recall Initiation | Michigan Texas
1996-11-20 | Drugs | III | d-036-7 | Nature's Formula Herbal Espresso In 1 Fluid Ounce 30 Ml Plastic Bottles Dietary Supplement | Manufacturer | Efs Herbal Drops San Diego California | September 18 1996 | Unapproved New Drug Product Labeled For Use As A Bronchodilato | Approximately 3 609 Bottles Were Distributed | Arizona Arkansas California Florida Louisiana Minnesota Mississippi Nevada Oklahoma Tennessee Texas Utah Washington State Wisconsin Canada
1996-11-20 | Drugs | III | d-037-7 | Udl Laboratories Brand Amantadine Hcl Capsules Usp 100 Mg Unit Dose Packaged In Blister Pack Strips Of 10 Rx Indicated In The Treatment Of Idiopathic Parkinson's Disease | Udl Laboratories | Banner Pharmacaps Formerly Known As Chase Laboratories Inc | October 30 1996 | Dissolution Failure At 9 Month Stability Timepoin | 2 093 Unit Cartons Were Distributed Firm Estimated That 20 Percent Of Product Remained On Market At Time Of Recall Initiation | Nationwide
1996-11-20 | Drugs | III | d-038-7 | Minocycline Hydrochloride Capsules Usp 50 Mg In Bottles Of 100 | Warner Chilcott | Warner Lambert Company Lititz Pennsylvania Responsible Firm | October 18 1996 | Lot Failed Content Uniformity Test One Capsule Was Less Than 67 Ls Rsd Of 10 Is 13 | 4 767 Bottles Were Distributed | Nationwide
1996-11-20 | Devices | II | z-134-7 | Servo O 2 Oxygen Control System Used To Provide Precise Control Of Oxygen Concentration For Infants In The Care Plus Incubator Or Under Oxygen Hoods | Ohmeda | Bpr Medical Systems Mansfield Notts United Kingdom | October 14 1996 | The Low Battery Audible Alarm May Not Sound Prior To Battery Depletion And Subsequent Device Shutdown And The Nominal Time Of Operation Remaining After Activation Of The Visual Low Battery Alarm Is Seven Minutes Rather Than The Specified Nominal Value Of Ten Minute | 33 Units Were Distributed | Tennessee California
1996-11-20 | Devices | III | z-090/095-7 | Coulter Md Automated Hematology Analyzers For In Vitro Diagnostic Testing 6 A Md 8 Model No | Manufacturer | Coulter Corporation Miami Florida | February 18 1994 | The Models With Software Version 1 F 1 Or Lower Could Report Erroneous Platelet Results On Specimens That Exceed The Linearity Rang | 1 353 Units Were Distributed | Nationwide
1996-11-20 | Devices | III | z-096-7 | Coulter Onyx Automated Differential Cell Counters Part 6705522 Model Numbers 0690 And 0692 | Manufacturer | Coulter Corporation Miami Florida | February 15 1994 | Software Program Level | 38 Units Were Distributed | Nationwide
1996-11-20 | Devices | III | z-097/098-7 | Coulter Maxm And Maxm With Autoloader Used As A Automated Hematology Cell Counters And Differential Cell Counter For In Vitro Diagnostic Use A Maxm Model 0041 B Maxm With Autoloader Model 0056 | Manufacturer | Coulter Corporation Miami Florida | April 26 1994 | Software Failure Could Report An Erroneous Platelet Resul | 953 Units Were Distributed | Nationwide
1996-11-20 | Devices | III | z-099/102-7 | Coulter Maxm Cell Counters Automated Hematology Cell Counters And Differential Cell Counter For In Vitro Diagnostic Use A Maxm Model No | Manufacturer | Coulter Corporation Miami Florida | June 15 1994 | The Reagent Can Run Out Without Giving Warning Caused By A Capacitor That Has The Polarity Reverse | 524 Units Were Distributed | Nationwide
1996-11-20 | Devices | III | z-103/119-7 | Coulter Automated Differential Cell Counters A Coulter S 550 B Coulter S 560 C Coulter S 770 D Coulter S 790 E Coulter S 7120 F Coulter S 880 G Coulter S 880 Micro H Coulter Jr Jr With Ir And Js I Coulter St J Coulter Stkr K Coulter Stks L Coulter S Plus I M Coulter S Plus Ii N Coulter S Plus Iii O Coulter S Plus Iv S Plus Iv With Diff P Coulter S Plus V Q Coulter S Plus Vi Recall Z 103 119 7 | Manufacturer | Coulter Corporation Miami Florida | January 19 1995 | The Diluent Pickup Tube Assembly May Not Ensure Sufficient Diluent Levels When The Low Level Alarm Is Triggere | 6 187 Units Were Distributed | Nationwide
1996-11-20 | Devices | III | z-120-7 | Automated Differential Cell Counters Manufactured And Marketed By Coulter Under The Names Coulter Onxy Model Number 0692 | Manufacturer | Coulter Corporation Miami Florida | November 28 1994 | Certain Analog Boards Produced An Elevated Frequency Of Aperture Error 16 During Control Recover | 2 Units Were Distributed | Illinois
1996-11-20 | Devices | III | z-121/124-7 | Coulter Onyx Automated Differential Cell Counters A Onyx Monochrome Model No | Manufacturer | Coulter Corporation Miami Florida | July 7 1995 | Product May Dilute Patient Samples In The Closed Vial Mode When The Needle Vent Line Becomes Obstructe | 701 Units Were Distributed | Nationwide
1996-11-20 | Devices | III | z-125/127-7 | Coulter Md Onxy Automated Differential Cell Counters A Onyx Model No | Manufacturer | Coulter Corporation Miami Florida | November 20 1995 | Software Errors Can Cause Misidentification Of Patient Results And Or Incorrect Transmission Of Specific Default Values For Eosinophil And Basophi | 786 Units Were Distributed 9 | Nationwide
1996-11-20 | Devices | III | z-128-7 | Coulter Cytotrol Control Cells Part 6604248 Reagent Used To Assess The Activity Of Monoclonal Antibodies By Flow Cytometry | Manufacturer | Coulter Corporation Miami Florida | December 7 1995 | An Increase Of Rbc Debris Was Found In The Control Cells Which Can Lead To An Inability To Achieve Expressed Assay Value | 90 Kits Were Distributed | Nationwide
1996-11-20 | Devices | III | z-129-7 | 4 C Plus Coulter Counter Cell Control Part 7546844 Used For Monitoring The Precision Of Coulter Instruments | Manufacturer | Coulter Corporation Miami Florida | May 6 1996 | Low Rdw Recovery For Abnormal High 4 C Plus Cell Control Can Be Obtaine | 61 Kits Were Distributed | Nationwide
1996-11-20 | Devices | III | z-130-7 | Sportrol Growth Promotion Test Suspensions Containing Clostridium Sporogenes 10 Ml Of Suspension Distributed In Plastic Vials One 1 Vial Per Package For Usp Growth Promotion Testing | Manufacturer | North American Science Associates Inc | October 29 1996 | There Was Reduction In The Spore Population Counts Below The Labeled Rang | 53 Vials Were Distributed | Nationwide Ireland Singapore Japan