## 1997-03-12
RECALL_CLASSIFICATION_DATE | PRODUCT_TYPE | CLASSIFICATION | RECALL_NUMBER | PRODUCT | RECALLING_FIRM | MANUFACTURER | RECALL_INITIALIZATION_DATE | REASON | VOLUME | DISTRIBUTION
1997-03-12 | Biologics | II | b-385-7 | Platelets | Manufacturer | Central Indiana Regional Blood Center Inc | June 11 1996 | Blood Products Were Prepared From Units Of Whole Blood Which Were Shipped At Unacceptable Temperatures | 10 Units Were Distributed | Indiana
1997-03-12 | Biologics | II | b-432-7 | Plasma Apheresis Bowl Model 00625 B With Ultra Sonic Weld Sterile Fluid Path Used With Haemonetics Plasma Collection System | Haemonetics Corporation | Haemonetics Corporation Leetsdale Pennsylvania | February 20 1997 | Plasma Apheresis Bowls Determined To Have An Incomplete Seal In The Ultrasonic Weld Joint Causing The Bowl To Leak During Normal Plasmapheresis Operations | 2 073 099 Bowls Were Distributed | California Florida Georgia Iowa Indiana Massachusetts New York Oklahoma Pennsylvania South Carolina Tennessee
1997-03-12 | Biologics | II | b-434-7 | Platelets | Manufacturer | Blood Systems Inc | January 22 1996 | Blood Product Was Prepared From A Unit Which Was Collected From A Donor Who Had Taken Aspirin Within Three Days Of Donation | 1 Unit Was Distributed | Louisiana
1997-03-12 | Biologics | II | b-435-7 | Platelets | Manufacturer | Central Indiana Regional Blood Center Indianapolis Indiana | January 10 1996 | Blood Product Was Prepared From A Unit Which Was Collected From A Donor Who Had Taken Aspirin Within Three Days Of Donation | 1 Unit Was Distributed | Utah
1997-03-12 | Biologics | II | b-459-7 | Bayrab Brand Rabies Immune Globulin Human | Manufacturer | Bayer Corporation Clayton North Carolina | December 2 1996 | Product Was Labeled As Preservative Free But Contained Low Levels Of Thimerosal | Approximately 6 000 Doses Of Single Dose Vials Were Distributed 5 | Nationwide
1997-03-12 | Biologics | II |  | S Division Rochester Michigan | The Parke Davis Division Of Warner Lambert Company |  | February 12 1997 | Parke Davis Fluogen Showed A Decrease In Potency Of One Component Of The Vaccine After Distribution | Approximately 551 000 Units Were Distributed | Nationwide
1997-03-12 | Biologics | III | b-433-7 | Platelets Pheresis | Manufacturer | American Red Cross Blood Services Baltimore Maryland | July 22 1996 | Blood Product Was Collected From A Donor With An Unacceptable Health History | 1 Unit Was Distributed | District Of Columbia
1997-03-12 | Biologics | III | b-436-7 | Monoclate P Antihemophilic Factor Human Factor Viii C Pasteurized Monoclonal Antibody Purified | Centeon L L C | Centeon L | February 19 1997 | Reconstitution Difficulties In That Vials Failed To Transfer All Of The Diluent Into A Lyophilized Product | 1 908 Vials Were Distributed Firm Estimated That Little If Any Product Remained On Market At Time Of Recall Initiation | Nationwide
1997-03-12 | Devices | II | z-297-7 | Sensi Touch Regional Anesthesia Delivery System Spinal Anesthesia Tray Catalog 8888 828111 22 Ga Plastic Hub Diamond Point Marcaine | Sherwood Davis Geck | Kelsar S | January 31 1997 | The Tray Incorrectly Identifies The Contents As Product Catalog No | 400 Trays Were Distributed Between 8 25 96 And 1 27 97 | Michigan Missouri Connecticut Texas New Jersey Kentucky Minnesota Pennsylvania Massachusetts
1997-03-12 | Devices | II | z-298/299-7 | Imatron Ct Scanner A Component Of The Patient Table Potentiometer Assembly Y Potentiometer A Imatron C 100 Ct Scanner B Imatron C 150 Series Scanners | Manufacturer | Imatron Inc | July 19 1996 | The Table Would Go Beyond Its Limits And Could Not Be Driven Back Due To A Blown Fuse | 81 Units Were Distributed | Nationwide And International
1997-03-12 | Devices | II | z-300-7 | Model Crc Laser Profilometer Used To Determine The Surface Characteristics Of Paper | Manufacturer Fda Approved The Firm's Corrective Action Plan On March 1 | International Paper Company Tuxedo New York | March 1 1997 | Noncompliance With Performance Standards For Laser Products In That The Products Failed To Incorporate Required Safety Interlocks On The Operator Access Door | 2 Units Were Distributed | Nationwide
1997-03-12 | Devices | II | z-303-7 | Eeg Tend Monitor Model Oee 7102 B Intended For Long Term Eeg Monitoring And Analysis | Nihon Kohden America | Nihon Kohden Corporation Shinjuku Ku Tokyo Japan | October 16 1996 | When The 12 V Dc Power Supply Conductor Is Broken And The Main Power Is Turned On A Higher Than Acceptable Current Leakage Is Applied To The Monitoring Electrodes And A Chemical Reaction Is Generated With Eeg Paste Which Can Discolor The Patient's Skin When The Electrodes Are Placed On The Patient Skin | 23 Units Were Distributed | Arizona California Illinois Kansas Louisiana South Dakota Texas Utah
1997-03-12 | Devices | III | z-293-7 | Neurosign Straight Concentric Bipolar Probe With Touch Proof Connectors Catalog 7033 0919 An Accessory To The Neurosign 100 Motor Nerve Monitor | Smith Nephew Ent | The Magstim Company Ltd | October 16 1996 | The Device Is Subject To Possible Extrusion Of Silicone From The End Of The Probe When The Device Was Autoclaved By The Manufacturer | 8 Units Were Distributed And Are Accounted For 8 | Arizona Kentucky New York Texas Washington State Wisconsin
1997-03-12 | Devices | III | z-294-7 | Incstar Vca Igg Clin Elisa Test Kit Product 4590 An In Vitro Diagnostic Test Kit For Detecting Antibodies In Human Serum To Epstein Barr Viral Capsid Antigen Vca | Manufacturer | Incstar Corporation Stillwater Minnesota | February 3 1997 | Some Negative Samples May Yield A Positive Assay Response | 131 Kits Were Distributed | California New Jersey New York International
1997-03-12 | Devices | III | z-302-7 | Cytomegalovirus Igm Elisa Test Kit Product 2325250 Used For The Qualitative Determination Of Igm Antibody To Cytomegalovirus In Human Serum | Manufacturer | Clark Laboratories Inc | January 23 1997 | The Absorbance Values For Controls And Specimens Are Lower Than Expected | 495 Kits Were Distributed | Nationwide And International
1997-03-12 | Devices | III | z-304-7 | C 150 Xp Ct Scanner Installed With Software Version 12 | Manufacturer | Imatron Inc | August 28 1996 | The Device's Software Version 12 | 30 Units Were Distributed | Nationwide And International
1997-03-12 | Devices | III | z-305/307-7 | Software Versions 11 | Manufacturer | Imatron Inc | November 20 1995 | During Reconstruction Of Images The Fast Reconstruction System Software Fsr May Produce An Error Condition Which The Computer May Interpret As A Warning Causing The Computer To Display Erroneous Data For The Image | 37 Units Were Distributed | Nationwide And International
1997-03-12 | Devices | III | z-308-7 | Nellcor Symphony N 3000 Pulse Oximeter Model N 3000 | Nellcor Puritan Bennett | Nellcor Puritan Bennett Inc | October 23 1996 | The Device May Be Potentially Subject To Printed Circuit Board Damage Which Can Result In The Monitor Reporting Sp 02 Values Slightly Outside The Normal Accuracy Range Stated In The Operations Manual | 111 Units Were Distributed | Nationwide And International
1997-03-12 | Devices | III |  | Sherpa Vector And Sherpa Vector X Coronary Guiding Catheters Designed To Provide A Pathway Through Which Therapeutic Devices Are Introduced A Vector 6 F Catalog Numbers Pa 5 Xx B Vector 7 F Catalog Numbers As 7 Xxx C Vector 8 F Catalog Numbers As 8 Xxx D Vector X 6 F Catalog Numbers S 26 Xxx E Vector X 7 F Catalog Numbers S 27 Xxx F Vector X 8 F Catalog Numbers S 28 Xxx |  | Medtronic Interventional Vascular Inc | N A | Resistance May Be Experienced While Passing Interventional Devices Through The Lumen Of The Guiding Catheter Which May Result In The Inability To Move The Interventional Device Within The Lumen Of The Guiding Catheter | Approximately 38 796 Units Were Distributed | Nationwide