## 1998-7-15
RECALL_CLASSIFICATION_DATE | PRODUCT_TYPE | CLASSIFICATION | RECALL_NUMBER | PRODUCT | RECALLING_FIRM | MANUFACTURER | RECALL_INITIALIZATION_DATE | REASON | VOLUME | DISTRIBUTION
1998-7-15 | Drugs | III | d-182-8 | Oxygen Medical Usp Compressed In High Pressure E Cylinders Rx | Manufacturer | Sieps Homecare Inc | May 27 1998 | Subpotent And Good Manufacturing Practice Deviations | 2 Cylinders Were Distributed | Minnesota
1998-7-15 | Drugs | II | d-172-8 | Compressed Oxygen Usp Transfilled In E And M 6 Size Aluminum Cylinders | Manufacturer | Shiv Respiratory Equipment Inc | May 17 1998 | Subpotency | 218 E Size Cylinders 125 M 6 Size Cylinders Were Distributed | New York
1998-7-15 | Devices | II | z-642-8 | Ohmeda Tec 6 Desflurane Vaporizers Note Excludes North American Drager Nad And Dragerwerk Tec 6 Variant Vaporizers | Manufacturer | Datex Ohmeda Formerly Ohmeda Medical Systems Division Madison Wisconsin | June 22 1998 | Gas Leaks | 6 207 Vaporizers Were Distributed | Nationwide And International
1998-7-15 | Devices | II | z-643/644-8 code lot numbers: a) 96 g 398; b) 96 g 348 | A 3 M Modular Shoulder System 14 Mm X 125 Mm Stems Product Number 17464 B 3 M Modular Should System 16 Mm X 125 Mm Stems Product Number 17466 A Metal Stem Designed To Fit Into A Surgically Prepared Cavity In A Humerus Bone During Shoulder Joint Replacement Surgery | 3 M Health Care | 3 M Health Care Rotherham England | June 3 1998 | One Stem With A 16 Mm Diameter Was In A Box Labeled As 14 Mm Diameter | A 42 Stems B 31 Stems Were Distributed | Nationwide And International
1998-7-15 | Devices | II | z-645-8 | Medex Injection Adapter With Male Leur Lock Intended For Use In The Administration Of Small Volume Fluids Using Syringes | Medex | Medex Inc | June 26 1998 | The Sterility Of The Device Cannot Be Assured Due To Defective Packaging | 390 Cases 100 Units Per Case Were Distributed | Arkansas Illinois New Jersey Pennsylvania New York Montana Connecticut California Maryland Missouri North Carolina Tennessee Texas Virginia Iowa Oregon Minnesota
1998-7-15 | Devices | II | z-646/647-8 | Newport Infant Adult Breeze Ventilator Designed For Respiratory Support Of Infants Pediatrics Or Adults A Model E 200 D Domestic Sales Ventilator B Model E 200 E Foreign Sales Ventilators | Newport Medical Instruments | Mer Mar Inc | March 5 1998 | A Bad Ic Chip Causes Inadvertent System Reset Continually During Power Up Or Randomly During Operation Wherein 1 The Ventilators Cease Operation Gas Flow Ceases And The Exhalation Valve Opens For About 4 Seconds Then Restarts At Values That Were Set Before The Reset Except As Follows 2 The Sigh Function Defaults To Off 3 The Nebulizer Defaults To Off 4 The Inspiratory Minute Volume Alarm Defaults To The 0 50 L | 61 Ventilators Were Distributed | Nationwide And International
1998-7-15 | Devices | II | z-649/650-8 | Victoreen Veridose Diodes Designed For Use As Radiation Detection Devices Encased Within A Polystyrene Material A Model 30 472 Positive Polarity Veridose Diodes B Model 30 472 8000 Negative Polarity Veridose Diodes | Manufacturer | Victoreen Inc | May 7 1998 | The Devices Were Manufactured With An Inappropriate Shielding Material Build Up Cap | 23 Diodes Were Distributed | Arizona Kentucky New York Pennsylvania Florida Virginia Maryland Michigan Wisconsin Ohio Tennessee Texas Alabama Georgia Mississippi
1998-7-15 | Devices | II | z-667/668-8 | Tracheal Intubation Fiberscope Instruction Manual A Catalog No | Olympus American | Olympus Optical Company Ltd | June 3 1998 | The Cover Pages For The Instruction Manuals Were For The Correct Intubation Fiberscope But The Rest Of The Instructions Were For The Other One | A 27 Units B 287 Units Were Distributed Were Distributed | Nationwide And International
1998-7-15 | Devices | II | z-669-8 | Model Tru Pulse C 02 Laser System Used In General Surgery | Manufacturer Fda Approved The Firm's Corrective Action Plan On July 1 | Tissue Medical Lasers Inc | July 1 1998 | Noncompliance With Performance Standards For Laser Products In That The Operators Manual Lacked Calibration Procedures | 275 Units Were Distributed | Nationwide And International
1998-7-15 | Devices | II | z-670-8 | I 125 Seeds Iodine 125 Therapeutic For Interstitial Brachytherapy Model 6711 Used As Radioactive Implants For The Treatment Of Prostate Cancer | Manufacturer | Medi Physics Inc | September 23 1997 | The Seeds Were Mislabeled With An Incorrect Apparent Radioactivity Range | 26 Seeds Were Distributed | Georgia And Florida
1998-7-15 | Devices | II | z-671-8 | I 125 Seeds Iodine 125 Therapeutic For Interstitial Brachytherapy Model 6711 Used As Radioactive Implants For The Treatment Of Prostate Cancer | Manufacturer | Medi Physics Inc | January 8 1998 | The Seeds Were Mislabeled With An Incorrect Apparent Radioactivity Range | 161 Seeds Were Distributed | Washington State California New York
1998-7-15 | Devices | II | z-672-8 | Disposable Manual Resuscitator Part Dmr 2 Used To Deliver Supplemental Oxygen | Nellcor Puritan Bennett | Nellcor Mexico Tijuana Mexico | May 22 1998 | The Duckbill Valve Can Become Unseated Allowing Rebreathing And Or Insufficient Inspiratory Air Pressure And 2 The Pressure Relief Valve Ball Can Rust Causing It To Become Inoperative | 67 490 Units Were Distributed | Nationwide And Italy