## Report-1995-01-18
RECALL_CLASSIFICATION_DATE | PRODUCT_TYPE | CLASSIFICATION | RECALL_NUMBER | PRODUCT | RECALLING_FIRM | MANUFACTURER | RECALL_INITIALIZATION_DATE | REASON | VOLUME | DISTRIBUTION
1995-01-18 | Drugs | II | d-044-5 | Artane Trihexyphenidyl Hcl Sequels 5 Mg Sustained Release Capsules In Bottles Of 60 Rx Used In The Treatment Parkinsonism | Manufacturer | Lederle Labs Inc | December 7 1994 | Product Does Not Meet Dissolution Specifications Through Expiration Date | 15 031 Bottles Were Distributed | Nationwide Chile
1995-01-18 | Drugs | II | d-045-5 | Asendin Amoxapine Tablets 100 Mg In Bottles Of 100 And In 10 Unit Dose Strips Totaling 100 Tablets Rx Indicated For The Relief Of Symptoms Of Depression | Manufacturer | Lederle Laboratories Inc | November 14 1994 | Product Does Not Meet Dissolution Specifications Through Expiration Date | 7 288 Bottles And Unit Dose Strips Were Distributed | Nationwide
1995-01-18 | Drugs | III | d-038-5, which appeared in the december 21, 1994 enforcement report listed a lot number as 1479smj064 |  |  |  | N A |  |  | N/A
1995-01-18 | Devices | II | z-169-5 | Positron Or Position These Terms Seem To Be Used Interchangeably Energy Processor Powerboard Part 3100128 000 For The Ecat Scanner 953 | Cti Pet Systems | Siemens Medical Systems Inc | October 15 1993 | A Component Not Permitted By The Manufacturer's Specifications Was Used In The Manufacture Of The Powerboards | 3 Units Firm Estimates None Remain On Market | Georgia Sweden Germany
1995-01-18 | Devices | II | z-233/234-5 | Transducer Protectors 0 | Millipore Corporation | Millipore Cidra Inc | January 27 1994 | Product May Exhibit Air Leakage At The Female Luer Lock Connection To The Male Luer Of The Hemodialysis Machine Under Normal Operating Conditions | 3 5 Million Were Distributed | Nationwide
1995-01-18 | Devices | II | z-235-5 | Diapure Dissolution System Used In Hemodialysis Treatment | Manufacturer | Fresenius Usa Inc | June 20 1994 | Product Does Not Have An Approved Application For Pre Market Approval Or An Investigational Device Exemption In Effect | 66 Units | California Louisiana Maryland New Jersey Oklahoma
1995-01-18 | Devices | II | z-236/237-5 | Siemens Linear Accelerators A Digital Mevatron Linear Accelerators With Therapy Data Management System Tdms B Conventional Mevatron Linear Accelerator With Tvs System | Manufacturer | Siemens Medical Systems Inc | September 15 1994 | The Tdms Software And Hardware Versions 3 | 110 Units Were Distributed | Nationwide And International
1995-01-18 | Devices | II | z-264/266-5 | Aequitron Ventilators A Model Lp 6 V Volume Ventilators B Lp 10 Volume Ventilators D Drager Model Ev 801 Electronic Ventilators Used To Provide Continuous Respiratory Support For Patients With Respiratory Insufficiencies Recall Z 264 266 5 | Manufacturer | Aequitron Medical Inc | December 9 1994 | Device Malfunctions Could Result Because Of A Failure To Apply Adhesive To Bearings In The Connecting Arm Assemblies Of The Ventilators | A 85 Units B 248 Units C 102 Units Were Distributed | Nationwide Canada Japan Germany
1995-01-18 | Devices | II | z-287-5 | Silicone Suction Reservoir 100 Cc Wound Drainage System | Stryker Instruments | Sil Med Corporation Taunton Massachusetts | May 23 1994 | Seals On The Pouches Containing The Devices Were Broken And The Sterility Of The Devices Had Been Compromised | 30 Boxes 10 Units Per Box Were Distributed | Nationwide
1995-01-18 | Devices | II | z-288/296-5 | 5 French Pigtail And Selective Angiographic Catheters | Manufacturer | Scimed Life Systems Inc | November 8 1994 | The Distal Tip Of The Devices Can Separate During Use | 1 944 Units Were Distributed | Nationwide And International
1995-01-18 | Devices | II | z-298-5 | Pulsemaster Dental Surgery Laser System | Manufacturer Firm Initiated Field Correction Ongoing Distribution Nationwide | Incisive Technologies Inc | N A | The Device Was Found To Have A Defective Design In That Certain Heat Exchangers That Had Been Distributed Were Non Passivated Exchangers That Corrode Easily Resulting In Decreased Laser Output Power | 39 Units Were Distributed | Nationwide Germany Japan
1995-01-18 | Devices | III | z-242/245-5 | Olympus Disposable Snares Used For Polypectomy Procedures A Model Sd 14 L B Model Sd 14 U C Model Sd 15 L D Model Sd 15 U | Olympus America | Shiraku Olympus Company Ltd | November 4 1994 | The Snare Wire Could Separate From The Interconnection To The Handle During The Removal Of Polyps | 3 396 Units Were Distributed | Nationwide
1995-01-18 | Devices | III | z-267-5 | Stratus Hlh Fluorometeric Enzyme Immunoassay Which Is An Automated Procedure For The Quantitative Determination Of Hlh In Serum And Plasma Samples | Baxter Diagnostic | Baxter Diagnostic Of Pr Inc | February 19 1994 | An Incorrect Checksum Value Was Calculated By Using The Tab Expiration Date Instead Of The Kit Expiration Date | 6 Kits Were Distributed In January 1994 | Colorado Florida Iowa Kansas Texas
1995-01-18 | Devices | III | z-268-5 | Stratus Hfsh Fluorometric Enzyme Immunoassay An Automated Procedure For The Quantitative Determination Of Hfsh Levels In Serum And Plasma Samples | Baxter Diagnostic | Baxter Diagnostic Of Pr Inc | November 19 1993 | An Incorrect Checksum Value Was Provided For The Analyzer Employed By This Test | 37 Kits Were Distributed From September 1993 To December 1993 | Nationwide And Venezuela
1995-01-18 | Devices | III | z-269-5 | Polypectomy Snare 7 French In Small Medium And Large Sizes | Manufacturer | Wiltek Medical Inc | August 1 1994 | The Snare Head May Become Detached From The Cable | 3 541 Snares Were Distributed | California Connecticut Georgia Michigan Missouri New York Pennsylvania Virginia
1995-01-18 | Devices | III | z-284/286-5 | Stockert Heater Cooler Units Used To Warm Cool Patient's Blood During Open Heart Surgery A Model 16 02 20 Heater Cooler Base Unit B Model 16 12 01 Heater Cooler Ii 115 V Base Unit Plus Controller C Model 16 21 01 Heater Cooler With Mast Base Unit Plus Controller | Sorin Biomedical | Stockert Instrumente Gmbh Munich Germany | October 24 1994 | The Devices Fail To Heat Or Cool During Use | 25 Units Were Distributed | California Indiana Ohio Pennsylvania Texas Washington State
1995-01-18 | Devices | III | z-297-5 | Stratus Ck Mb Fluorometric Enzyme Immunoassay An Automated Procedure For The Quantitative Determination Of Ck Mb Levels In Serum And Plasma Samples | Baxter Diagnostic | Baxter Diagnostic Of Pr Inc | December 19 1993 | Products Are Being Recalled Due To Kit Cross Configurations With One Or More Components Being Assigned The Same Lot Number | 802 Kits Were Distributed From December 1993 To January 1994 | Nationwide And Canada