## Report-1995-02-01
RECALL_CLASSIFICATION_DATE | PRODUCT_TYPE | CLASSIFICATION | RECALL_NUMBER | PRODUCT | RECALLING_FIRM | MANUFACTURER | RECALL_INITIALIZATION_DATE | REASON | VOLUME | DISTRIBUTION
1995-02-01 | Drugs | III | d-053-5 | Pseudoephedrine 60 Mg Guaifenesin Maxifed G 500 Mg Tablets Rx Decongestant In Bottles Of 100 | American Pharmaceuticals | Mikart Inc | November 2 1994 | A Bottle Was Found To Contain A Foreign Tablet Pyrazinamide 500 Mg | 2 488 Bottles Were Distributed | Georgia Alabama Tennessee Louisiana Pennsylvania North Carolina Mississippi Virginia South Carolina Florida Illinois
1995-02-01 | Drugs | III | d-054-5 | Isocom Capsules Rx Used For The Treatment For Relief Of Vascular And Tension Headaches In Bottles Of 50 100 And 250 | Manufacturer | Nutripharm Laboratories Inc | August 31 1994 | Insufficient Stability Data To Support Expiration Date | Firm Estimates Few Bottles Remain On Market | Nationwide
1995-02-01 | Drugs | III | d-055-5 | Century Pharmaceuticals Brand Potassium Chloride 20 Solution In Pint Bottles Rx Source Of Potassium | Manufacturer | Century Pharmaceuticals Inc | January 14 1994 | Presence Of Mold In Product | 932 Pints Were Distributed | Maryland Mississippi Arizona South Dakota Nebraska Minnesota Wyoming New Mexico
1995-02-01 | Biologics | II | b-152-5 | All Products Distributed From Units Collected Between March 23 1994 And March 27 1994 | Manufacturer | The American National Red Cross Charlotte North Carolina | March 29 1994 | Blood Products Incorrectly Tested For Antibodies To Human Immunodeficiency Virus Types 1 And 2 Anti Hiv 1 2 Were Distributed | Undetermined | Undetermined
1995-02-01 | Biologics | II | b-153-5 | Red Blood Cells Leukocytes Removed | Manufacturer | United Blood Services Blood Systems Inc | February 2 1994 | Blood Products In Which Filtration And The Addition Of Additive Solution Occurred More Than Eight Hours After Collection Was Distributed | 2 Units | Texas
1995-02-01 | Devices | II | z-145/146-5 | Motorized Three Wheeled Vehicle A Battery Powered Device Intended For Medical Purposes That Is Used For Outside Transportation By Disabled Persons A Legend Model Sc 300 B Sunrunner Model Sc 400 | Manufacturer | Pride Health Care Inc | April 5 1993 | The Base Of The Seat Component Of The Device Was Not Welded Properly Which Could Cause The Seat Component To Break Off The Device | Firm Estimates There Are No Uncorrected Scooters In Commerce | Arkansas California Colorado Florida Connecticut Illinois Iowa Minnesota New York North Carolina Nebraska Ohio Pennsylvania Texas Wisconsin Louisiana Kentucky Washington State Oregon Canada
1995-02-01 | Devices | II | z-314-5 | Stryker Thermaltek 300 Portable Patient Warmer | Manufacturer | Stryker Patient Handling Kalamazoo Michigan | September 27 1994 | When The Warming Device Is Set On High The Device Does Not Automatically Cool Down After A Period Of Time Resulting In Patient Burns | 293 Units Were Distributed | Nationwide And International
1995-02-01 | Devices | II | z-331/332-5 | Vector Mti Ambulatory Infusion Pumps A Model 102040 B Model 102045 | Manufacturer | Infusion Technology Inc | November 7 1994 | When The Infusion Pump Is Set To The Total Parenteral Nutrition Mode The Pump Allows The Potential For An Incorrect Infusion Rate But Will Not Over Infuse | 88 Units Were Distributed | Nationwide
1995-02-01 | Devices | II | z-147/148-5 | C Taper Head Modular Femoral Bearing Head A 26 Mm Catalog 06 2605 B 28 Mm Catalog 06 2805 | Manufacturer | Osteonics Inc | March 1 1994 | Packaging Mix Up The 26 Mm Labeled Package May Contain The 28 Mm Unit Or Vice Versa | 80 Units Were Distributed | Nationwide And Japan
1995-02-01 | Devices | II | z-300-5 | Aldetect Formaldehyde Test Reagent For In Vitro Diagnostic Use Product Is Sold As 0 | Dewallace Technical Sales | Pochemco Inc | June 13 1994 | Product Exhibits Discoloration | 44 Cases Of Lot 5240 46 Case Of Lot 5241 45 Cases Of 5242 Were Distributed Firm Estimates None Of These Lots Remain On The Market 18 Cases Of Lot 5243 41 Cases Of Lot 5245 And 44 Cases Of Lot 5247 Were Distributed | Nationwide
1995-02-01 | Devices | II | z-305-5 | Abbott Vision Glycated Hb Calibrator Kit Used For Calibrating The Abbott Vision Glycated Hb Test On The Abbott Vision System To Generate A Calibration Curve Used To Determine The Percent Hemoglobin A 1 C In A Specimen | Manufacturer | Abbott Laboratories Diagnostics Division Abbott Park Illinois | May 9 1994 | This Lot Of The Calibrator Kit Had A Higher Value Than Assigned | 2 567 Sets Were Distributed | Nationwide Canada Brazil Italy Japan Singapore Taiwan Australia Germany
1995-02-01 | Devices | II | z-306-5 | Blood Agar Tsa With Sheep Blood An In Vitro Diagnostic Plated Media Used As A General Purpose Medium For The Growth Of A Wide Variety Of Microorganisms | Manufacturer | Remel Limited Partnership Lenexa Kansas | June 13 1994 | The Plates Had Bacterial Contamination | Firm Estimated That None Of Lot 2042 Remained On The Market However Only 1 100 Of The 1 600 Plates From Segment 123 Were Confirmed To Be Recovered At Time Of Recall Initiation | Florida Nebraska Missouri Texas Illinois Wisconsin Michigan Minnesota Alabama South Carolina Georgia North Carolina Indiana Ohio Pennsylvania Oklahoma Iowa
1995-02-01 | Devices | II | z-307-5 code catalog #06-2412 | Phenol Red Broth With 1 Sorbitol An In Vitro Diagnostic Recommended For Use In Qualitative Procedures For The Determination Of Fermentation Reactions Of Micoroganisms | Manufacturer | Remel Limited Partnership Lenexa Kansas | June 9 1994 | The Box Labeled As Phenol Red Broth W 1 Sorbitol Contained The Product Cta Dextrose | 640 Tubes Were Distributed Firm Estimates None Remains On The Market | Wisconsin California Texas Georgia Ohio Connecticut Missouri New York Virginia
1995-02-01 | Devices | II | z-308-5 | Bordetella Pertussis Atcc 9340 Bacti Disk An In Vitro Diagnostic Product Recommended For Use In Laboratory Quality Assurance Procedures | Manufacturer | Remel Limited Partnership Lenexa Kansas | April 1 1994 | The Bacti Disk Labeled As Bordetella Pertussis Atc 9340 Actually Contained The Organism Bordetella Parapertussis | 17 Vials Were Distributed Firm Estimates None Remains On The Market | Iowa Florida California Ohio Indiana Michigan New York Illinois Arkansas Maryland Kansas
1995-02-01 | Devices | II | z-312/313-5 | Penile Implants A Ultrex Plus 18 Cm Infrapubic Implant Part 72401467 B Ultrex Plus 18 Cm Penile Scrotal Implant Part 72401466 | Manufacturer | American Medical Systems Inc | October 5 1994 | The Tubing Between The Cylinders And The Pump Is Too Short In Lot Aj 767 And Too Long In Lot Aj 766 | Firm Reported That As Of 11 17 94 8 Of The 10 Penile Scrotal Implants The Other Two Had Been Implanted And 9 Of 9 Infrapubic Implants Had Been Returned To Ams | Texas North Carolina California
1995-02-01 | Devices | II | z-317/318-5 | Ultracast Non Convertible Buccal Tubes Used With Dental Braces To Translate Torque And Rotational Forces From The Band To The Teeth A Maxillary Double Tube With Mesial Hook Catalog Numbers 003 291 Cr 003 291 Cl 003 293 Cr 003 293 Cl B Mandibular Double Tube With Ball Hook Catalog Numbers 004 213 Cr 004 213 Cl 004 214 Cr 004 214 Cl | Manufacturer | American Orthodontics Corporation Sheboygan Wisconsin | March 19 1993 | Tubes Are Made Of Stainless Steel With A High Carbon Content Which Causes Them To Be Brittle | Approximately 5 400 Tubes Were Distributed Firm Estimates None Remains On The Market | Nationwide
1995-02-01 | Devices | II | z-327-5 | Abbott Vision Crp C Reactive Protein Reagent Used With The Abbott Vision System For The Quantitative Determination Of C Reactive Protein In Anticoagulated Whole Blood Plasma Or Serum | Manufacturer | Abbott Laboratories Abbott Park Illinois | December 19 1994 | This Lot Of The Abbott Vision C Reactive Protein Reagents The Calibrator A Has A Higher Value Than Assigned | 1 572 10 Packs Were Distributed | Alabama Arkansas California Florida Georgia Illinois Indiana Iowa Michigan Montana New York North Carolina Oklahoma Pennsylvania Tennessee Texas
1995-02-01 | Devices | II |  | Digiflator Digital Inflation Devices A Mansfield Product Number 1520 20 Cc Angioplasty B Microvasive Product Number 5050 60 Cc Gastrointestinal C Microvasive Product Number 5055 Alimentary Tract Use D Microvasive Product Number 705 108 20 Cc Urological E Microvasive Product Number 705 109 20 Cc Urological 12 Box F Medi Tech Product Number 15 150 20 Cc Percutaneous Angioplasty G Medi Tech Product Number 15 151 20 Cc Percutaneous Angioplasty |  | South East Instruments Medical Inc | N A | Loss Of Electronic Display And Accuracy Of Pressure Display | Firm Estimated That 2 386 Units Remain On Market | Nationwide And International