## 1994-02-23
RECALL_CLASSIFICATION_DATE | PRODUCT_TYPE | CLASSIFICATION | RECALL_NUMBER | PRODUCT | RECALLING_FIRM | MANUFACTURER | RECALL_INITIALIZATION_DATE | REASON | VOLUME | DISTRIBUTION
1994-02-23 | Drugs | II | d-130/131-4 | Sus-phrine (epinephrine) Suspension 1:200 For Subcutaneous Injection Packaged In Multidose 5-ml Vials And 0.3-ml Ampuls. Recall #d-130/131-4. | Forest Pharmaceuticals Inc. St. Louis Missouri | Steris Laboratories Inc. Phoenix Arizona. | December 4, 1993 | Potential for presence of mold in product containers resulting from storage conditions at recalling firm | Approximately 45 147 0.3-ml Ampuls And 39 842 5-ml Vials Were Distributed. | Nationwide.
1994-02-23 | Drugs | II | d-132-4 | Ergoloid Mesylates Tablets (oral) 1 Mg Rx Used For The Symptomatic Relief Of Idiopathic Decline In Mental Capacity In Bottles Of 60 And 100. Recall #d-132-4. | Allscrips Phrmaceuticals Inc. Vernon Hills Illinois (repacker) | Danbury Pharmacal Inc. Carmel New York. | January 19, 1994 | Product does not meet usp specifications for content uniformity | 31 Bottles Of 100 And 22 Bottles Of 60 Tablets Were Distributed; Firm Estimates 25% Of Product Remains On The Market. -2- | Ohio Florida.
1994-02-23 | Drugs | III | d-129-4 | Loestrin Fe (norethindrone Acetate Combined With Ethinyl Estradiol Tablets Usp And Ferrous Fumarate Tablets Usp) 1.5/30. Recall #d-129-4. | Parke-davis Division Of Warner Lambert Morris Plains New Jersey | Warner Lambert Fajardo Puerto Rico. | December 10, 1993 | Product does not meet process validation specifications | Approximately 39 002 Units Were Distributed. | Nationwide.
1994-02-23 | Biologics | II | b-139-4 | Saliva Specimen Collection Kit (orasure) For The Collection Of Human Mucosal Transdudata (saliva Component) To Be Used In A Test For The Detection Of The Antibody To The Human Immunodeficiency Virus Type 1 (anti-hiv-1). Recall #b-139-4. | Manufacturer | Home Office Reference Laboratory Lenexa Kansas. | July 12, 1991 | An unapproved device. | 312 197 Kits. | Nationwide.
1994-02-23 | Biologics | II | b-157/158-4 | (a) Red Blood Cells; (b) Platelets. Recall #b-157/158-4. | Manufacturer | W.e. And Lela I. Stewart Regional Blood Center Tyler Texas. | May 10, 1993 | Blood products tested repeatedly reactive for the antibody to hepatitis b core antigen (anti-hbc) | 1 Unit Of Each Component. | Texas.
1994-02-23 | Biologics | II | b-159-4 | Red Blood Cells. Recall #b-159-4. | Manufacturer | East Texas Blood Center Inc. Nacogdoches Texas. -3- | May 18, 1993 | Blood product tested repeatedly reactive for the antibody to human t-lymphotropic virus type i (anti-htlv-i) | 1 Unit. | Texas.
1994-02-23 | Biologics | III | b-138-4 | Red Blood Cells. Recall #b-138-4. | Manufacturer On August 26 1991. Firm-initiated Recall Complete. | United Blood Services Blood Systems Inc. Scottsdale Arizona. | August 26, 1991 | Autologous blood product tested positive for syphilis but was not labeled with a biohazard label | 1 Unit. | Arizona.
1994-02-23 | Devices | II | z-736-3 | Model Lp10 Ventilator. Recall #z-736-3. | Manufacturer | Aequitron Medical Inc. Minneapolis Minnesota. | August 25, 1993 | The pressure limit control valve may stick in some ventilators, which may allow patient pressure to increase above the control setting | 915 Units. | Nationwide And International.
1994-02-23 | Devices | II | z-225/226-4 | Bird Adult/pediatric Volume Ventilators Product #15020 Electrically Powered Microprocessor Controlled Volume Cycled Pressure Limited Critical Care Ventilator: (a) Model 8400st; (b) Model 8400sti. Recall #z-225/226-4. | Manufacturer | Bird Products Corporation Palm Springs California. | September 15, 1993 | The products' exhalation valve is malfunctioning due to the premature failure of the component flex circuit which provides electrical energy to the voice coil | 3 112 Units. | Nationwide And International.
1994-02-23 | Devices | II | z-253/254-4 | Baxter Flo-gard Volumetric Infusion Pumps Used To Deliver A Wide Variety Of Fluids: (a) Flo-gard 6201 Volumetric Infusion Pump Model 2m8063; (b) Flo-gard 6301 Dual Channel Volumetric Infusion Pump Model No. 2m8064 (two Channels Allow The Delivery Of Two Fluids Simultaneously). Recall #z-253/254-4. | Baxter Healthcare Corporation Round Lake Illinois | Sharp Corporation Nara Japan. | December 9, 1993 | Reported failure of the devices to detect air in the iv tubing | (a) 16 300 Units; (b) 4 400 Units Were Distributed. | Nationwide And International.
1994-02-23 | Devices | II | z-258-4 | Laser Illuminator Model-alp-1 Used To Illuminate Object In Low Light Level With Infrared Radiation For Viewing Through A Night Vision Scope. Recall #z-258-4. | Manufacturer. Fda Approved The Firm's Corrective Action Plan January 27 1994. Firm-initiated Recall Ongoing. | Life Plus Management Inc. New York New York. | January 27, 1994 | Product lacked a certification label; lacked an identification label; failed to incorporate a remote interlock connector; failed to incorporate a key control; failed to incorporate an emission indicator; failed to incorporate a beam attenuator; failed to have a class iiib warning logotype affixed to the product; failed to have an aperture label affixed to the product; user information failed to include the required caution statement and a reproduction of all required warning labels | 15 Units. | International.
1994-02-23 | Devices | III | z-770-3 | Matrx Centurion Mixer Rx Medical Device Used In Dental Analgesia Which Electronically Mixes Nitrous Oxide And Oxygen. Recall #z-770-3. | Manufacturer | Matrx Medical Inc. Orchard Park New York. | September 10, 1993 | The devices lack a prescription device label and the user manual contains a typographical error which resulted in column headings being switched on page 8 of the operations manual dealing with performance and accuracy specifications | 342 Units Were Distributed. | Nationwide Japan Australia Finland Colombia Sweden Italy Brazil Israel.
1994-02-23 | Devices | III | z-208-4 | Hypoguard Supreme Blood Glucose Test Strips 50 Test Strips Per Bottle Used With Supreme Blood Glucose Meters By Diabetic Patients To Monitor Their Blood Glucose Levels. Recall #z-208-4. | Manufacturer | Chronimed Inc. Minnetonka Minnesota. | November 15, 1993 | Incorrect meter codes on the test strips resulted in glucose readings which were as much as 17 | 314 Bottles Of 50 Test Strips Per Bottle Were Distributed. | Nationwide.
1994-02-23 | Devices | III | z-223-4 | Apex Hole Eliminator Product #1246-01 A Metal Plug That Is Screwed Into A Threaded Hole In The Bottom Of Acetabular Cup Shells. Recall #z-223-4. | Depuy Division Boehringer Mannheim Corporation Warsaw Indiana | Thortex Inc. Clackamas Oregon. | October 27, 1992 | The minor dimension on the plug, which is screwed into a threaded hole in the bottom of acetabular cup shells, was oversized and would not mate properly | 144 Pieces Were Distributed. | Nationwide And International.