## 1993-11-10
RECALL_CLASSIFICATION_DATE | PRODUCT_TYPE | CLASSIFICATION | RECALL_NUMBER | PRODUCT | RECALLING_FIRM | MANUFACTURER | RECALL_INITIALIZATION_DATE | REASON | VOLUME | DISTRIBUTION
1993-11-10 | Drugs | II | d-018-4 | Atenolol With Chlorthalidone Tablets 50 Mg/25 Mg Rx Used As A Combination Antihypertensive/diuretic Packaged In Bottles Of 100. Recall #d-018-4. | Manufacturer | Mutual Pharmaceutical Company Inc. Philadelphia Pennsylvania. -2- | September 15, 1993 | Product does not meet disintegration specifications | Lot #25322 -- 1 662 Bottles (rugby) 1 675 Bottles (mutual); Lot #25323 -- 352 Bottles (mutual) 144 Bottles (url) 4 320 Bottles (geneva); Lot #25324 -- 288 Bottles (url) 49 Bottles (goldline) 809 Bottles (mutual). | Nationwide.
1993-11-10 | Drugs | II | d-019-4 | Compressed Medical Oxygen Usp In Size E Cylinders. Recall #d-019-4. | Manufacturer | County Health Services Inc. Doing Business As American Medical Equipment (ame) Georgetown Texas. | September 22, 1993 | Current good manufacturer practice deficiencies | Firm Estimates None Remains On The Market. | Texas.
1993-11-10 | Drugs | II | product lacrisert, a sterile rx 5 mg hydroxypropyl cellulose ophthalmic insert and applicator, indicated in patients with moderate to severe dry eye syndromes. recall #d-020- 4. | Lacrisert A Sterile Rx 5 Mg Hydroxypropyl Cellulose Ophthalmic Insert And Applicator Indicated In Patients With Moderate To Severe Dry Eye Syndromes. Recall #d-020- 4. | Manufacturer | Merck Manufacturing Division Division Of Merck & Company Inc. West Point Pennsylvania. | November 4, 1993 | Lack of assurance of sterility | 4 077 Units Were Distributed; Firm Estimates That Very Little If Any Of The Product Remains On The Market. | Nationwide And International.
1993-11-10 | Drugs | II | d-021/022-4 | Cortisone Acetate Tablets Usp 10 Mg And 25 Mg In Bottles Of 100 Used As Replacement Therapy In Adrenocortical Deficiency States. Recall #d-021/022-4. | Manufacturer | The Upjohn Company Kalamazoo Michigan. | December 4, 1992 | Product does not meet dissolution specifications through expiration date | 29 790 Bottles Of Lot Numbers 982df 733hs 561fd And 616hj And 14 539 Bottles Of Lot Numbers 535dw Were Distributed. | Nationwide And International.
1993-11-10 | Biologics | III | b-032-4 | Olympus Pk 7100 Data Diskettes Used To Store Data From The Hard Drive Of The Olympus Pk 7100 Automated Pretransfusion Blood Testing System. Recall #b-032-4. | Olympus America Inc. Lake Success New York | Formats Unlimited East Farmingdale New York. | June 2, 1993 | Diskettes were defective. | 46 Diskettes. | Nationwide.
1993-11-10 | Biologics | III | b-034/036-4 | (a) Red Blood Cells; (b) Platelets; (c) Recovered Plasma. Recall #b-034/036-4. | Manufacturer | American Red Cross Blood Services Dedham Massachusetts. | September 24, 1992 | Blood products, which tested negative for the antibody to human immunodeficiency virus type 1 (anti-hiv-1), but were collected from a donor who previously tested repeatedly reactive for anti-hiv-1, were distributed | 1 Unit Of Each Component. | Massachusetts.
1993-11-10 | Devices | II | z-046/047-4 | Software Programs: (a) Ups Version 10.22 Dos; (b) Unix Systems Used For The Re-entering Of Catheter Description Points In The Nucletron Planning System. Recall #z-046/047-4. | Nucletron Corporation Columbia Maryland | Nucletron B.v. Waardegelder 1 The Netherlands. -4- | October 15, 1992 | In a situation with implants of two or more cathetrs, an error situation can occur when the catheters are described with catheter description points | 231 Units. | Nationwide Brazil Canada West Indies.
1993-11-10 | Devices | II | z-064/073-4 | Adams Compact Ii Centrifuge Distributed As A Primary Care Diagnostic Product And Through Several Private Label Distributors And Clinical Laboratories Which Include: (a) Catalog #0225 - Becton Dickinson (b) Catalog #0226 - Allied Clinical Laboratories (c) Catalog #0230 - Methpath (d) Catalog #0231 - Roche (e) Catalog #0232 - Damon Clinical Laboratories (f) Catalog #0234 - Mawd Medical Laboratories (g) Catalog #0236 - Eastside Medical Laboratories (h) Catalog #0237 - Laboratory Of Clinical Medicine (i) Catalog #0238 - Metwest Clinical Laboratories (j) Catalog #0239 - Clinical Diagnostic Services. Recall #z-064/073-4. | Manufacturer | Becton Dickinson Sparks Maryland. | March 24, 1993 | Under certain operating conditions, the rotating tube shields may contact the wiring within the housing resulting in abraded insulation or exposed wires | Approximately 13 177 Units Were Distributed; Firm Estimates 2 215 Units Remain On The Market. | Nationwide.
1993-11-10 | Devices | III | z-018/019-4 | Dermatology Laser Systems: (a) Candela Sptl-1a Dermatology Laser System (p/n 9904-10-0300/0302); Candela Sptl-1ag (germany) Dermatology Laser System (p/n 9904-10-0302); (b) Candela Sptl Iec Dermatology Laser System (p/n 9904-0010-0250). Recall #z-018/019-4. | Manufacturer | Candela Laser Corporation Wayland Massachusetts. | March 17, 1993 | The design allowed for possible internal fires | (a) 64 Units; (b) 54 Units. | Nationwide And International.
1993-11-10 | Devices | III | z-061-4 | Techmate 1000 An Automated Immunostaining System For Tissue Samples Placed On Microscope Slides. Recall #z-061-4. | Manufacturer | Biotek Instruments Inc. Santa Barbara California. | June 24, 1993 | The device's mechanical arm will stop after the failure of an electronic component which senses the position of the mechanical arm | 40 Units. | California Colorado Connecticut Illinois Massachusetts Maryland Michigan Minnesota North Carolina New Jersey New York Ohio Oregon Pennsylvania Texas Washington Vermont District Of Columbia.
1993-11-10 | Devices | III | z-062-4 | Bioscan Reagent Strips Dip And Read Test For Measuring Glucose Ketone Bilirubin Protein Ph And Blood In Urine In Bottles Each Holding 100 Test Strips. Recall #z-062-4. | Dia-screen Corporation Minneapolis Minnesota | Yeong Dong Pharmaceutical Corporation Seoul Korea. | September 20, 1993 | The test strips are not detecting blood urine in the specified time | 120 Bottles. | Massachusetts Michigan Ohio.
1993-11-10 | Devices | III | z-063-4 | Ortho Brand Round Bottom Mrc-5 Human Embryonic Lung (diploid) Cell Culture Tubes Used For The Isolation Of Infectious Agents And As An Indicator System For The Presence Of Toxins. Recall #z-063-4. | Manufacturer | Ortho Diagnostic Systems Inc. Carpinteria California. | June 28, 1993 | Tubes are contaminated with the organism xanthonas (pseudomonas) maltophilia which causes the device not to function as designed | 1 995 Tubes Were Distributed; Firm Estimates None Remains On The Market. | California Connecticut Illinois Maine Missouri New Jersey New York Ohio Oklahoma Oregon.
1993-11-10 | Devices | III | n-063/064-3 | Acufex Orthopedic Screwdrivers: (a) Part #013108 7 Mm Cannulated Driver Large Handle; (b) Part #013109 9 Mm Cannulated Driver Large Handle; (c) Part #013426 7 Mm Cannulated Driver Small Handle; (d) Part #013427 9 Mm Cannulated Driver Small Handle. Safety Alert #n-063/064-3 & N-087/088-3. |  | Acufex Microsurgical Inc. Mansfield Massachusetts. Alerted By Manufacturer By Letter Dated October 15 1992. Distribution Nationwide And International. Quantity (a) 1 235 Units; (b) 793 Units; (c) 734 Units; (d) 454. Reason There Is A Potential For Tip Breakage | NOT-FOUND | There is a potential for tip breakage | (a) 1 235 Units; (b) 793 Units; (c) 734 Units; (d) 454. | Nationwide And International.
1993-11-10 | Devices | III | product pickled gouramy fish (93-683-331). charges adulterated - the article consists in part of a filthy substance by reason of the presence therein of insects, and rodent, human, and cat/dog hair | Pickled Gouramy Fish (93-683-331). Charges Adulterated - The Article Consists In Part Of A Filthy Substance By Reason Of The Presence Therein Of Insects And Rodent Human And Cat/dog Hair |  |  | NOT-FOUND | Of the presence therein of insects, and rodent, human, and cat/dog hair