## 1993-01-27
RECALL_CLASSIFICATION_DATE | PRODUCT_TYPE | CLASSIFICATION | RECALL_NUMBER | PRODUCT | RECALLING_FIRM | MANUFACTURER | RECALL_INITIALIZATION_DATE | REASON | VOLUME | DISTRIBUTION
1993-01-27 | Drugs | II | d-103-3 | Propylthiouracil Tablets 50 Mg A Rx Drug Used For Treatment Of Hyperthyroidism In Bottles Of 100 And 1 000. Recall #d-103-3. | Barr Laboratories Inc. Pomona New York | Barr Laboratories Inc. Northvale New Jersey. | January 6, 1993 | Product does not meet dissolution specifications | 19 530 Bottles Of Lot 0i089ak And 16 347 Bottles Of Lot 1c089ev Were Distributed. | Nationwide.
1993-01-27 | Drugs | III | d-102-3 | Seba-clear Therapeutic Shampoo In 1 Fluid Ounce (professional Sample) And 4 Fluid Ounce Bottles Non-rx Product Labeled As An Antiseborrheic. Recall #d-102-3. | Manufacturer | Heran Pharmaceutical Inc. San Antonio Texas. | January 8, 1993 | Product is subpotent for salicylic acid ingredient and also not in compliance in formulation for dandruff and psoriasis products | 1 116 4-fluid Ounce Bottles And An Unknown Number Of Sample Size Bottles Were Distributed. | Unknown.
1993-01-27 | Biologics | III | b-128-3 | Source Plasma. Recall #b-128-3. | Premier Bioresources Inc. Fort Worth Texas | Premier Bioresources Inc. Corpus Christi Texas. | October 7, 1992 | Source plasma, exposed to unacceptable temperatures, was distributed | 81 Units. | Indiana.
1993-01-27 | Devices | II | z-001/004-3 | Drills Used For Integral Biointegrated Dental Implant Systems: (a) 10 Mm Intermediate Drill Catalog #0203; (b) 8 Mm Final Spade Drill Catalog #0207; (c) Surgical Instrumentation Kit Catalog #0350; (d) Surgical Autoclave Kit Catalog #0326. Recall #z-001/004-3. | Manufacturer | Calcitek Inc. Carlsbad California. | November 17, 1992 | Drills packaged and labeled as 8 mm final spade drills were actually 10 mm intermediate drills and vice versa | (a) 50 Units; (b) 14 Units; (c) 1 Kit; (8) 1 Kits. All Product Involved Was Returned And Reworked. | Nationwide Uk.
1993-01-27 | Devices | II | z-164-3 | Luxar Model Lx-15 And Lx-20 C02 Surgical Lasers. Recall #z-164-3. | Manufacturer. Fda Approved The Firm's Corrective Action Plan December 8 1992. Firm-initiated Recall Ongoing. | Luxar Corporation Bothell Washington. | December 8, 1992 | Noncompliance with performance standards for laser products in that the mechanical mounting of the laser beam shutter could become loose possibly resulting in failure to block emission from the device | 639 Units Were Distributed. | Nationwide. -4-
1993-01-27 | Devices | II | z-167-3 | Ac Adapter/charger Date | Manufacturer To Sole Direct Account Laerdal Medical Corporation On August 14 1992. Laerdal Sent 2 "important Notifications" Dated August 12 1992 And August 17 1992 To Distributors. Firm-initiated Recall Ongoing. | Devilbiss Health Care Inc. Somerset Pennsylvania. | August 14, 1992 | The adapter/charger may fail to adequately charge the internal battery, resulting in the battery failing to operate the device as long as stated in the labeling | 497 Units. | Nationwide.
1993-01-27 | Devices | II | z-168-3 | Adapter/charger With Date | Manufacturer To Distributors Via A Service Bulletin In October 1992. Firm-initiated Recall Ongoing. | Devilbiss Health Care Inc. Somerset Pennsylvania. | October 19, 1992 | The adapter/charger may fail to adequately charge the internal battery, resulting in the battery failing to operate the device as long as stated in the labeling | Undetermined. | Nationwide.
1993-01-27 | Devices | II | z-204/205-3 | Vha+plus Brand Non-sterile Latex Examination Gloves: (a) Small/medium Product #v7266; (b) Medium Product #v7857. Recall #z-204/205-3. | Ansell Inc. Dothan Alabama | Ansell Inc. Byagama Sri Lanka. | December 12, 1992 | Gloves failed to comply with requirements for patient examination gloves due to leakage defects | 140 Cases (20 Boxes Of 100 Gloves Per Box) Were Distributed. | Nationwide.
1993-01-27 | Devices | II | z-220/230-3 | Baxter Custom Sterile Surgical Packs Which Contain Argyle 18 French Ferguson Left Vent Catheters: (a) Cv Accessory Pack Catalog Numbers Scv41bsamw Scv41bsamx Scv41bsamy Scv41bsamz (b) Cv Pack Catalog Nos. Scv46ohhcc Scv46ohhcb Scv46ohhca Scv46ohshh Scv46ohshi Scv46ohshg Scv46ohshf Scv46ohshc Scv46ohshe Scv46ohshd Scv46ohmca (c) Cheatam/kanaly Cv Pack Catalog Nos. Scv47bsmhf Scv47bsmhe Scv47bsmhd Scv47bsmhc (d) Toal/kanaly/randolph Cv Pack Catalog No. Scv47mjnra (e) Cv Set Up Pack Catalog Nos. Scv47bsphl Scv47bsphm Scv47bsphk Scv47bsphi Scv47bsphj Scv47bsphh Scv47bsphg Scv47bsphe (f) Open Heart Pack Catalog Nos. Scv43oh691 Scv43oh692 Scv43oh69c Scv43oh69a Scv43oh69b (g) Knee Pack Catalog No. Sop39knmcf (h) Coronary Tray Catalog Nos. San21ccmlc San23ccmld San23ccmle San21ccmlb San21ccmla (i) Valve Tray Catalog Nos. Scv21hvmlb Scv23hvmlc Scv21hvmla (j) Open Heart Set Up Pack Catalog Nos. Scv47bsccc Scv47bsccd Scv47bsccb Scv47bscca (k) Cv Pack Ii Catalog Nos. Scv46bsmmc Scv46bsmmb Scv46bsmma. Recall #z-220/230-3. | Baxter Healthcare Corporation Mcgaw Park Illinois | Sherwood Medical St. Louis Missouri (catheter). | October 20, 1992 | The argyle 18 french ferguson left vent catheters (sherwood catalog no | 7 260 Custom Packs Were Distributed; Firm Estimates 100 Packs Remain On The Market. | Texas Oklahoma Iowa North Carolina Florida.
1993-01-27 | Devices | II | z-219-3 | Gas Analyzer Used With Certain Lots Of The Bird Sentry Blender Part #10209. The Blender Provides Precise Mixing Of Medical Grade Air And Oxygen Which Is Delivered Via Breathing Circuits To The Patient. Recall #z-219-3. | Bird Products Corporation Palm Springs California | Vascular Technology Inc. Chelmaford Massachusetts (component). | September 10, 1992 | A stray segment, period, or colon may appear on the analyzer display panel | 59 Units Were Distributed. | California Florida Illinois Louisiana New York Minnesota Tennessee Utah England Sweden.
1993-01-27 | Devices | II | z-246-3 | Quartz Micropipette Puller Model P-2000 Used To Pull Glass Or Quartz Tubing Under Controlled Heating Conditions To Make Micropipettes With The Desired Characteristics. Recall #z-246-3. | Manufacturer. Fda Approved The Firm's Corrective Action Plan Dated December 23 1992. Firm-initiated Field Correction Ongoing. | Sutter Instrument Company Novato California. | December 23, 1992 | Noncompliance with performance standard for light-emitting products | Between 11 And 20 Units Were Distributed. | Nationwide And International.
1993-01-27 | Devices | III | z-203-3 | Endotyer Suturing Aid. Recall #z-203-3. | Birtcher Medical Systems Irvine California | Birtcher/solos Medical Systems Duluth Georgia. | November 23, 1992 | The device was marketed without an approved 510(k) | Approximately 69 Units Were Distributed. | Nationwide And International.
1993-01-27 | Devices | III | z-217-3 | Cards O.s. Strep A Test Kit An In-vitro Diagnostic Device And Its Purpose Is To Detect Group A Streptococcal Antigen Directly From Throat Swabs. Recall #z-217-3. | Manufacturer | Pacific Biotech Inc. San Diego California. | December 1, 1992 | The product may exhibiting false positive results | 2 717 Kits Were Distributed. | Nationwide Austria Canada Denmark Finland Italy Spain Sweden Switzerland.