## 1993-05-05
RECALL_CLASSIFICATION_DATE | PRODUCT_TYPE | CLASSIFICATION | RECALL_NUMBER | PRODUCT | RECALLING_FIRM | MANUFACTURER | RECALL_INITIALIZATION_DATE | REASON | VOLUME | DISTRIBUTION
1993-05-05 | Drugs | II | d-245-3 | Peritrate Sa (pentaerythritol Tetranitrate) Sustained Action Tablets In Bottles Of 100 Indicated For The Relief Of Angina Pectoris. Recall #d-245-3. | Parke Davis Division Of Warner Lambert Company Morris Plains New Jersey | Warner Lambert Inc. Vega Baja Puerto Rico. | April 12, 1993 | Lot 16101va does not meet disintegration specifications | Approximately 38 968 Units Were Distributed. | Nationwide.
1993-05-05 | Drugs | II | d-246-3 | Cytotec (misoprostol) Tablets 200 Mcg In Unit Dose Bottles Of 60 And 100 In Cartons Of 100 Unit Dose Tablets And In Physician Sample Packs Of 12 A Rx Oral Tablet Indicated For The Prevention Of Nsaid-induced Gastric Ulcers. Recall #d-246-3. | G.d. Searle & Company Skokie Illinois | G.d. Searle & Company Gaguas Puerto Rico. | April 22, 1993 | Presence of pseudomonas bacteria | 39 276 Physician Samples 368 Unit Dose Cartons 21 621 Bottles Of 100 And 6 128 Bottles Of 60 Were Distributed; Firm Estimates 60 Percent Of The Product Remains On The Market. | Nationwide.
1993-05-05 | Devices | II | z-490-3 | Air-tite Disposable Syringes Sterile 5 Ml (cc) For Human And Veterinary Use Needle Size "21gx1" (.80x25 Mm) For Human And Veterinary Use. Recall #z-490-3. | Air-tite Of Virginia Inc. Virginia Beach Virginia | Mesawa Medical Industry Company Inc. Tokyo Japan. | January 20, 1992 | The syringes may leak. | 230 000 Pieces Were Distributed; Firm Estimates 9 700 Pieces Remain On The Market. | Nationwide.
1993-05-05 | Devices | II | z-491-3 | Sterile Argyle Yankauer Suction Tube With High Capacity Tip. Recall #z-491-3. | Sherwood Medical Company St. Louis Missouri | Sherwood Medical Argyle New York. -2- | March 11, 1993 | Movement of the plastic against cloth garments is cloth garments is capable of generating static electricity, which may obscure the electro- cardiographic signal, making it difficult for the monitor to appropriately time boot compressions | 3 200 Units Were Distributed. The Packages May Contain Microscopic Holes Which May Compromise Their Sterility. _______________ Product Circulator Boot System A Pneumatic Compression Boot System That Is Used Externally On The Leg To Augment Distal Blood Flow. Recall #z-498-3. Code All Serial Numbers. Manufacturer Circulator Boot Corporation Ardmore Pennsylvania. Recalled By Manufacturer By Letter July 31 1992. Firm-initiated Recall Ongoing. Distribution Pennsylvania Virginia Massachusetts Illinois Tennessee New York Florida New Jersey Oklahoma Maryland Missouri Georgia California North Carolina Alabama Washington D.c. Canada Saudi Arabia. Quantity 75 Units Were Distributed. | Tennessee Missouri Arkansas Texas Florida.
1993-05-05 | Devices | II | z-504/505-3 | Imed Gemini Volumetric Infusion Pump And Controller: (a)model Pc-1; (b) Pc-2. Recall #z-504/505-3. | Manufacturer | Imed Corporation San Diego California. | February 3, 1993 | These devices are subject to pumping mechanism failures which may cause uncontrolled excessive fluid delivery | Approximately 62 000 Units Were Distributed. | Nationwide Except North Dakota Rhode Island Wyoming International
1993-05-05 | Devices | II | z-506-3 | Aerochamber Flowsignal With Large Mask Part #80510 A Nebulizer Used To Deliver Liquids In Aerosol Form To Patients For Breathing. Recall #z-506-3. | Manufacturer | Monaghan Medical Corporation Plattsburgh New York. | January 8, 1993 | Stress cracks were found near the "gate" of the molded adapter where the mask is mated to the body of the aerochamber, which could result insignificant leakage of medication from the unit or simply malfunction of the unit if it cracks completely | 209 Cases (10 Units Per Case) Were Distributed. | Nationwide.
1993-05-05 | Devices | II | z-508-3 | Cidex Formula 7* Long-life Activated Dialdehyde Solution In 1 Gallon Containers A Cold Sterilizing Solution. Recall #z-508-3. | Manufacturer | Packaging Corporation Of America Los Angeles California. | February 17, 1993 | Foreign material was introduced into some of the solution | 208 Gallons Were Distributed. | Alabama Arkansas Arizona Colorado Florida Georgia Kansas Louisiana Missouri Mississippi North Carolina New Mexico Oklahoma South Carolina Tennessee Texas Utah.
1993-05-05 | Devices | III | z-488-3 | Colottrac Cwb Controls (citrated Whole Blood Controls) A Lyophilized Citrated Whole Blood Sample Which Can Be Used To Verify The Performance Of The Automated Coagulation Timer And The Applicable Test Cartridges. Recall #z-488-3. | Manufacturer | Medtronic Hemotec Inc. Englewood Colorado. | February 5, 1993 | The controls will produce results which are outside the clotting time ranges printed on the package insert for each cartridge type | 1 035 Controls Were Distributed. | New York Indiana Texas Louisiana California Tennessee Arizona Colorado Illinois South Carolina New Mexico Florida Oregon Kentucky Pennsylvania Montana Canada.
1993-05-05 | Devices | III | z-489-3 | Goldenberg Incus Prosthesis 4.2 Mm Long 1.15 Mm Cannulation Hydroxylapatite/plastic-pore An Implantable Inner Ear Prosthesis. Recall #z-489-3. -4- | Smith & Nephew Richards Inc. Memphis Tennessee | Smith & Nephew Richards Inc. Medical Specialties Division Bartlett Tennessee. | November 12, 1992 | The immediate package container was mislabeled as a shea porp prosthesis | 25 Units Were Distributed. | Nationwide Canada France Germany England Holland.
1993-05-05 | Devices | III | z-492/493-3 | Novalens (rosilfocon A) Rigid Gas Permeable Contact Lenses And Lens Buttons Used In The Manufacture Of Novalens. Recall #z-492/493-3. | Manufacturer | Ocutec Corporation Morrisville North Carolina. | March 18, 1993 | Products were distributed without an approved supplement to their premarket approval for the manufacturing site | 11 757 Buttons Were Used In Making Contact Lenses Distributed By Ocutec Corporation. | Nationwide.
1993-05-05 | Devices | III | z-503-3 | Baxter S/p Caraway Capillary Tubes Heparinized 350 Microliter Volume Used For Micro Blood Sample Collection. Recall #z-503-3. | Manufacturer | Chase Instruments Corporation Rockwood Tennessee. | March 26, 1993 | The ph of heparin solution that is applied to the inner surface of the tubes was found to be out of established ph limits, due to contaminated water | 36 Cases (100 Tubes Per Package 6 Packages Per Case) Were Distributed; Firm Estimates None Remains On The Market. | California Georgia Illinois Louisiana Maryland North Carolina Ohio Texas.
1993-05-05 | Devices | III | z-507-3 | Tryptic Soy Agar (rodac) Contact Plate. Recall #z-507-3. | Manufacturer | Micro Diagnostics Inc. Lombard Illinois. | December 15, 1992 | The plates had trace contamination due to poor aspetic technique during manufacture | 33 Cartons (10 Plates Per Carton) Were Distributed. Firm Estimates That None Of The Product Remains On The Market. | Illinois.