## 1993-07-28
RECALL_CLASSIFICATION_DATE | PRODUCT_TYPE | CLASSIFICATION | RECALL_NUMBER | PRODUCT | RECALLING_FIRM | MANUFACTURER | RECALL_INITIALIZATION_DATE | REASON | VOLUME | DISTRIBUTION
1993-07-28 | Drugs | III | d-296-3 | Ancef Sterile Cefazolin Sodium (lyophilized) Equivalent To 1 Gram Cefazolin Packaged In Glass Vials Of 1 Gram/100 Ml 10 Units Per Packer And 4 Packers Per Outside Shipper. Recall #d-296-3. | Manufacturer | Smith Kline Beecham Pharmaceuticals Conshohocken Pennsylvania. | June 26, 1993 | Some shipping cases are mislabeled as 10 gram | 14 170 Units Were Distributed. | Nationwide.
1993-07-28 | Biologics | III | b-321/322-3 | (a) Platelets; (b) Recovered Plasma. Recall #b-321/322-3. | Manufacturer | American Red Cross Blood Services Miami Florida. | June 19, 1991 | Blood products, corresponding to a unit of red blood cells which was possibly contaminated with microorganisms, were distributed | 1 Unit Of Each Component. | Florida California.
1993-07-28 | Devices | II | z-563-3 | Genesys Gamma Camera For Nuclear Medicine Models 2129- 3011a. Recall #z-563-3. | Manufacturer | Adac Laboratories Milpitas California. | June 1, 1992 | Wear occurs in the older genesys gamma camera's acme head screw assemblies resulting in reduced friction | 317 Units. | Nationwide Europe Indonesia China.
1993-07-28 | Devices | II | z-564-3 | Adac Labs Genesys Gamma Camera Model 21219-3011a For Nuclear Medicine. Recall #z-564-3. | Manufacturer | Adac Laboratories Milpitas California. | February 22, 1992 | A problem exists within the radius drive mechanism of the camera | 298 Units. | Nationwide Europe Indonesia China.
1993-07-28 | Devices | II | z-565-3 | Software Option Pegasys/arc Acquisition Sw V1.81 Rev A. Software Is Used With The Adac Arc-300 Digital Nuclear Imaging System Device. The Software Was Distributed All Under Adac Labels In Each Of The Following Software Packages: 9200 - 0184 Software Option Peg/arc Swv 1.81 Rev. A; 9200 - 0185a Software Update Peg/arc Swv 1.81 Rev A; 9200 - 0174 Software Option Arc Camera V 1.75 Rev A. Recall #z-565-3. | Manufacturer | Adac Laboratories Milpitas California. | July 25, 1991 | The gamma camera for imaging the patient will continue to carry out full or partial instructions causing the camera to move unexpectedly while the computer is storing the acquisition | 43 Packages Were Distributed. | Nationwide.
1993-07-28 | Devices | II | z-589/590-3 | Acs Digoxin Test Kits: (a) Acs Digoxin Combos 50 Test Kit Part #672249; (b) Acs Digoxin Combos 300 Test Kit Part #672248 Used To Monitor Heart Arrhythmias. Recall #z-589/590-3. | Manufacturer | Ciba Corning Diagnostics Corporation E. Walpole Massachusetts. | May 11, 1992 | Discrepant values exist between the acs digoxin method and reference methods | 292 Kits Were Distributed. | Nationwide Australia Italy Finland England Belgium Canada Venezuela.
1993-07-28 | Devices | II | z-611/612-3 | Serono Hcg Maiaclone Used For The Quantitative Determination Of Human Chorionic Gonadotropin In Serum And Plasma And The Qualitative Determination Of Hcg In Serum Plasma And Urine: (a) 150 Test Kit Part #472402; (b) 300 Test Kit Part #472400. Recall #z-611/612-3. | Ciba Corning Diagnostics Corporation E. Walpole Massachusetts | Serono Diagnostics S.a. Swtizerland. | August 21, 1991 | The products may yield falsely high test results with certain types of patient samples | 3 049 Kits Were Distributed. | Nationwide.
1993-07-28 | Devices | II | z-615-3 | Mfhd Formatter Model 46-308179p2 Used With Ge Nuclear Medicine Systems And For Making Hardcopy Film Images Of Video Output. Recall #z-615-3. | Ge Medical Systems Waukesha Wisconsin | Camtronix Ltd. Medical Systems Hartland Wisconsin. | September 22, 1992 | Possible display of incorrect information | 150 Units Were Distributed. | Nationwide And International.
1993-07-28 | Devices | II | z-617/619-3 | Various Sterile Devices: (a) Snap Kovers Sterile Products Used In Surgical Procedures To Establish A Sterile Barrier Between The Operating Site And Any Instruments-of Usage (non-sterile) That May Be Used In This Surgical Procedure; (b) Decanting Devices - Bag Decanter Vial Decanters And Transfer Device Sterile Devices Used In Decanting Sterile I.v. Solutions To Patients; (c) Drape Sheets Sterile Devices Used To Establish A Sterile Barrier Between Patient And Operating Site During Surgery. Recall #z-617/619-3. | Manufacturer | Advance Medical Designs Inc. Marietta Georgia. | April 12, 1993 | Labeling of some product as sterile prior to undergoing sterilization process making the product nonsterile | Undetermined. | Nationwide.
1993-07-28 | Devices | II | z-648-3 | Westmed Pacing Coil Temporary Bipolar Epicardial Pacing Device Product #wpc-1. Recall #z-648-3. | Deknatel Inc. Fall River Massachusetts | West Med Osborne Park Western Australia. | January 29, 1993 | Some of the electrodes fail to conduct | 2 456 Units Were Distributed. | Nationwide Canada.
1993-07-28 | Devices | III | z-591/594-3 | Magic T4 Standards And Kits Used For The Quantitative Determination Of Thyroxine In Serum Or Plasma: (a) Magic T4 Standard Part #472304; (b) Magic T4 Kits Part #472300; (c) Magic T4 Kits Part #472318; (d) Magic T4 Kits Part #472349. Recall #z-591/594-3. | Manufacturer | Ciba Corning Diagnostics Corporation E. Walpole Massachusetts. | NOT-FOUND | The product is unstable and may deteriorate | 1 008 Units And 933 Vials Sold As Part Of Kits Were Distributed. | Nationwide And International.
1993-07-28 | Devices | III | z-595/596-3 | Acs Multicalibrator B2 & B4 Packs Used As A High And Low Calibrator For The Acs System For Different Assayed Products: (a) B2 Pack Part #672171; (b) B6 Pack Part #672181. Recall #z-595/596-3. | Manufacturer | Ciba Corning Diagnostics Corporation E. Walpole Massachusetts. | March 19, 1992 | The high control values are in the range of 26, not the required 28 | 721 Kits Were Distributed. | Nationwide And International.
1993-07-28 | Devices | III | z-597/599-3 | Magic T3 Uptake Radioassay Used For The Assessment Of Unsaturated Thyroxine Binding Proteins (t3-uptake) In Serum: (a) Magic T3 Uptake Radioassay Standards Part #472289; (b) Magic T3 Uptake Radioassay Standards Part #472322; (c) Magic T3 Uptake Radioassay Kit Part #472320. Recall #z-597/599-3. | Manufacturer | Ciba Corning Diagnostics Corporation E. Walpole Massachusetts. | November 3, 1992 | Product has low control values that may result in a decrease in patient and control values of approximately 6% | 1 210 Kits Of Standards 741 And 702 Kits Of Standards 624 Were Distributed. | Nationwide And International.
1993-07-28 | Devices | III | z-600/601-3 | Acs/tsh Assay An In-vitro Diagnostic: (a) Acs/tsh Assay Combos 50 Test Kit Part #672221; (b) Acs/tsh Assay Combos 300 Test Kit Part #672220. Recall #z-600/601-3. | Manufacturer | Ciba Corning Diagnostics Corporation E. Walpole Massachusetts. | February 26, 1992 | Testing commercial control recoveries of this product ran higher than the firm's specification | 3 882 Kits Were Distributed; Firm Estimates None Remains On The Market. | Nationwide Germany France Hong Kong Canada Belgium Italy Australia Sweden.
1993-07-28 | Devices | III | z-602-3 | Magic Mab Tsh Control In 10 Ml Vials Low Control Used For The Quantitative Determination Of Thyroid Stimulating Hormone In Serum. Recall #z-602-3. | Manufacturer | Ciba Corning Diagnostics Corporation E. Walpole Massachusetts. | November 12, 1992 | The product was mislabeled with values of 0 | 10 Vials Were Distributed. | New Jersey Indiana Florida New Hampshire Minnesota.
1993-07-28 | Devices | III | z-603-2 | Magic Ferritin Bulk Antibody And Tracer Used For The Quantitative Determination Of Ferritin In Serum And Plasma. Recall #z-603-2. | Manufacturer | Ciba Corning Diagnostics Corporation E. Walpole Massachusetts. | March 31, 1991 | The antibody did not meet the high dose hook claim (60,000) mg/ml as stated in the product insert | 114 Units Were Distributed. | Colorado Texas Arizona.
1993-07-28 | Devices | III | z-604/610-3 | The Immophase (125i) Folate Radioassay Used For The Quantitative Determination Of N5-methyltetrahy-drofolic Acid(mthfa) In Serum Plasma And Red Blood Cells: (a) Immophase B12 Folate Combined Kits Part # 474272 (b) Immophase Separate B12 Kits Part # 474276 (c) Immophase Separate Folate Kits Part # 474273 (d) Baxter B12/folate Kits Part # 474781 (e) Baxter (japanese) B/12 Folate Combined Kit Part # 474787 (f) B/12 Folate Standards Part # 474267 (g) B/12 Folate Controls Part # 474271. Recall #z-604/610-3. | Manufacturer | Ciba Corning Diagnostics Corporation E. Walpole Massachusetts. | 1990-09-27 | B-12 values were 15/20% low in the deficient and indeterminate range of the assay | Firm Estimates None Remains On The Market. | Nationwide Japan France.
1993-07-28 | Devices | III | z-613/614-3 | The Magic T3 (125i) Radioimmunoassay Intended For The Quantitative Determination Of Triiodothyronine (t3) In Serum Or Plasma: (a) Standards Part #472308; (b) Kit Part #472305. Recall #z-613/614-3. | Ciba Corning Diagnostics Corporation E. Walpole Massachusetts | Ciba Corning Diagnostics Corporation Medfield Massachusetts. | May 30, 1991 | The 0.25 ng/dl standard value of certain lots was misassigned and caused a poor backfit. This low standard may result in inaccurate test results. | Standards Lot 10151 Were Either Shipped As Standards (513 Units) Or Part Of A Kit As Follows: Lot 02907 - 98 Kits; 05901 - 741 Kits; 08603 - 617 Kits. Total 1 456. Total Standards Shipped From Lot 10151 - 513. Standards Lot 10631 Were Either Shipped As Standards (589 Units) Or As Part Of A Kit As Follows: Lot 08405 - 20 Kits; 10206 - 29 Kits; 10709 - 35 Kits; 11302 - 22 Kits; 11508 581 Kits. Total 687 Kits. Standards Shipped From Lot 10631 = 589. | Nationwide And International.