## Report-1996-05-15
recall_classification_date | product_type | classification | recall_number | product | recalling_firm | manufacturer | recall_initiation_date | reason | volume | distribution
1996-05-15 | Drugs | II | d-136-6 | Medline Epi-clenz Instant Antiseptic Hand Cleanser, An Otc Liquid For Handwashing To Decrease Bacteria On The Skin | Medline Industries | Polychem Corporation, New Haven, Connecticut | April 20, 1996 | Some Bottles Were Mislabeled As Tlc Mouth Rinse | 297 cases were distributed; firm estimated that 25 percent of product remained on market at time of recall initiation | Nationwide
1996-05-15 | Drugs | II | d-137-6 | Co2 Compressed Gas | Manufacturer | Mg Industries, Inc | April 23, 1996 | Presence Of Trace Quantities Of Hcn (hydrogen Cyanide) | undetermined | Pennsylvania, Maryland, West Virginia, Virginia, Tennessee, Kentucky, Georgia, North Carolina, New Jersey, Bermuda
1996-05-15 | Drugs | II | d-138-6 | Carbon Dioxide Usp, 50 Pound And 7 Pound Cylinder Capacity | Repacker | Sunox, Inc | May 2, 1996 | Presence Of Trace Quantities Of Hcn (hydrogen Cyanide) | approximately 1,200 cylinders (50 pound capacity) and 159 cylinders (7 pound capacity) were produced | North Carolina, Virginia
1996-05-15 | Drugs | III | d-133-6 | Cytarabine Injection 20 Mg/ml -- 1000mg/50ml; Cytarabine Injection 20 Mg/ml -- 100mg/5ml, Rx Drug For Iv Or Subcutaneous Use | Faulding Pharmaceutical Company (purepac) | Faulding Pharmaceutical Company, David Bull Laboratories, Mulgrave, Victoria 3170, Australia | March 8, 1996 | Particulate Matter Formation (calcium And Silicate) | 32,400 bottles were distributed; firm estimated that 50-60 percent of product remained on market at time of recall initiation | Nationwide
1996-05-15 | Drugs | III |  | Barre-national (a) Bromanate Elixir In 4 Ounce Bottles, Otc Decongestant/antihistamine, Under The Albertsons And Janet Lee Labels; -6- (b) Tussex Cough Syrup, In 4 Ounce Bottles, Otc Cough Syrup, Cough Suppressant, Nasal Decongestant, Expectorant, Under The Moore And Major Labels | Manufacturer | Barre-national, Inc | April 26, 1996 | Product Labeling Not In Accord With Otc Monograph (missing Warning Statement) | firm estimates none remains on the market | California, Connecticut, Florida, Michigan, Texas
1996-05-15 | Biologics | II | b-393-6 | Platelets | Manufacturer | United Blood Services, Chicago, Illinois | February 12, 1996 | Blood Product, Prepared From A Unit Of Whole Blood Designated As A "difficult" Collection, Was Distributed. | 1 unit | Illinois
1996-05-15 | Biologics | III | b-366-6 | Red Blood Cells | Manufacturer | United Blood Services, Chicago, Illinois | February 8, 1996 | Red Blood Cells, Stored At Room Temperature For An Unknown Amount Of Time, Were Distributed | 2 units | Illinois
1996-05-15 | Biologics | III | b-395-6 | Recovered Plasma | Manufacturer | American Red Cross Blood Services, Peoria, Illinois | October 30, 1995 | Blood Product, Corresponding To A Unit Of Red Blood Cells Implicated In A Febrile Transfusion Reaction, Was Distributed | 1 unit | California
1996-05-15 | Devices | II | z-640/646-6 | Sterile Intravenous Administration Iv Sets With 1 | Mcgaw | Mcgaw Of Puerto Rico, Inc | January 19, 1996 | An Inadequate Bond/design May Cause Tube/filter Separation, Resulting In Leakage | 21,145 cases (895,324 units) were distributed | Nationwide, Puerto Rico, Canada
1996-05-15 | Devices | II | z-724-6 | Tanfit 24 Suntanning Bed | Manufacturer | Puretan International/asterisk, Inc | November 14, 1995 | Eighteen (18) Gauge Wire Was Used In The Timers In Lieu Of 12 Gauge Which Caused The Timers To Shut The Lamps Off Prematurely | 38 units | Nationwide
1996-05-15 | Devices | II | z-727-6 | Lumin Laparoscopic Uterine Manipulator Injector, Catalog #mis-100, Used For Manipulation Of The Uterus During Diagnostic And Surgical Procedures | Manufacturer | Utah Medical Products, Inc | March 25, 1996 | The Sterility Of The Device May Be Compromised Due To Loss Of Package Integrity | 305 units were distributed | Arizona, California, Indiana, Kansas, Maine, Michigan, Tennessee, Texas, Virginia, And International
1996-05-15 | Devices | II | z-742/746-6 | Puritan Bennett 7200 Series Ventilators, Indicated For Delivery Of Gas To Patients Dependent On Artificial Respiration: (a) Model No | Manufacturer | Nellcor Puritan Bennett (npb), Carlsbad, California | October 2, 1995 | Device Contains A Defective Pressure Transducer, Causing It To Stop Ventilation At Set Parameters | 2,987 units were distributed | Nationwide And International
1996-05-15 | Devices | II | z-748/749-6 | Biostar Strep A Oia Test, Intended For In- Vitro Diagnostic Use As An Aid In The Diagnosis Of Group A Streptococcal Pharyngitis: (a) Catalog No | Manufacturer | Biostar, Inc | April 18, 1996 | The Test Kits May Produce False Positive Results When Performed At The Maximum Assay Times Stated In The Package Insert | firm estimates 125,000 test devices or approximately 1,250 100-test kits were distributed | Nationwide And International
1996-05-15 | Devices | II | z-750/751-6 | Action Cat Brand Power Scooter, Three-wheeled Wheel Chair: (a) Cat Traditional; (b) Cat Aviation | Manufacturer | Invacare Corporation, Elyria, Ohio | January 19, 1996 | The Brake Key (pin) Which Connects The Motor Shaft To The Scooter Brake May Slide Out Of Position Causing The Brake To Become Non- Functional | 671 scooters were distributed | Nationwide
1996-05-15 | Devices | II | z-752/753-6 | Uroview Diagnostic Imaging Systems | Manufacturer. Fda Approved The Firm's Corrective Action Plan November 24 | Oec Medical Systems, Salt Lake City, Utah | November 24, 1993 | The Devices Were Found To Be In Noncompliance With 21 Cfr 1032(a) And 21 Cfr 1020 | 61 units | Nationwide And International
1996-05-15 | Devices | II | z-754-6 | Model 9400 Mobile Fluoroscopy Imaging System | Manufacturer. Fda Approved The Firm's Corrective Action Plan November 17 | Oec Medical Systems, Salt Lake City, Utah | November 17, 1993 | The Device Was Found To Be In Noncompliance With 21 Cfr 1032(c) In That Wires In The X-ray Control May Cause An Accidental Radiation Occurrence (aro) Resulting In The X-ray Unit | 767 units | Nationwide And International
1996-05-15 | Devices | II | z-755-6 | Aruba Suntanning Bed | Manufacturer | Puretan International/asterisk, Inc | May 16, 1994 | The Shaft On Timer Was Too Short And This Caused The Knob Of The Timer To Jam When The Temperature Increased And Expanded The Plastic Canopy Of The Device | 368 units were distributed | Nationwide And Canada
1996-05-15 | Devices | II | z-756-6 | Ophthalas 532 Laser System, Used In Ophthalmic Surgery | Manufacturer. Fda Approved The Firm's Corrective Action Plan May 2 | Alcon Laboratories, Inc | May 2, 1996 | The Device Was Defective Within 21 Cfr 1003 | 848 units | Nationwide And International
1996-05-15 | Devices | II | z-757-6 | Tri-rolls Brand Power Scooter, A Three-wheeled Wheelchair | Manufacturer | Invacare Corporation, Elyria, Ohio | February 21, 1994 | Under Severe Use Or Heavy Abuse Tiller Spring That Turns The Tiller Throttle Can Break And/or Slip Off, Due To A Metal Burr, Causing The Scooter To Remain In Drive Unless The Drive Control Is Manually Moved Into Neutral Position | 5,631 scooters were distributed | Nationwide
1996-05-15 | Devices | II | z-764/765-6 | Hydromat Hydrovalve And Tubing Set For Laparoscopic Suction And Irrigation: (a) Hydromat Hydrovalve & Tubing Set, Catalog #263101-61, Series 1 Set; (b) Hydromat Hydrovalve & Tubing Set, Catalog #263104-61, Series 4 Set | Karl Storz Endoscopy | Infinity Extrusion And Engineering, Inc | January 30, 1996 | Device Sterility May Be Compromised Due To A Defective Packaging Seal | 720 sets were distributed | California, Colorado, Florida, Georgia, Kentucky, Massachusetts, Maryland, New York, Ohio, Oklahoma, Pennsylvania, South Carolina, Texas, Virginia, Washington State, West Virginia
1996-05-15 | Devices | III | z-581-6 | Sentinel-seal Autotransfusion Chest Drainage System (cdu), Catalog #8888-571489, A Sterile, Single Use, Prescription Device | Manufacturer | Sherwood Medical Company, Commerce, Texas | October 12, 1994 | Device Is Subject To Air Leak Which May Occur Between The Patient Connection Site And The Chest Drainage System Collection Chamber | 2,680 units were distributed | Belgium, France, Greece, Italy, Spain
1996-05-15 | Devices | III | z-726-6 | Baxter Myelogram Tray, Containing An 18 Gauge X 3-1/2" Spinal Needle With Stylet And Nine- Holed Aspiration Needle, Used To Obtain Spinal Fluid Samples And Removal Of Oil-based Contrast Media During Myelograms | Baxter Healthcare Corporation | Baxter Healthcare Corporation, Mannford, Oklahoma | April 19, 1996 | The Aspiration Cannula May Not Fit Inside The Spinal Needle As Intended | 470 cases were distributed; firm estimated that 50 percent of product remained on market at time of recall initiation | Nationwide
1996-05-15 | Devices | III |  | Schneider Wallstent Esophageal Prosthesis, Used To Keep The Esophagus Of A Patient Open When There Are Strictures Or Fistulas Resulting From Cancer: (a) Model E38-2040-95, 40mm Long; (b) Model E38-2060-95, 60mm Long; (c) Model E38-2090-95, 90mm Long |  | Schneider (usa), Inc | N/A | There Is A Possibility Of Increased Risk Of Perforation And Bleeding In Patients Who Have Received Or Will Receive Radiation Therapy Following Stent Placement, And To Those Patients With Hiatal Hernia And/or Gastric Prolapse Into The Esophageal Area For Stents Implanted Near The Gastroesophageal Junction | 4,577 prostheses were distributed | Nationwide, Argentina, Australia, Canada, Hong Kong