## Report-1996-06-05
recall_classification_date | product_type | classification | recall_number | product | recalling_firm | manufacturer | recall_initiation_date | reason | volume | distribution
1996-06-05 | Drugs | III | d-150-6 | J. b. Williams Cepacol Maximum Strength Sore Throat Spray (cool Menthol) | J. b. Williams Company | Les Emballages Knowlton, Knowlton, Quebec, Canada | September 5, 1995 | Subpotency. | 58,212 units were distributed | Nationwide
1996-06-05 | Devices | I = | z-729/741-6 | Avian Transport Adult And Pediatric Ventilator (a) Avian Ventilator, Part No | Manufacturer | Bird Products Corporation, Palm Springs, California | April 1, 1996 | All Device Models Have Oxygen Bleeding Or Leaking Internally, Which Is Not Being Adequately Vented To The Outside Of The Case. This Causes Elevated Oxygen Concentrations Within The Device When The Unit Is Operated With An Enriched Gas Supply (e. g. , 40% - 100% Oxygen). Furthermore, One Particular Brand Of Capacitor With An Inadequate Voltage Specification, Found In Approximately 120 To 150 Of The Devices, At One Circuit Board Location, Fails Due To Overheating, Which In The Presence Of The Elevated Oxygen Concentrations, Causes Combustion Within The Case. | approximately 3,200 units were distributed | Nationwide And International
1996-06-05 | Devices | II | z-781/786-6 | Receptal Suction Liner System, Used With The Receptal Hard Canister For Vacuum Suction Waste Body Fluid Management In Hospital Operating Rooms: (a) List 43024- 01: 2000 Ml Liner/lid, 1/4" Port V3024-01: Vha+plus Label; (b) List 43027-01: 2000 Ml Liner/lid With Pour Spout (arthroscopy) V3027-01- Vha+plus Label; (c) List 43042-01: 2000 Ml Liner/lid With Vac-gard Internal Non-mechanical Shut Off Valve/filter V3042-01-vha+plus Label; (d) List 43044-05: 2000 Ml Liner/lid With Vac-gard Internal Non-mechanical Shut Off Valve/filter, Pour Spout V3044-05- Vha+plus Label; (e) List 43043-01: 1500 Ml Liner/lid With Vac-gard Internal Non-mechanical Shut Off Valve/filter V3043-01 Vha+plus Label; -2- (f) List 43044-01: 1500 Ml Liner/lid With Vac-gard Internal Non-mechanical Shut Off Valve/filter, Pour Spout | Abbott Laboratories | Abbott Laboratories, Salt Lake City, Utah | November 8, 1995 | The Receptal Liner Lids May Pop Off The Suction Canisters Compromising The Function Of The Device | approximately 8,000,000 units are distributed annually; firm estimates none remains on the market | Nationwide
1996-06-05 | Devices | II | z-841-6 | Calibration Gas Cylinders (10% 02, Balance N2), Used To Calibrate And Check An Oxygen/carbon Dioxide Analyzer | Manufacturer | Scott Specialty Gases, Inc | April 13, 1995 | The Cylinders Labeled As 10% Oxygen/balance Nitrogen, Actually Contain 10% Carbon Dioxide/balance Nitrogen | 36 cylinders were distributed | Washington State, Indiana, New York, Iowa, Kansas, Wisconsin, Korea, Canada
1996-06-05 | Devices | II | z-842/845-6 | Bird Ventilators, Indicated For Delivery Of Medical Gases To Patients Dependent On Artificial Respiration: (a) Model No | Bird Products Corporation | Vernay/sfernice, Div | December 21, 1995 | The Devices Contain A Defective Potentiometer, Which Can Cause The Ventilators To Cease Operating While Alarms Remain Functional | 799 ventilators and upgrade kits are subject to recall | Nationwide And International