## Report-1996-10-02
recall_classification_date | product_type | classification | recall_number | product | recalling_firm | manufacturer | recall_initiation_date | reason | volume | distribution
1996-10-02 | Drugs | I = | d-270-6 | Apothe'cure Dmso (dimethylsulfoxide) For Injection, In 30 Ml Vials | Manufacturer | Apothe'cure, Inc | August 5, 1996 | Fda Analysis Found Samples Contaminated With Bacteria; And Fda Inspection Uncovered Deficiencies Of Good Manufacturing Practices. | undetermined | Nationwide And International
1996-10-02 | Drugs | II | d-271-6 | Altaire Pharmaceutical Brand Clear Cough Cough Suppressant Expectorant (guaifenesin/ Dextromethorphan), Oral Liquid, In 4 Fluid Ounce Plastic Bottles | Manufacturer | Altaire Pharmaceutical, Inc | August 22, 1996 | Product Is Misbranded Because The Dosing Cup Does Not Bear Measurements That Are Consistent With The Labeled Dosing Instructions | 12,316 units were distributed; firm estimated that 200 units remained on market at time of recall initiation | New York, New Jersey, Ohio, North Carolina, Texas, Indiana, Wisconsin, Michigan, Alabama, California, District Of Columbia
1996-10-02 | Drugs | II | d-273/274-6 | Altaire Pharmaceuticals Otc Pain Relievers, In 4 Fluid Ounce Plastic Bottles: (a) Apra Cherry-flavored Children's Acetaminophen Elixir; (b) Altenol Grape-flavored Children's Acetaminophen Elixir | Manufacturer | Altaire Pharmaceutical, Inc | August 22, 1996 | Products Are Misbranded Because The Dosing Cup Does Not Bear Measurements That Are Consistent With The Labeled Dosing Instructions | (a) 67,945 units; (b) 23,012 units were distributed; firm estimated that 12,000 units combined remained on market at time of recall initiation | Alaska, Arizona, Arkansas, California, Florida, Illinois, Indiana, Michigan, Minnesota, New York, Texas, Puerto Rico
1996-10-02 | Drugs | III | d-272-6 | Tussilan Liquid, Rx, In 1 Ounce, 4 Ounce And 128 Ounce Containers, Antitussive, Expectorant, Nasal Decongestant, Antihistamine, Used In The Treatment Of Colds | Manufacturer | Sein Mendez Laboratories, Rio Piedras, Puerto Rico | August 14, 1996 | Product Does Not Meet Potency Specifications. | 12,731 1-ounce units; 9,405 4-units; 172 units were distributed | Puerto Rico
1996-10-02 | Devices | II | z-1202-6 | Lifestyles Condoms Lubricated With Nonoxynol- 9 | Manufacturer | Ansell, Inc | April 19, 1996 | The Condoms Failed The Water Leak Test Apparently Due To Deterioration Of The Latex Because Of Improper Storage Conditions | approximately 78,000 condoms were distributed; firm estimates none remains on the market | California, Florida, Georgia, Massachusetts, New Jersey, New York, Ohio, Oregon, Texas, Vermont, Washington State
1996-10-02 | Devices | II | z-1228/1233-6 | Hand Held Controllers Used With Maxxus And Optima Nuclear Medicine Imaging Systems, Used For Patient Imaging: -3- (a) Model No | General Electric Company | Ige (international General Electric) Ltd | August 8, 1994 | After The Motion Button Is Released, The Imaging System's Table Or Detector Head May Fail To Stop | 139 units were distributed | Nationwide And International
1996-10-02 | Devices | II | z-1234-6 | Software Version 4 | Manufacturer | Althin Medical, Inc | August 5, 1996 | The Formaldehyde And Actril Chemical Disinfect Modes May Not Function Properly Due To Software Problems | 437 units and 52 kits were distributed | Arizona, California, Connecticut, Florida, Georgia, Indiana, Iowa, Maine, Massachusetts, Missouri, New Jersey, New York, North Carolina, Ohio, Tennessee, Texas, Washington State, And International
1996-10-02 | Devices | III | z-1110-6 | Sabratek Model 3030 Volumetric Infusion Pumps, Used In Hospitals And Home Care For Delivery Of Intravenous Fluids For Inotropics, Antibiotics, Tpn, Igg, Cytogram, Pentamidine And Hydration Therapies | Sabratek Corporation | Sabratek Corporation, Niles, Illinois | August 2, 1995 | Two Capacitors, C2 And C8, Were Misaligned During Assembly Of The Logic Printed Circuit Board, Causing Reversed Polarity, Which Over A Period Of Time Could Cause Capacitors To Initiate Breakdown, Resulting In A Change In The Motor Signal | 45 pumps were distributed | Ohio, Massachusetts, Illinois, California, Florida, Texas, Pennsylvania
1996-10-02 | Devices | III | z-1111-6 | Newvues (vifilcon A) Soft Contact Lens, Hydrophilic Contact Lens | Manufacturer | Ciba Vision Corporation, Atlanta, Georgia | August 16, 1996 | Lenses May Contain A Sphere Power Other Than That Reflected In The Labeling | 1,648 multipacks (9,888 lenses) were distributed | Nationwide And International
1996-10-02 | Devices | III | z-1215-6 | Lifepak 10 Defibrillator/monitor/pacemaker | Manufacturer | Physio Control Corporation, Redmond, Washington | August 7, 1996 | There Was A Potential For Pacing Function To Cease Operation | 13 units were distributed | Iowa, Massachusetts, North Carolina, New York, Ohio, Pennsylvania, Rhode Island, Wisconsin
1996-10-02 | Devices | III | z-1220-6 | Incstar Cmv-vue Test Kit, Catalog #24300, For The Direct Qualitative Detection Of Early Structural Protein Of Cmv (cytomegalovirus) | Manufacturer | Incstar Corporation, Stillwater, Minnesota | September 9, 1996 | There Was A Decreased Staining Intensity In The Positive Control | 67 kits were distributed | Arkansas, Illinois, Maryland, New Mexico, New York, North Carolina, Tennessee, Australia, Germany, Italy, Korea, Portugal, Spain
1996-10-02 | Devices | III | z-1221/1227-6 | Vasoseal Vascular Hemostasis Device (vhd) Kits, Indicated For Use In Reducing Time To Hemostasis At The Femoral Arterial Puncture Side In Patients Who Have Undergone Diagnostic Angiography Or Percutaneous Transluminal Coronary Angioplasty Procedures Using An 8 French Or Smaller Procedural Sheath: Catalog Numbers: 75301, 75302, 75303, 75304, 75305, 75306, 75307 | Datascope Corporation | Bioplex Medical B | February 20, 1996 | There May Be A Missing Component In The Kit | 5,970 kits were distributed | Nationwide
1996-10-02 | Drugs | I = | d-270-6 | Apothe'cure Dmso (dimethylsulfoxide) For Injection, In 30 Ml Vials | Manufacturer | Apothe'cure, Inc | August 5, 1996 | Fda Analysis Found Samples Contaminated With Bacteria; And Fda Inspection Uncovered Deficiencies Of Good Manufacturing Practices. | undetermined | Nationwide And International
1996-10-02 | Drugs | II | d-271-6 | Altaire Pharmaceutical Brand Clear Cough Cough Suppressant Expectorant (guaifenesin/ Dextromethorphan), Oral Liquid, In 4 Fluid Ounce Plastic Bottles | Manufacturer | Altaire Pharmaceutical, Inc | August 22, 1996 | Product Is Misbranded Because The Dosing Cup Does Not Bear Measurements That Are Consistent With The Labeled Dosing Instructions | 12,316 units were distributed; firm estimated that 200 units remained on market at time of recall initiation | New York, New Jersey, Ohio, North Carolina, Texas, Indiana, Wisconsin, Michigan, Alabama, California, District Of Columbia
1996-10-02 | Drugs | II | d-273/274-6 | Altaire Pharmaceuticals Otc Pain Relievers, In 4 Fluid Ounce Plastic Bottles: (a) Apra Cherry-flavored Children's Acetaminophen Elixir; (b) Altenol Grape-flavored Children's Acetaminophen Elixir | Manufacturer | Altaire Pharmaceutical, Inc | August 22, 1996 | Products Are Misbranded Because The Dosing Cup Does Not Bear Measurements That Are Consistent With The Labeled Dosing Instructions | (a) 67,945 units; (b) 23,012 units were distributed; firm estimated that 12,000 units combined remained on market at time of recall initiation | Alaska, Arizona, Arkansas, California, Florida, Illinois, Indiana, Michigan, Minnesota, New York, Texas, Puerto Rico
1996-10-02 | Drugs | III | d-272-6 | Tussilan Liquid, Rx, In 1 Ounce, 4 Ounce And 128 Ounce Containers, Antitussive, Expectorant, Nasal Decongestant, Antihistamine, Used In The Treatment Of Colds | Manufacturer | Sein Mendez Laboratories, Rio Piedras, Puerto Rico | August 14, 1996 | Product Does Not Meet Potency Specifications. | 12,731 1-ounce units; 9,405 4-units; 172 units were distributed | Puerto Rico
1996-10-02 | Devices | II | z-1202-6 | Lifestyles Condoms Lubricated With Nonoxynol- 9 | Manufacturer | Ansell, Inc | April 19, 1996 | The Condoms Failed The Water Leak Test Apparently Due To Deterioration Of The Latex Because Of Improper Storage Conditions | approximately 78,000 condoms were distributed; firm estimates none remains on the market | California, Florida, Georgia, Massachusetts, New Jersey, New York, Ohio, Oregon, Texas, Vermont, Washington State
1996-10-02 | Devices | II | z-1228/1233-6 | Hand Held Controllers Used With Maxxus And Optima Nuclear Medicine Imaging Systems, Used For Patient Imaging: -3- (a) Model No | General Electric Company | Ige (international General Electric) Ltd | August 8, 1994 | After The Motion Button Is Released, The Imaging System's Table Or Detector Head May Fail To Stop | 139 units were distributed | Nationwide And International
1996-10-02 | Devices | II | z-1234-6 | Software Version 4 | Manufacturer | Althin Medical, Inc | August 5, 1996 | The Formaldehyde And Actril Chemical Disinfect Modes May Not Function Properly Due To Software Problems | 437 units and 52 kits were distributed | Arizona, California, Connecticut, Florida, Georgia, Indiana, Iowa, Maine, Massachusetts, Missouri, New Jersey, New York, North Carolina, Ohio, Tennessee, Texas, Washington State, And International
1996-10-02 | Devices | III | z-1110-6 | Sabratek Model 3030 Volumetric Infusion Pumps, Used In Hospitals And Home Care For Delivery Of Intravenous Fluids For Inotropics, Antibiotics, Tpn, Igg, Cytogram, Pentamidine And Hydration Therapies | Sabratek Corporation | Sabratek Corporation, Niles, Illinois | August 2, 1995 | Two Capacitors, C2 And C8, Were Misaligned During Assembly Of The Logic Printed Circuit Board, Causing Reversed Polarity, Which Over A Period Of Time Could Cause Capacitors To Initiate Breakdown, Resulting In A Change In The Motor Signal | 45 pumps were distributed | Ohio, Massachusetts, Illinois, California, Florida, Texas, Pennsylvania
1996-10-02 | Devices | III | z-1111-6 | Newvues (vifilcon A) Soft Contact Lens, Hydrophilic Contact Lens | Manufacturer | Ciba Vision Corporation, Atlanta, Georgia | August 16, 1996 | Lenses May Contain A Sphere Power Other Than That Reflected In The Labeling | 1,648 multipacks (9,888 lenses) were distributed | Nationwide And International
1996-10-02 | Devices | III | z-1215-6 | Lifepak 10 Defibrillator/monitor/pacemaker | Manufacturer | Physio Control Corporation, Redmond, Washington | August 7, 1996 | There Was A Potential For Pacing Function To Cease Operation | 13 units were distributed | Iowa, Massachusetts, North Carolina, New York, Ohio, Pennsylvania, Rhode Island, Wisconsin
1996-10-02 | Devices | III | z-1220-6 | Incstar Cmv-vue Test Kit, Catalog #24300, For The Direct Qualitative Detection Of Early Structural Protein Of Cmv (cytomegalovirus) | Manufacturer | Incstar Corporation, Stillwater, Minnesota | September 9, 1996 | There Was A Decreased Staining Intensity In The Positive Control | 67 kits were distributed | Arkansas, Illinois, Maryland, New Mexico, New York, North Carolina, Tennessee, Australia, Germany, Italy, Korea, Portugal, Spain
1996-10-02 | Devices | III | z-1221/1227-6 | Vasoseal Vascular Hemostasis Device (vhd) Kits, Indicated For Use In Reducing Time To Hemostasis At The Femoral Arterial Puncture Side In Patients Who Have Undergone Diagnostic Angiography Or Percutaneous Transluminal Coronary Angioplasty Procedures Using An 8 French Or Smaller Procedural Sheath: Catalog Numbers: 75301, 75302, 75303, 75304, 75305, 75306, 75307 | Datascope Corporation | Bioplex Medical B | February 20, 1996 | There May Be A Missing Component In The Kit | 5,970 kits were distributed | Nationwide