## Report-1996-10-22
recall_classification_date | product_type | classification | recall_number | product | recalling_firm | manufacturer | recall_initiation_date | reason | volume | distribution
1996-10-22 | Drugs | I = | d-254-6, adrenal cortex injection, which appeared in the october 16, 1996 enforcement report, should read as follows: _______________ product all sizes and strengths of adrenal cortex injection, a rx injectable drug, labeled as distributed by hallmark labs, chicago, il | All Sizes And Strengths Of Adrenal Cortex Injection, A Rx Injectable Drug, Labeled As Distributed By Hallmark Labs, Chicago, Il | Phyne Pharmaceuticals | (primary) Ontor Beauty Products, Inc | October 1, 1996 | Mycobacterium Abscessus Contamination Of Unopened Vial Of Hallmark Labs' Labeled Product Distributed By Phyne Pharmaceuticals And Epidemiological Data Collected By Cdc Correlating The Use Of This Product With Mycobacterium Abscesses At The Site Of Injection. | undetermined | Nationwide
1996-10-22 | Devices | II | z-001-7 | Gemini Pc-1 Volumetric Infusion Pump/controller, European 220 Volt Model V8 | Manufacturer | Imed Corporation, San Diego, California | June 18, 1996 | The Pumps Contain A Faulty Toroid (inductor) Which May Heat Up Adjacent Wires Causing An Electrical Short | 645 units were distributed | International
1996-10-22 | Devices | II | z-007/014-7 | Max-i-probe Endodontic Irrigation Syringe: (a) Catalog # Maxe125, 3cc Syringe, 25g Probe, 1-2% Naoc1 Solution (b) Catalog # Maxe625, 3cc Syringe, 25g Probe, 4-6% Naoc1 Solution (c) Catalog # Maxe628, 3cc Syringe, 28g Probe, 4-6% Naoc1 Solution (d) Catalog # Maxe630, 3cc Syringe, 30g Probe, 4-6% Naoc1 Solution (e) Catalog # Maxe650, Trial Kit Of 2 Trays-- 1 Tray Of 3cc Syringes And 1 Tray Of 10cc Syringes, With Each Tray Containing 1- 25g, -28g And 2-30g Probes, With 4-6% Naoc1 Solution (f) Catalog # Maxe825, 3cc Syringe, 25g Probe, 2-3% Naoc1 Solution -5- (g) Catalog # Maxe828, 3cc Syringe, 28g Probe, 2-3% Naoc1 Solution (h) Catalog # Maxe830, 3cc Syringe, 30g Probe, 2-3% Naoc1 Solution | Manufacturer | Mpl Technologies, Inc | September 9, 1996 | The External Finger Grip Flanges May Separate From The Barrel Of The Syringe, Rendering The Syringe Inoperable | 12,455 syringes were distributed | Nationwide And Belgium
1996-10-22 | Devices | II | z-017-7 | Lifesaver Portable Oxygen Units, Catalog No | Manufacturer | Hudson Respiratory Care, Inc | June 14, 1996 | Device Was Distributed With A Regulator With A Flow Range Of 1-8 Lpm, Instead Of The Correct Regulator With A Flow Range Of 2-15 Lpm | 42 units were distributed | California, Texas, Illinois, Pennsylvania, Utah, Missouri, Massachusetts, Canada, Honduras
1996-10-22 | Devices | II | z-018-7 | The Thairapy Vest Airway Clearance System, Used To Provide Airway Clearance Therapy | Manufacturer | American Biosystems, Inc | February 8, 1996 | A Shorting In An Output Transistor Could Allow The Device Blower To Operate At Pressures Higher Than Specification, Resulting In Vest Pressures Constricting The Thorax And Impairing Inspiration | 173 devices were distributed | Nationwide
1996-10-22 | Devices | II |  | Low Porosity Woven Vascular Prosthesis, Catalog #olp1206, Indicated For Replacement Or Bypass Procedures In Aneurysmal And Occlusive Disease Of The Abdominal Aorta, Visceral Arteries And Proximal Peripheral Arteries Exclusive Of The Coronary Arteries | Manufacturer | Intervascular Inc | August 21, 1996 | One Thread Is Missing In The Involved Woven Vascular Prothesis Which May Potentially Increase Water Permeability Along The Line Of The Pulled Thread | 65 unimplanted devices were in commerce at time of recall initiation | Tennessee, New York, Louisiana, Illinois, Alabama, Florida, California, Pennsylvania, Arizona, Georgia, Texas, New Jersey, Ohio, Japan
1996-10-22 | Devices | II | z-020-024-7 | Input Ps Catheter Introducer Sheath With Hydro/pel Coating, Used To Facilitate Placing A Catheter Through The Skin Into A Vein Or Artery: (a) 5f 1000511 5f 11 Cm Input 2 (freeway) 050011 5f 11 Cm Input Introducer Intl | Usci Manufacturing Facility | C. r. Bard Irl. | August 13, 1996 | Introducer Hub Is Separating At Or Near The Weld Site Of The Upper And Lower Hub Body Parts | approximately 980,539 units were distributed | Nationwide And International
1996-10-22 | Devices | II | z-027/028-7 | Microvision-c Radiographic Film, Used For Mammography | Sterling Diagnostic Imaging | Sterling Diagnostic Imaging, Inc | August 30, 1996 | The Devices Have Demonstrated Uncharacteristic Aging, That Is Below Working Limits For Speed For This Product, Which May Result In The Mis-diagnosis In Mammography Readings | firm estimated that 19,388 boxes of film remained on market at time of recall initiation | Nationwide, Mexico, Brazil, Latin America
1996-10-22 | Devices | II | z-1216/1218-6 | (a) Mobilaire Iii Oxygen Concentrator, Model No | Manufacturer | Invacare Corporation, Sanford, Florida | March 8, 1996 | Kinked Oxygen Transfer Hoses In The Units Caused By Defective Exterior Foam Insulation May Affect The Flow Of Oxygen, And Therefore, Affect The Concentration Of Oxygen The Patient Receives | 193 units were distributed | Nationwide