## Report-1997-03-12
recall_classification_date | product_type | classification | recall_number | product | recalling_firm | manufacturer | recall_initiation_date | reason | volume | distribution
1997-03-12 | Biologics | III | b-433-7 | Platelets, Pheresis | Manufacturer | American Red Cross Blood Services, Baltimore, Maryland | July 22, 1996 | Blood Product Was Collected From A Donor With An Unacceptable Health History | 1 unit was distributed | District Of Columbia
1997-03-12 | Biologics | III | b-436-7 | Monoclate-p Antihemophilic Factor (human), Factor Viii:c Pasteurized, Monoclonal Antibody Purified | Centeon L. l. c. | Centeon L | February 19, 1997 | Reconstitution Difficulties In That Vials Failed To Transfer All Of The Diluent Into A Lyophilized Product | 1,908 vials were distributed; firm estimated that little if any product remained on market at time of recall initiation | Nationwide
1997-03-12 | Devices | II | z-297-7 | Sensi-touch Regional Anesthesia Delivery System, Spinal Anesthesia Tray, Catalog #8888- 828111, 22 Ga Plastic Hub, Diamond Point, Marcaine | Sherwood Davis & Geck | Kelsar S | January 31, 1997 | The Tray Incorrectly Identifies The Contents As Product/catalog No | 400 trays were distributed between 8/25/96 and 1/27/97 | Michigan, Missouri, Connecticut, Texas, New Jersey, Kentucky, Minnesota, Pennsylvania, Massachusetts
1997-03-12 | Devices | II | z-298/299-7 | Imatron Ct Scanner, A Component Of The Patient Table Potentiometer Assembly (y-potentiometer): (a) Imatron C100 Ct Scanner; (b) Imatron C150 Series Scanners | Manufacturer | Imatron, Inc | July 19, 1996 | The Table Would Go Beyond Its Limits And Could Not Be Driven Back Due To A Blown Fuse | 81 units were distributed | Nationwide And International
1997-03-12 | Devices | II | z-300-7 | Model Crc Laser Profilometer, Used To Determine The Surface Characteristics Of Paper | Manufacturer. Fda Approved The Firm's Corrective Action Plan On March 1 | International Paper Company, Tuxedo, New York | March 1, 1997 | Noncompliance With Performance Standards For Laser Products In That The Products Failed To Incorporate Required Safety Interlocks On The Operator Access Door | 2 units were distributed | Nationwide
1997-03-12 | Devices | II | z-303-7 | Eeg Tend Monitor, Model Oee-7102b, Intended For Long Term Eeg Monitoring And Analysis | Nihon Kohden America | Nihon Kohden Corporation, Shinjuku-ku, Tokyo, Japan | October 16, 1996 | When The 12v Dc Power Supply Conductor Is Broken, And The Main Power Is Turned On, A Higher Than Acceptable Current Leakage Is Applied To The Monitoring Electrodes And A Chemical Reaction Is Generated With Eeg Paste, Which Can Discolor The Patient's Skin When The Electrodes Are Placed On The Patient Skin | 23 units were distributed | Arizona, California, Illinois, Kansas, Louisiana, South Dakota, Texas, Utah
1997-03-12 | Devices | III | z-293-7 | Neurosign Straight Concentric Bipolar Probe With Touch Proof Connectors, Catalog #7033- 0919, An Accessory To The Neurosign 100 Motor Nerve Monitor | Smith & Nephew Ent | The Magstim Company Ltd | October 16, 1996 | The Device Is Subject To Possible Extrusion Of Silicone From The End Of The Probe When The Device Was Autoclaved By The Manufacturer | 8 units were distributed and are accounted for | Arizona, Kentucky, New York, Texas, Washington State, Wisconsin
1997-03-12 | Devices | III | z-294-7 | Incstar Vca Igg Clin-elisa Test Kit, Product #4590, An In-vitro Diagnostic Test Kit For Detecting Antibodies In Human Serum To Epstein-barr Viral Capsid Antigen (vca) | Manufacturer | Incstar Corporation, Stillwater, Minnesota | February 3, 1997 | Some Negative Samples May Yield A Positive Assay Response | 131 kits were distributed | California, New Jersey, New York, International
1997-03-12 | Devices | III | z-302-7 | Cytomegalovirus Igm Elisa Test Kit, Product #2325250, Used For The Qualitative Determination Of Igm Antibody To Cytomegalovirus In Human Serum | Manufacturer | Clark Laboratories, Inc | January 23, 1997 | The Absorbance Values For Controls And Specimens Are Lower Than Expected | 495 kits were distributed | Nationwide And International
1997-03-12 | Devices | III | z-304-7 | C-150xp Ct Scanner Installed With Software Version 12 | Manufacturer | Imatron, Inc | August 28, 1996 | The Device's Software Version 12 | 30 units were distributed | Nationwide And International
1997-03-12 | Devices | III | z-305/307-7 | Software Versions 11 | Manufacturer | Imatron, Inc | November 20, 1995 | During Reconstruction Of Images, The Fast Reconstruction System Software (fsr) May Produce An Error Condition Which The Computer May Interpret As A Warning, Causing The Computer To Display Erroneous Data For The Image | 37 units were distributed | Nationwide And International
1997-03-12 | Devices | III | z-308-7 | Nellcor Symphony N-3000 Pulse Oximeter, Model N-3000 | Nellcor Puritan Bennett | Nellcor Puritan Bennett, Inc | October 23, 1996 | The Device May Be Potentially Subject To Printed Circuit Board Damage Which Can Result In The Monitor Reporting Sp02 Values Slightly Outside The Normal Accuracy Range Stated In The Operations Manual | 111 units were distributed | Nationwide And International
1997-03-12 | Devices | III |  | Sherpa Vector And Sherpa Vector X Coronary Guiding Catheters, Designed To Provide A Pathway Through Which Therapeutic Devices Are Introduced: (a) Vector 6f, Catalog Numbers Pa5xx; (b) Vector 7f, Catalog Numbers As7xxx; (c) Vector 8f, Catalog Numbers As8xxx; (d) Vector X 6f, Catalog Numbers S26xxx; (e) Vector X 7f, Catalog Numbers S27xxx; (f) Vector X 8f, Catalog Numbers S28xxx |  | Medtronic Interventional Vascular Inc | N/A | Resistance May Be Experienced While Passing Interventional Devices Through The Lumen Of The Guiding Catheter, Which May Result In The Inability To Move The Interventional Device Within The Lumen Of The Guiding Catheter | approximately 38,796 units were distributed | Nationwide