## Report-1997-12-03
recall_classification_date | product_type | classification | recall_number | product | recalling_firm | manufacturer | recall_initiation_date | reason | volume | distribution
1997-12-03 | Biologics | III | b-143-8 | Source Plasma | Manufacturer | Decatur Plasma, Inc | June 13, 1997 | Blood Products Were Collected From Donors Who Were Not Tested For Syphilis | 3 units were distributed | North Carolina
1997-12-03 | Biologics | III | b-146-8 | Fresh Frozen Plasma | Manufacturer | United Blood Services, Chicago, Illinois | August 22, 1997 | Blood Product Was Prepared From A Unit Of Whole Blood Identified As A Difficult Draw | 1 unit was distributed | Illinois
1997-12-03 | Biologics | III | b-147-8 | Fresh Frozen Plasma | Manufacturer | United Blood Services, Chicago, Illinois | August 26, 1997 | Blood Product Was Prepared From A Unit Of Whole Blood Identified As A Difficult Draw | 1 unit was distributed | Illinois
1997-12-03 | Biologics | III | b-150-8 | Fresh Frozen Plasma | Manufacturer | United Blood Services, Chicago, Illinois | September 2, 1997 | Blood Product Was Prepared From A Unit Of Whole Blood Identified As A Difficult Draw | 1 unit was distributed | Illinois
1997-12-03 | Biologics | III | b-151-8 | Cryo-poor Plasma | Manufacturer | United Blood Services, Chicago, Illinois | September 5, 1997 | Blood Product Was Prepared From A Unit Of Whole Blood Identified As A Difficult Draw | 1 unit was distributed | Illinois
1997-12-03 | Biologics | III | b-182-8 | Recovered Plasma | Manufacturer | American Red Cross Blood Services, Miami, Florida | January 7, 1997 | Blood Product Corresponded To A Unit Of Clotted Red Blood Cells | 1 unit was distributed | Switzerland
1997-12-03 | Biologics | III | b-202/206-8 | A) Red Blood Cells; B) Platelets; C) Fresh Frozen Plasma; D) Cryoprecipitated Ahf; E) Recovered Plasma | Blood Systems | United Blood Services, Scottsdale, Arizona | March 12, 1996 | Blood Products Were Collected From Donors Whose Body Temperatures At The Time Of Donation May Have Been Greater Than 99 | a) 21 units; b) 11 units; c) 3 units; d) 3 units; e) 8 units were distributed | Arizona, Mississippi, Switzerland
1997-12-03 | Biologics | III | b-209-8 | Platelets, Pheresis | Manufacturer | American Red Cross Blood Services, Dedham, Massachusetts | February 12, 1997 | Blood Product For Which The Date And Time Of Receipt Was Not Documented | 1 unit was distributed | New Hampshire
1997-12-03 | Biologics | III | b-213/214-8 | A) Red Blood Cells; B) Platelets | Manufacturer | Florida Blood Services, Inc | April 5, 1996 | Blood Products Were Collected From A Donor With An Elevated Body Temperature | 1 unit of each component was distributed | Florida
1997-12-03 | Biologics | III | b-219/221-8 | A) Red Blood Cells; B) Platelets; C) Recovered Plasma | Blood Systems | United Blood Services, Scottsdale, Arizona | May 5, 1997 | Blood Products Had An Acceptable Alt But Were Collected From A Donor Who Previously Had An Elevated Alt On Two Separate Occasions | 1 unit of each component was distributed | Arizona And Switzerland
1997-12-03 | Biologics | III | b-229-8 | Red Blood Cells | Manufacturer | American Red Cross, Atlanta, Georgia | September 2, 1997 | Blood Product Was Collected From A Donor Who Reported Cold And Flu Symptoms After Donation | 1 unit was distributed | Florida
1997-12-03 | Biologics | III | b-242-8 | Fresh Frozen Plasma | Manufacturer | American Red Cross Blood Services, Miami, Florida | February 3, 1997 | Blood Products Were Not Frozen Within 8 Hours Of Whole Blood Collection | 2 units were distributed | Florida
1997-12-03 | Biologics | III | b-271-8 | Red Blood Cells | Blood Systems | United Blood Services, Mcallen, Texas | September 17, 1997 | Blood Product May Have Remained At Room Temperature For More Than 30 Minutes | 1 unit was distributed | Texas
1997-12-03 | Biologics | III | b-272-8 | Whole Blood | Manufacturer | American Red Cross Blood Services, Portland, Oregon | August 21, 1997 | Blood Product May Have Remained At Room Temperature More Than 30 Minutes | 1 unit was distributed | Oregon
1997-12-03 | Biologics | III | b-278/279-8 | A) Red Blood Cells; B) Fresh Frozen Plasma | Manufacturer | American Red Cross Blood Services, Miami, Florida | June 18, 1997 | Blood Products Were Collected From A Donor Who Took Antibiotics Within 48 Hours Of Donation | 1 unit of each component was distributed | Florida
1997-12-03 | Devices | II | z-198-8 | Pediatric Flow Meter/humidifier Holder, Model #pv5lp-627 | Manufacturer | Sunrise Medical Hhg, Inc | October 23, 1997 | The Labeling Failed To Warn Users That The Device Is Intended To Be Directly Attached To A Devilbiss Model Pv5l Oxygen Concentrator Only, And That When Using With Other Models, The Device Should Be Used Only As A Remote Holder And Must Be Securely Fastened To A Sturdy Horizontal Surface | 424 units were distributed | Nationwide, Australia, Canada
1997-12-03 | Devices | II | z-201/202-8 code all lot codes/serial numbers | Fresenius Hemodialysis Delivery System With Software Version 8 | Manufacturer | Fresenius Medical Care, Walnut Creek, California | November 5, 1997 | In The Unlikely Event Of A Sudden Massive Blood Leak, The System May Not Give An Alarm | 21,083 units were distributed | Nationwide, Canada, Mexico, Caribbean
1997-12-03 | Devices | II | z-204-8 | Ruby Laser System, Model #5000, Used In Dermatology | Manufacturer. Fda Approved The Firm's Corrective Action Plan On November 26 | Sharplan Lasers, Inc | November 26, 1997 | The Operator's Manual Lacked Calibration Procedures | 121 units were distributed | Nationwide And International
1997-12-03 | Devices | III | z-199-8 | Fisher Diagnostics Brand Leuko-stat Stain Kit, Catalog #cs430d, In-vitro Diagnostic | Manufacturer | Richard-allan Scientific, Kalamazoo, Michigan | August 18, 1997 | One Of The Three Kit Components Was Mislabeled So That Two Of The Components Bear The Same Label | 139 kits were distributed | Nationwide
1997-12-03 | Devices | III | z-200-8 | Innofluor Amikacin System Calibrators, Catalog No | Manufacturer | Oxis International, Inc | October 2, 1997 | The Device May Not Perform According To Specifications During The Labeled Shelf Life | 50 sets were distributed | Belgium, France, Germany, Greece, Spain, And The United Kingdom