## Report-1998-7-15
recall_classification_date | product_type | classification | recall_number | product | recalling_firm | manufacturer | recall_initiation_date | reason | volume | distribution
1998-7-15 | Drugs | III | d-182-8 | oxygen, Medical, Usp, Compressed, In High Pressure E Cylinders, rx | manufacturer | sieps Homecare, Inc | May 27, 1998 | subpotent And Good Manufacturing Practice Deviations. | 2 cylinders were distributed | minnesota
1998-7-15 | Drugs | II | d-172-8 | compressed Oxygen Usp, Transfilled In E And M6 Size Aluminum cylinders | manufacturer | shiv Respiratory Equipment, Inc | May 17, 1998 | subpotency. | 218 e-size cylinders; 125 m6-size cylinders were distributed | new York
1998-7-15 | Devices | II | z-642-8 | ohmeda Tec 6 Desflurane Vaporizers (note: Excludes North American drager (nad) And Dragerwerk Tec 6 Variant Vaporizers) | manufacturer | datex-ohmeda (formerly Ohmeda Medical Systems Division), Madison, wisconsin | June 22, 1998 | gas Leak | 6,207 vaporizers were distributed | nationwide And International
1998-7-15 | Devices | II | z-643/644-8 code lot numbers: a) 96 g 398; b) 96 g 348 | a) 3m Modular Shoulder System, 14 Mm X 125 Mm Stems, Product number 17464; B) 3m Modular Should System, 16 Mm X 125 Mm Stems, product Number 17466, A Metal Stem Designed To Fit Into A surgically-prepared Cavity In A Humerus Bone During Shoulder joint Replacement Surgery | 3m Health Care | 3m Health Care, Rotherham, England | June 3, 1998 | one Stem With A 16 Mm Diameter Was In A Box Labeled As 14 Mm diameter | a) 42 stems; b) 31 stems were distributed | nationwide And International
1998-7-15 | Devices | II | z-645-8 | medex Injection Adapter, With Male Leur Lock, Intended For Use In the Administration Of Small Volume Fluids, Using Syringes | medex | medex, Inc | June 26, 1998 | the Sterility Of The Device Cannot Be Assured Due To Defective packaging | 390 cases (100 units per case) were distributed | arkansas, Illinois, New Jersey, Pennsylvania, New York, Montana, connecticut, California, Maryland, Missouri, North Carolina, tennessee, Texas, Virginia, Iowa, Oregon, Minnesota
1998-7-15 | Devices | II | z-646/647-8 | newport Infant/adult "breeze" Ventilator, Designed For respiratory Support Of Infants, Pediatrics, Or Adults: a) Model E200d (domestic Sales) Ventilator; b) Model E200e (foreign Sales) Ventilators | newport Medical Instruments | mer-mar, Inc | March 5, 1998 | a Bad Ic (chip) Causes Inadvertent System Reset, Continually during Power-up, Or Randomly During Operation, Wherein (1) The ventilators Cease Operation (gas Flow Ceases And The Exhalation valve Opens) For About 4 Seconds, Then Restarts At Values That were Set Before The Reset (except As Follows), (2) The Sigh function Defaults To Off, (3) The Nebulizer Defaults To Off, (4) the Inspiratory Minute Volume Alarm Defaults To The 0-50 L | 61 ventilators were distributed | nationwide And International
1998-7-15 | Devices | II | z-649/650-8 | victoreen Veridose Diodes, Designed For Use As Radiation detection Devices Encased Within A Polystyrene Material: a) Model 30-472 Positive Polarity Veridose Diodes; b) Model 30-472-8000 Negative Polarity Veridose Diodes | manufacturer | victoreen, Inc | May 7, 1998 | the Devices Were Manufactured With An Inappropriate Shielding material (build-up Cap) | 23 diodes were distributed | arizona, Kentucky, New York, Pennsylvania, Florida, Virginia, maryland, Michigan, Wisconsin, Ohio, Tennessee, Texas, Alabama, georgia, Mississippi
1998-7-15 | Devices | II | z-667/668-8 | tracheal Intubation Fiberscope Instruction Manual: a) Catalog No | olympus American | olympus Optical Company, Ltd | June 3, 1998 | the Cover Pages For The Instruction Manuals Were For The Correct intubation Fiberscope, But The Rest Of The Instructions Were For the Other One | a) 27 units; b) 287 units were distributed were distributed | nationwide And International
1998-7-15 | Devices | II | z-669-8 | model Tru-pulse C02 Laser System, Used In General Surgery | manufacturer. Fda Approved The Firm's Corrective Action Plan On july 1 | tissue Medical Lasers, Inc | July 1, 1998 | noncompliance With Performance Standards For Laser Products In that The Operators Manual Lacked Calibration Procedures | 275 units were distributed | nationwide And International
1998-7-15 | Devices | II | z-670-8 | i-125 Seeds, (iodine-125) Therapeutic For Interstitial brachytherapy, Model 6711, Used As Radioactive Implants For The treatment Of Prostate Cancer | manufacturer | medi-physics, Inc | September 23, 1997 | the Seeds Were Mislabeled With An Incorrect Apparent radioactivity Range | 26 seeds were distributed | georgia And Florida
1998-7-15 | Devices | II | z-671-8 | i-125 Seeds, (iodine-125) Therapeutic For Interstitial brachytherapy, Model 6711, Used As Radioactive Implants For The treatment Of Prostate Cancer | manufacturer | medi-physics, Inc | January 8, 1998 | the Seeds Were Mislabeled With An Incorrect Apparent radioactivity Range | 161 seeds were distributed | washington State, California, New York
1998-7-15 | Devices | II | z-672-8 | disposable Manual Resuscitator, Part #dmr2, Used To Deliver supplemental Oxygen | nellcor Puritan Bennett | nellcor Mexico, Tijuana, Mexico | May 22, 1998 | the Duckbill Valve Can Become Unseated, Allowing Rebreathing and/or Insufficient Inspiratory Air Pressure; And (2) The pressure Relief Valve Ball Can Rust, Causing It To Become inoperative | 67,490 units were distributed | nationwide And Italy