## Report-1999-11-17
recall_classification_date | product_type | classification | recall_number | product | recalling_firm | manufacturer | recall_initiation_date | reason | volume | distribution
1999-11-17 | Devices | II | z-130-0 | spline Ha-coated Cylinder Dental Implants (endosseous Implants), part #1852 | manufacturer | sulzer Calcitek Inc | May 3, 1999 | 13mm And 15mm Implants Were Interchanged, Each Size In Trays intended (labeled) For The Other | 1,010 implants were distributed | nationwide And Internationl
1999-11-17 | Devices | II | z-135/139-0 | a) Percutaneous Transluminal Angioplasty (pta) Catheter, Catalog numbers: Pta-os253, 303 And 403 b) Z-med Dilation Catheter, Catalog #pdz-s c) Z-med Ii Dilatation Catheter, Catalog Numbers Pds-630, 643 And 655 d) Pta-os Balloon Catheter, Catalog Numbers: Vsb-030 And 040 e) Ghost Ii Dilation Catheter, Catalog Numbers Gii-109, 113, 121, 129, 130, 137, 138, 141 | manufacturer | numed, Inc | January 29, 1999 | non-sterilit | 302 catheters were distributed | massachusetts, United Kingdom, South Africa
1999-11-17 | Devices | II | z-154/159-0 | immulite Anti-thyroglobulin Antibody And Immulite Anti-thyroid peroxidase Antibody, For In-vitro Diagnostic Use: a) Immulite Anti-thyroglobulin Antibody - 50 Tests; b) Immulite Anti-thyroglobulin Antibody - 100 Tests; c) Immulite Anti-thyroglobulin Antibody - 500 Tests; d) Immulite Anti-thyroid Peroxidase Antibody - 50 Tests; e) Immulite Anti-thyroid Peroxidase Antibody - 100 Tests; f) Immulite Anti-thyroid Peroxidase Antibody - 500 Tests | diagnostic Products Corporation | euro/dpc Ltd | July 15, 1999 | the Sample Diluent Is Non-homogeneous And May Produce Inaccurate test Results | 1,269 kits were distributed | nationwide And United Kingdom
1999-11-17 | Devices | II | z-160/169-0 | port-a-cath Implantable Access Systems In Kits And Trays | manufacturer | sims Deltec, Inc | October 18, 1999 | fracture Of Device Catheter Can Result In Leaks | 17,485 implantable access systems were distributed reason fracture of device catheter can result in leak | nationwide And International
1999-11-17 | Devices | II | z-170/171-0 | revision J Software In The Nellcor Puritan Bennett, 740 ventilator System And 760 Ventilator System | mallinckrodt | nellcor Puritan Bennett Ireland, Ltd | September 13, 1999 | software Can Cause Under Or Over-delivery Of Oxygen | 223 units | nationwide And International
1999-11-17 | Devices | II | z-172-0 | medi-vac Non-conductive Suction Tube With Maxi-grip Connectors, A sterile Non-conductive Tubing Used With A Vacuum-powered Suction apparatus | allegiance Healthcare Corporation | allegiance Healthcare Corporation, Mexicali-baja Cal, Mexico | October 28, 1999 | the Outer Package Seal May Not Hold, Thus The Sterility Of The product May Be Compromised | 8,340 units were distributed | nationwide And International
1999-11-17 | Devices | II | z-176/189-0 | alm Prismalix (prx) Ceiling Mounted Surgical Lighting System, used To Provide Visible Illumination Of Surgical patients And Other Patient Examination Procedures: a) Prx 6401 Series; b) Prx 8401 Series; c) Prx 6601 Series; d) Prx 8601 Series; e) Prx 4401 Series; f) Prx 6441 Series; g) Prx 8441 Series; h) Prx 4441 Series; i) Prx 6001 Series; j) Prx 8001; k) Prx 4001 Series; l) Prx 4401 Series; m) Prx 4441 Series; n) Prx 6601 | alm Surgical Equipment | alm Sa, Parc De Limere, Ardon Orleans Cedex | September 17, 1999 | the Main Arm Of The Device May Fall From The Shaft Because Of missing Shaft-retaining Clips | 491 devices were distributed | nationwide, Guam, Canada
1999-11-17 | Devices | II | z-190/198-0 | sri Brand Of Surgical Procedure Packs, An Assembly Of Finished medical Devices And Other Items Used During The Preparation and/or Performance Of Surgical/medical Procedures: a) Pack No | manufacturer | sterile Recoveries, Inc | September 15, 1999 | the Devices Were Labeled And Distributed As Sterile Devices without Assurance That They Were Sterile | 410 units were distributed | california
1999-11-17 | Devices | II | z-199-0 | tyshak Brand Peripheral Balloon Dilation Catheter, Model 610532, rated Burst Pressure: 3 | manufacturer | numed, Inc | August 26, 1999 | the Firm Inadvertently Attached A Tyshak Ii Balloon To A Tyshak I catheter | 7 catheters were distributed | pennsylvania
1999-11-17 | Devices | II |  | oxy-peep Oxygen Diluter Kit, Model P-63000, Dilutes Oxygen Being administered To A Patient | manufacturer | pegasus Research Corporation, Santa Ana, California | October 11, 1999 | the Exhalation Valve In The Mask May Stick | 516 units were distributed | nationwide
1999-11-17 | Devices | II | z-201/204-0 | blood Collection Safety Systems As Follows: a) Safe Point M-d, Multi-draw Needle Cover System, 21 Gauge X 1 1/2" Needle, Packaged 150 Or 200 Pieces Per Box, Or Individually packaged | north American Medical Products | north American Medical Products, Inc | September 23, 1999 | lack Of Assurance Of Sterility | 28,900 catheters (bulk lots) | nationwide
1999-11-17 | Devices | II | z-205-0 | proclear Compatibles Daily Wear Contact Lens | manufacturer | biocompatibles Eyecare, Inc | November 3, 1999 | leaking Blister Seal Of Contact Lens Packaging Compromised sterility | 119,028 lenses were distributed in the united states | nationwide And International