## Report-2000-05-03
recall_classification_date | product_type | classification | recall_number | product | recalling_firm | manufacturer | recall_initiation_date | reason | volume | distribution
2000-05-03 | Biologics | II |  | A) Red Blood Cells; B) Fresh Frozen Plasma; C) Recovered Plasma | Manufacturer | , By Letter Dated January 21 Or 24, 2000 | January 21, 2000 | Blood Products Were Collected From A Donor With A History Of Having Tested Positive For Hepatitis | a) 2 units; b) 1 unit; c) 1 unit | Illinois And Switzerland
2000-05-03 | Biologics | II | b-580/581-0 | A) Red Blood Cells; B) Recovered Plasma | Manufacturer | , By Letters Dated January 6, 2000 Or December 12, 1999 | January 6, 2000 | Blood Products Were Collected From An Unsuitable Donor | 1 unit of each component was distributed | Illinois And Switzerland
2000-05-03 | Biologics | II | b-582-0 | Platelets | Manufacturer | , By Letter Dated January 11, 2000 | January 11, 2000 | Blood Product Was Collected From A Donor With An Elevated Alt | 1 unit was distributed | Illinois
2000-05-03 | Biologics | II | b-584-0 | Platelets Pheresis, Leukocytes Reduced | Manufacturer | , By Letter Dated December 27, 1999 | December 27, 1999 | Blood Product Had A High White Cell Count | 1 unit was distributed | Illinois
2000-05-03 | Biologics | II |  | A) Red Blood Cells, Leukocytes Reduced; B) Cryoprecipitated Ahf | Manufacturer | , By Telephone On February 22, 2000 | February 22, 2000 | Blood Products Were Collected From A Donor Taking The Drug Famciclovir | 1 unit of each component was distributed | Illinois
2000-05-03 | Biologics | II | b-591-0 | Red Blood Cells, Leukoreduced | Manufacturer | , By Telephone On January 31, 2000, And By Letter Dated February 4, 2000 | January 31, 2000 | Blood Product Was Collected From A Donor Who Reported Travel To An Area Designated As Endemic For Malaria | 1 unit was distributed | North Carolina
2000-05-03 | Biologics | II | b-593/594-0 | A) Red Blood Cells; B) Recovered Plasma | Manufacturer | , By Fax On January 26, 2000, And March 15, 2000 | January 26, 2000 | Blood Products Tested Negative For The Antibody To The Human T- Lymphotropic Virus Type I And Ii (anti-htlv-i/ii, But Was Collected From A Donor Who Previously Tested Reactive For Anti-htlv I/ii | 1 unit of each component was distributed | Illinois And Switzerland
2000-05-03 | Biologics | II | b-595/597-0 | A) Red Blood Cells; B) Platelets For Manufacturing; C) Recovered Plasma | Manufacturer | , By Letter Dated February 4, 2000 | February 4, 2000 | Blood Products Were Collected From An Unsuitable Donor | 1 unit of each component was distributed | Ohio, Tennessee, New York
2000-05-03 | Biologics | II | b-598-0 | Red Blood Cells | Manufacturer | , By Letter Dated May 14, 1999 | May 14, 1999 | Blood Products Were Exposed To Unacceptable Shipping And Storage Temperatures | 5 units were distributed | New York
2000-05-03 | Biologics | II | b-599/600-0 | A) Red Blood Cells; B) Fresh Frozen Plasma | Manufacturer | , By Letter Dated November 29, 1999 | November 29, 1999 | Blood Products Were Collected From A Donor Taking The Drug Proscar | 1 unit of each component was distributed | Ohio And Florida
2000-05-03 | Biologics | II | b-602-0 | Source Plasma | Manufacturer | , By Fax Dated April 29, 1999, And By A Corrected Fax Dated May 5, 1999 | April 29, 1999 | Blood Products Were Collected From A Donor Who Had Been Incarcerated | 3 units were distributed | Illinois
2000-05-03 | Biologics | III | b-505/507-0 | A) Red Blood Cells; B) Red Blood Cells, Leukocyte Reduced; C) Fresh Frozen Plasma | Manufacturer | , By Letter Dated May 19, 1999 | May 19, 1999 | Blood Products Corresponded To A Unit Of Pooled Platelets Which Was Positive For Propionbacterium Species | a) 8 units; b) 1 unit; c) 3 units were distributed | Ohio
2000-05-03 | Biologics | III | b-572/573-0 | A) Red Blood Cells; B) Platelets | Blood Systems |  | January 14, 2000 | Blood Products Were Collected From A Donor Taking The Antibiotic Cipro | 1 unit of each component was distributed | Alabama And Mississippi
2000-05-03 | Biologics | III | b-574/576-0 | A) Red Blood Cells; B) Platelets; C) Recovered Plasma | Manufacturer | , By Letter Dated July 2, 1999 | July 2, 1999 | Blood Products Tested Reactive For Syphilis | 1 unit of each component was distributed | Minnesota And Switzerland
2000-05-03 | Biologics | III | b-583-0 | Red Blood Cells, Leukocytes Reduced | Manufacturer | , By Letter Dated December 17, 1999 | December 17, 1999 | Red Blood Cells Were Leukoreduced By Filtration Greater Than Five Days Past The Time Of Collection | 2 units were distributed | Illinois
2000-05-03 | Devices | II | z-575-0 | Pro Med Kits Assembles First Aid Kits On Key Chains Each Containing A Clinipad Alcohol Prep, Antiseptic | Manufacturer | , By Letter On March 20, 2000, Followed By Telephone | March 20, 2000 | Kits Contain Clinipad Products Labeled As Sterile For Which Clinipad Is Unable To Assure The Sterility | 2,500 kits | California
2000-05-03 | Devices | II | z-577-0 | Peg Kits, T-j Tube Kits, Epidural Trays, Spinal Procedure, Universal Block Trays, Paracentesis Trays Containing Recalled Clinipad Products: Models 0150-14, 0150-20, 0150-24, 0160-14, 0160-20, 0160-24, 0260-16, 0260-18, 0260-22, 182125, 182131, 182184, 182224, 182229, 182230, 18236, 18259, 18283, 18289, 18668, 18178, 18184, 18202, 60432, 60433 | Manufacturer | , By Letter On April 7, 2000, Instructing Customers To Remove And Destroy The Clinipad Components Or To Return The Kit To Ballard | April 7, 2000 | Kits Contain Clinipad Products Labeled As Sterile For Which Clinipad Is Unable To Assure The Sterility | 41,231 kits/trays | Nationwide And Puerto Rico
2000-05-03 | Devices | II | z-578-0 | Procedure Kits Containing Recalled Clinipad Products: A) Leep Redikit - Product Number: 6060 B) Hysteroscopy Kit -product Number: 6160 C) Cervical Block Kit- Product Number: 6161 | Manufacturer | , By Letter The Week Of April 14, 2000, And Providing Stickers To Be Placed On Kits Advising Users Not To Use The Alcohol Prep Pads | April 14, 2000 | Kits Contain Clinipad Products Labeled As Sterile For Which Clinipad Is Unable To Assure The Sterility | 32,145 kits | Nationwide
2000-05-03 | Devices | II | z-579/582-0 | T. E. D. Sequential Compression Sleeves: A) Model No. 5329 Scd Blue Sleeves | The Kendall Co. Lp | The Kendall Co | February 24, 2000 | The Compression Sleeves May Have Occluded Airway Tubes | 6,292 cases were distributed | Nationwide And International
2000-05-03 | Devices | II | z-583/589-0 | Dual Stage Venous Return Cannula With Bullet Tip, Intended For Use As A Single Tube Venous Drainage From The Right Atrium And Vena Cava During Extracorporeal Circulation For Open-heart Surgery: A) Model No | Chase Medical |  | February 29, 2000 | The Cannula May Separate Between The Tip And Body During Use | 1,787 units were distributed | Pennsylvania, Texas, Nebraska, Wisconsin, Virginia, Louisiana, Massachusetts, Washington State, Virginia, New Jersey, Connecticut, And International
2000-05-03 | Devices | II | z-590/601-0 | 24french Aortic Arch Cannula, Indicated For Use For Perfusion Of The Ascending Aorta During Surgical Procedures Requiring Cardiopulmonary Bypass: A) Model No | Chase Medical |  | February 29, 2000 | The Cannula May Leak Under Certain Flow And Device Manipulation Conditions | 3,860 units were distributed | Texas, California, Florida, Minnesota, Alabama, Massachusetts, Mississippi, Georgia, Virginia, New Jersey, Connecticut, And International
2000-05-03 | Devices | II |  | Bayer Setpoint Unconjugated Estriol Reagent (ue3) For Bayer Immuno 1 System, An Assay To Measure Unconjugated Estriol In Human Serum | Bayer Corporatiaon |  | February 25, 2000 | The Device Is Associated With A Failure To Calibrate The Assay On The System Rss Limits, Erratic Quality Control Results, And A Calibration Interval Of Less Than Seven Days | 330 units were distributed | Nationwide And International
2000-05-03 | Devices | II | z-611-0 | Architect I System With Software Version 1 | Manufacturer | , By Telephone On February 25, 2000, And By Letters Dated February 24, 1999 | February 25, 2000 | Software Errors Can Potentially Result In Printed Patient Report To Incorrectly Associate Tests Results With Patients Ids Under Certain Conditions | 22 units were distributed | California, Florida, Illinois, Massachusetts, Michigan, Missouri, North Carolina, New Hampshire, New York, Ohio, Pennsylvania, Texas, Utah, Virginia, International
2000-05-03 | Devices | II | z-612-0 | Roche/hitachi 917 Disk System Chemistry Analyzer, Catalog #917000 | Roche Diagnostics Corporation |  | March 17, 2000 | There Is A Potential For Mismatch Of Sample Identification Information And Results On All Versions Of Software | 500 units were distributed | Nationwide
2000-05-03 | Devices | II | z-613-0 | Roche Hitachi 914 Chemistry Analyzer, Used For Diagnostic Testing For Routine And Specialty Chemistry Analytes | Manufacturer | , By Letter Dated March 17, 2000 | March 17, 2000 | There Is A Potential For Mismatch Of Sample Identification Information And Results On This Version Of Software | 56 units were distributed reason there is a potential for mismatch of sample identification information and results on this version of softwar | Alabama, Arizona, California, Colorado, Florida, Georgia, Hawaii, Indiana, Kansas, Michigan, Missouri, Nebraska, New Jersey, New York, Ohio, Oklahoma, Pennsylvania, Puerto Rico And Texas
2000-05-03 | Devices | II | z-616/621-0 | Lcs Total Knee System - Ps Tibial Insert: A) Small Size, Catalog Nos | Manufacturer | , By Letter Dated December 20, 1999 | December 20, 1999 | The Devices Can Cause The Escape Of Fluid From Blood Vessels Into The Tissues Or A Cavity And Cause Knee Pain | 236 units are subject to recall | Nationwide
2000-05-03 | Devices | II | z-622/627-0 | Sarns Mp-6 Integrated Cardioiplegia Sets, Used For Delivery Of Temperature Controlled Cardioplegia Solutions To The Heart During Cardiopulmonary Bypass Surgery: A) Mp-6 Integrated Cardioplegia Set, Catalog No | Manufacturer | , By Letters On March 31, And April 3, 2000, And By Telephone | March 31, 2000 | Holes In The Mylar Side Of The Passdown Line Pouch Are Inadequate Which Results In Compromising The Sterility Of The Device | 708 units were distributed | Nationwide And Germany
2000-05-03 | Devices | III | z-532/533-0 | A) Ankle Contracture Boot And Boot Sole Accessory, Model Numbers: 4300, 4301, 4302, 4305, 4306, And R4307; B) Boot Sole Accessory #4300/05wk | Manufacturer | , By Letters Dated September 23, 1999 And February 14, 2000 | September 23, 1999 | The Device Fails To Bear A Caution Label That It Is Not Intended For Weight Bearing Purposes, And As A Result, The Bolt On The Bottom Of The Boot Holding The Boot Sole In Place, Cracks And Breaks, Causing The Sole To Dislodge | 22,538 units were distributed | Nationwide
2000-05-03 | Devices | III | z-570/571-0 | Model 250 Neurocybernetic Prostheses (ncp) Programming Software Versions A) 4/1; B) 4 | Manufacturer | , By Telephone And Letter On August October 6, 1999 | August 6, 1999 | The Label Insert Sheets Were Printed With The Incorrect Lot Number | a) 1 unit of lot number 23952c; b) 10 units of lot number, 26023c, 10 units of lot 27137c and 50 units of lot 27735c were distributed | A) Canada; B) Belgium And Europe
2000-05-03 | Devices | III | z-576-0 | Immunocard H | Manufacturer | , By Certified Mail, Fax And Telephone On March 28, 2000 | March 28, 2000 | Variations Of The Positive Control And Test Ports May Lead To Invalid Results | 546 kits were distributed | Nationwide, Canada, Japan, Italy, Germany
2000-05-03 | Devices | III | z-606/608-0 | Boehringer Mannheim Systems Creatinine Plus, For The Quantitative Determination Of Creatinine In Serum, Plasma, And Urine: A) Catalog No | Roche Diagnostics Corporation |  | February 23, 2000 | These Lots Exhibit Positive Interference From Lactate | approximately 1,500 kits were distributed | Nationwide
2000-05-03 | Devices | III | z-615-0 | Carolon Trinity Pneumatic Pump, Model 480000 | Carolon Company |  | December 19, 1999 | A Plug Portal Area With A Hairline Crack And Recession Of Receptor Prong Into Casing Could Result In A Potential Fire Hazard Within The Cord | 499 pumps were distributed | Nationwide Quantity 499 Pumps Were Distributed