## Report-1995-01-24
recall_classification_date | product_type | classification | recall_number | product | recalling_firm | manufacturer | recall_initiation_date | reason | volume | distribution
1995-01-24 | Drugs | II | d-077-6 | Myochrysine (gold Sodium Thiomalate) Injection 50 Mg/ml, In 10 Ml Multi-dose Vials, Used For Rheumatoid Arthritis, Under The Merck And King Pharmaceuticals Label | Manufacturer | Merck And Company, Inc | December 18, 1995 | Product Does Not Meet Antimicrobial Effectiveness Test. | approximately 90,000 unit packages remained on market at time of recall initiation | Nationwide, Spain, Peru
1995-01-24 | Drugs | III | d-078-6 | Sodium Iodide I-131 Solution For Oral Therapeutic Use, In 100 Mci And 50 Mci Vials, Used For The Treatment Of Hyperthyroidism And Selected Cases Of Carcinoma Of The Thyroid | Cis-us | Cis Bio International, Subsidiary Of Compagnie Oris Industrie Sa, Cedex, France | December 15, 1995 | Some Packages Contained A Package Insert For Sodium Iodide I-131 Capsules. | 71 100-mci vials and 55 50-mci vials were distributed | Nationwide
1995-01-24 | Devices | II | z-321/325-6 | I. v. Extension Sets With "t" Connectors: (a) Abbott Extension Set With "t" Connector | Abbott Laboratories | Abbott Laboratories, Laurinburg, North Carolina (hg Lot Suffix); Abbott Health Products, Barceloneta, Puerto Rico (h1 Lot Suffix) | December 8, 1995 | Some Of The Extension Sets Leak And Disconnect Between The Male Fitment Of The Device And The Luer Fitment Of The I | 2,120,180 units were distributed; firm estimated that 135,000 - 150,000 units remained on market at time of recall initiation | Nationwide And International
1995-01-24 | Devices | II | z-332/334-6 | System 97, A Helium Charged, Mobile Intra- Aortic Balloon Pump (iabp), Catalog #0998-00- 0104-x, For Use As A Patient Aid During Pre-, Intra-, Or Post- Operative Open Heart Surgery; Use Within Patients Demonstrating Unstable Angina; Use Within Patients With Left Main Artery Occlusion Or Poor Left Ventricle Function: (a) Power Supply; (b) Front End Board; (c) Solenoid Driver Board | Manufacturer | Datascope Corporation, Paramus, New Jersey | August 31, 1995 | The System May Fail Because Of The Following Component Failures: Power Supply: Failure Of A Zener Diode, Designated At Vr9 In The System Power Supply, Which Can Result In A "power On" Condition Which Cannot Be Switched Off By Normal Means | (a) 259 units; (b) 554 units; (c) 153 units were distributed; reason the system may fail because of the following component failures: power supply: failure of a zener diode, designated at vr9 in the system power supply, which can result in a "power on" condition which cannot be switched off by normal mean | Nationwide And International
1995-01-24 | Devices | II | z-335-6 | 3m Sarns Brand Perfusion System 9000 Gas Flow System, Used To Provide And Monitor Co2 Flow To The Patient During Cardiopulmonary Bypass | Manufacturer | Sarns 3m Health Care, Ann Arbor, Michigan | July 24, 1995 | The Relay Controlling The Co2 Flow Can Experience Random Failure And Latch In The "on" Position, Causing Co2 Gas To Continue To Flow After Completion Of The Pre-bypass Co2 Flush Of The Perfusion Circuit | approximately 500 units | Nationwide And International
1995-01-24 | Devices | II | z-336/337-6 | Duopulse And Unipulse Dental Laser System, Used In Dental Surgery | Manufacturer. Fda Approved The Firm's Corrective Action Plan December 28 | Excel/quantronix Corporation, Hauppauge, New York | December 28, 1995 | Noncompliance With Performance Standards For Laser Products In That The Operator's Manuals Lacked Adequate Calibration Procedures And The Devices Had Several Labeling Noncompliances | 213 units were distributed | Nationwide, Germany, Korea
1995-01-24 | Devices | II | z-346-6 | Tempo/tingle Timers Installed With Diagnostic X-ray Systems, Used To Regulate The Duration, And To Some Extent, The Amount Of Exposure To X-radiation | Manufacturer. Fda Approved The Firm's Corrective Action Plan January 3 | Tingle X-ray Products, Inc | January 3, 1996 | Devices Were Not Tested For Accuracy In Accordance With The Specifications, And Did Not Bear Proper Certification And Identification Labels, As Required By 21 Cfr 1020 | 17 timers | Nationwide
1995-01-24 | Devices | II | z-347-6 | Immunocard Helicobacter Pylori Test Kit, Catalog #710030, A Rapid Enzyme Immunoassay For The Detection Of Igg Antibodies To Helicobacter Pylori In Human Serum And Plasma | Manufacturer | Meridian Diagnostics, Inc | November 14, 1995 | The Enzyme Conjugate In The Kit Was Prepared At The Wrong Dilution, Resulting In An Improper Sensitivity Of The Tests | 175 test kits of lot #710030 | Nationwide, Italy, Russia, Puerto Rico
1995-01-24 | Devices | II | z-348-6 | Captia Rubella-m Eia Test Kit, Product #801-165 | Centocor | Centocor, Uk Ltd | June 23, 1995 | There Is A Higher Incidence Of Equivocal/positive Results Than Anticipated | 200 kits were distributed | Nationwide And International
1995-01-24 | Devices | II | z-349/357-6 | Critikon Brand Protectiv And Protectiv-plus I | Johnson And Johnson | Johnson And Johnson Company, Formerly Known As Critikon, Inc | April 10, 1995 | An Opening May Be Present In The Clear Plastic Portion Of The Blister Package, Therefore Compromising The Sterile Barrier Of The Package | 778,191 catheters were distributed | Nationwide And International
1995-01-24 | Devices | III | z-358-6 | Cellfree Interleukin-2 Receptor (il-2r) Bead Assay Kit, Catalog #ak3120, For The Quantitative Measurement Of Interleukin-2 Receptor Level In Human Serum | Manufacturer | T Cell Diagnostics, Inc | January 28, 1994 | Controls Provided With The Product Were Not Meeting The Assigned Ranges And Fell Below Specified Limits | 20 kits were distributed | Massachusetts, Texas, Connecticut, Louisiana, Greece
1995-01-24 | Devices | III | z-360-6 | Glucometer Encore Test Strips, Used To Measure Glucose In Whole Blood | Manufacturer | Bayer Corporation, Elkhart, Indiana | July 7, 1995 | Subject Lots May Fail To Give Blood Glucose Reading, And The Glucometer Encore Meter Will Display An E-3 Error Code | 45,451 cartons | Nationwide