## Report-1995-03-29
recall_classification_date | product_type | classification | recall_number | product | recalling_firm | manufacturer | recall_initiation_date | reason | volume | distribution
1995-03-29 | Devices | II | z-455-5 | Monoject Hypodermic Needles, Single Use, Epoxy Needle Attachment, Polypropylene Hub, Luer Lock Hub, 19 Ga | Sherwood Medical Company | Sherwood Medical Company, Deland, Florida | December 21, 1994 | The Needles May Contain A Coating Of Silicone Oil In The Interior Of The Needle Hub | 642,600 needles of lot 237593 and 276,900 needles of lot 238057 were distributed | Nationwide And Canada
1995-03-29 | Devices | II | z-516-5 | Chemspor Biological And Chemical Indicator, A Bi For Monitoring Steam Sterilization Processes In Wet Environments | American Sterilizer Company | Amsco International, Inc | February 22, 1995 | The Device Is Labeled For 48-hour Incubation And 510(k) Clearance Is For 7 Days | 9,988 boxes of product were distributed | Nationwide And International
1995-03-29 | Devices | II | z-517-5 | Proof Flash Self-contained Biological And Chemical Indicator, A Bi For Monitoring Steam 270xf Steam Gravity Sterilization Cycles | American Sterilizer Company | Amsco International, Inc | February 22, 1995 | The Product Is Labeled For A 24-hour Or 48-hour Incubation Period | 25,462 cases (50 vials per box) were distributed | Nationwide And International
1995-03-29 | Devices | II | z-518-5 | Saxon Gold Ultra Spermicidal Lubricated Latex Condoms With Nonoxynol-9, Packed In Plastic Film | Manufacturer | Safetex, Division Of Aladan Corporation, Colonial Heights, Virginia | September 21, 1994 | Condoms Were Found To Have A Lower Concentration Of The Spermicide, Nonoxynol-9, Than Was Represented In The Device's Labeling And Premarket Notification | 20 gross were distributed; firm estimates none remains on the market | New York, Kansas
1995-03-29 | Devices | II | z-540/541-5 | Mrinnervu Disposable Endorectal Coils For Mri Imaging For The Colon And Cervix: (a) Cervix Coil Model Bcr-15; (b) Colorectal Coil Model Pcc-15 | Manufacturer | Medrad, Inc | January 31, 1995 | The Packaged Devices Had Broken Shafts | approximately (a) 812 units; (b) 915 units were distributed | Nationwide And International
1995-03-29 | Devices | II | z-542/543-5 | Coveen Urisheath/uriliner, Male External Catheters: (a) Pediatric/geriatric Catheter, Product #5120h-25mm; (b) Large Catheter, Product #5135h, 35mm | Coloplast | Coloplast A/s, Humlebaek, Denmark | January 31, 1995 | The Labeling Incorrectly Identifies The Product As Non-latex Product | 3,500 units of lot 86811 and 200 units of lot 86764 were distributed | Nationwide
1995-03-29 | Devices | II | z-545/547-5 | Adult, Child, And Infant Cpr Resuscitator Bags, Single Patient Use, Disposable, With Related Accessories (mask, Peep Valve, Peep Adapter And Manometer Port | Mercury Medical | Plaxtron Industrial (m) Sdn | August 29, 1994 | The Devices Have Excess Glue Around The Intake Valve On The Interior Of The Bag | 23,610 units were distributed | Nationwide And Puerto Rico
1995-03-29 | Devices | II | z-548/549-5 | Transflator Ii Balloon Inflation Syringes: (a) Transflator Balloon Inflation Syringe, Catalog #05-0802; (b) Transflator Ii Balloon Syringe With Ptca Accessory Kit, Catalog #05-0804 | Manufacturer | Surgical Technologies, Inc | February 2, 1995 | Devices Have Broken Transmitting Led Solder Joints Which May Result In Loss Of Transmission And The "signal Lost" Read- Out On The Monitors | 1,732 units were distributed; firm estimates none remains on the market | Nationwide, Belgium, Canada, Italy, The Netherlands, United Kingdom
1995-03-29 | Devices | II | z-550-5 | Medical Research Laboratory Dpd Defibrillator, Sold Under Datascope Part #0998-00-0004, Used In Conjunction With Datascope Passport Patient Monitoring System | Datascope Corporation | Medical Research Laboratories, Buffalo Grove, Illinois | March 2, 1995 | There Is A Potential Of The Devices -- When Used In The Synchronous Mode With The Datascope Passport Monitor For The Purpose Of Cardioversion -- To Discharge As Late As Approximately 160 Ms After The Peak Of The R Wave Of The Electrocardiogram | 656 units were distributed | Nationwide And International
1995-03-29 | Devices | II | z-558-5 | Encore Orthopedic Foundation Brand Knee Tibial Insert, 13mm Size 4 | Manufacturer | Encore Orthopedics, Inc | October 7, 1994 | The Label On The Interior Pouch May Incorrectly Identify The Insert Thickness As 15mm, However, The Actual Thickness Is 13mm Which Is Correctly Identified On The Exterior Box Label | 15 inserts were distributed; firm estimates none remain on market | New York, Oklahoma, Pennsylvania, South Carolina, Utah, Texas, Washington State
1995-03-29 | Devices | II | z-559-5 | Slit Rotation Knob On The Retroskop Cs, An Illumination Module (microscope Attachment) | Manufacturer | Carl Zeiss, Inc | February 6, 1995 | The Knob, Originally Designed To Be Removable So It Could Be Sterilized, Has The Potential To Loosen And Fall Off From The Retroskop After Being Sterilized | 80 attachments were distributed | Nationwide, Canada
1995-03-29 | Devices | II | z-565-5 | D4600 Air Exchange Line, Sterile/single Patient Use, An Accessory For Ophthalmic Use | Manufacturer | Storz Instrument Company, St | November 28, 1994 | Due To A Packaging Deficiency, The Pouches Were Improperly Sealed, Therefore Compromising Sterility | 2,213 boxes (10 pouches per box) were distributed | Nationwide And International