## Report-1995-07-26
recall_classification_date | product_type | classification | recall_number | product | recalling_firm | manufacturer | recall_initiation_date | reason | volume | distribution
1995-07-26 | Devices | II | z-853-5 | Bard Ptfe Vascular Grafts, Thinwall Straight, Intended For Use As A Bypass Conduit In The Reconstruction Of Occluded Or Diseased Blood Vessels | Bard Vascular Systems Division | Bard Vascular Systems Division, Billerica, Massachusetts | August 5, 1994 | Longitudinal "crack" In The Wall | 1,397 units were distributed | Nationwide, Ireland, Italy
1995-07-26 | Devices | II | z-874/877-5 | Cardiac Pulse Generators, Models 8220e, 8223e, 8224, 8232 | Telectronics Pacing Systems | Telectronics Pacing Systems, Miami Lakes, Florida | April 19, 1995 | Devices May Contain Defective Flex Circuits, Which Can Cause An Intermittent Or Permanent "no Output" Condition | 23 units | Nationwide, Canada, Europe
1995-07-26 | Devices | II | z-879/881-5 | Port-a-cath Dual Lumen Venous Access System With Silicone Catheter: (a) Reorder Numbers: 21-8010 And 21-8010-22; (b) Reorder Numbers: 21-8050 And 21-8050-22; (c) Reorder Number: 21-8052 | Manufacturer | Sims Deltec, Inc | May 3, 1995 | Possible Leaks | 2,378 units were distributed | Nationwide And International
1995-07-26 | Devices | II | z-944/945-5 | Safco Supergrade Autochuck Handpiece Headcaps, Product #wphc, A Component Of The 4-hole Supergrade Plus Autochuck Dental Handpiece, Product #wp4 | Safco Dental Supply Company | Nakanishi Dental Manufacturing Company, Ltd | June 16, 1995 | The Headcap May Come Off The Handpiece During Use | 791 handpieces and 383 replacement cartridges were distributed affecting a total of 1,174 headcaps | Nationwide
1995-07-26 | Devices | II | z-946-5 | Service Manuals Distributed With The Plumesafe 1202 (part #pss1202) And Porta Plumesafe 602 (part Ppss602) Smoke Evacuators; And The Instruction Sheet Distributed With The Virosafe 12 (part #vs12011 And Vs12013) And Virosafe 6 (part Numbers: Sing Vs6, Vs0606011, Vs06013) Smoke Evacuator Filters | Manufacturer | Medetek Devices, Inc | June 9, 1995 | Service Manual And Instructions For These Devices Make Unsupported Claims Of Anti-microbial Properties | 236 lume/porta plumesafe devices and 88 virosafe filters | Connecticut, Florida, Georgia, Illinois, Iowa, Michigan, Minnesota, Missouri, Nebraska, New Jersey, New York, North Carolina, Oklahoma, Pennsylvania, Tennessee, Texas, Virginia, Canada, Austria, Iceland, Italy, Malaysia
1995-07-26 | Devices | II | z-954/955-5 | Single Chamber Cardiac Pulse Generators, Models 1206 And 1206m | Telectronics Pacing Systems | Telectronics Pacing Systems, Miami Lakes, Florida | April 19, 1995 | The Pulse Generators May Contain Defective Ceramic Substrate Pads, Which Can Cause An Open Contact Condition In The Pacer Circuit And Result In Sudden Loss Of Output | 1 unit of model 1206m and 179 units of model 1206 | Nationwide And International