## Report-1996-01-10
recall_classification_date | product_type | classification | recall_number | product | recalling_firm | manufacturer | recall_initiation_date | reason | volume | distribution
1996-01-10 | Drugs | II | d-063/064-6 | Metronidazole, Rx Oral Synthetic Antiprotozoal And Antibacterial Agent, Under The Zenith And Goldline Labels: (a) 250 Mg In Bottles Of 100, 250, And 500; (b) 500 Mg In Bottles Of 50 And 100 | Manufacturer | Zenith Laboratories, Inc | unknown recall date | Superpotency. | lot #configuration amount dosage 2971-833 bottles of 100's 8731 250 mg 2971-837 bottles of 250's 1217 250 mg 3528-706 bottles of 250's 1404 250 mg 3528-706 bottles of 500's 469 250 mg 3007-820v bottles of 50's 394 500 mg 3007-820v bottles of 100's 6000 500 mg 3007-824 bottles of 100's 4859 500 mg | Nationwide
1996-01-10 | Drugs | III | d-056/059-6 | Cardizem Cd (diltiazem Hcl), Packaged In The Following Strengths And Sizes: (a) 120 Mg | Manufacturer | Hoechst Marion Roussel, Inc | November 2, 1995 | Several Capsules May Contain Metal Particles | 105,522/90-capsule bottles and 295/5,000-capsule containers were distributed | Nationwide
1996-01-10 | Drugs | III | d-060-6 | Bulk Magnesium Sulfate Anhydrous, In 200 Pound Fiber Drums, For Drug Use | Manufacturer | Mallinckrodt Chemical, Inc | November 16, 1995 | Product Does Not Meet Heavy Metal Specifications | 61 drums were distributed | Illinois, Missouri, Florida, New Jersey, Indiana, Ohio, Georgia, Delaware, Puerto Rico, Holland
1996-01-10 | Drugs | III | d-061-6 | Carbamide Peroxide 6 | Manufacturer | 's (clay-park) And The Following Private Labels: Consumer Value Stores (cvs) [18 Mo | December 6, 1995 | Subpotenc | 519,330 bottles of the lots with the 24 month expiry period, and 536,739 bottles with the 18 month expiry period | Nationwide, Puerto Rico, South America, West Indies
1996-01-10 | Drugs | III | d-062-6 | Fujisawa Cytarabine Injection, 20 Mg/ml, Rx Antineoplastic Sterile Lyophilized Powder For Reconstitution And Intravenous Powder Used In Combination With Other Approved Anticancer Drugs | Fujisawa Usa | F | December 29, 1995 | Presence Of Visual Silicone Particulates | 10,263 vials were distributed; firm estimated that less than 400 vials remained on the market at time of recall initiation | Nationwide
1996-01-10 | Drugs | III | d-065/066-6 | (a) Cetacaine Topical Anesthetic Liquid, In 56 Gram Bottles, Rx, Indicated For Use To Control Pain; (b) Cetacaine Topical Anesthetic Ointment, In 37 Gram Jars, Rx, Indicated For Use To Control Pain | Manufacturer | Cetylite Industries, Inc | April 14, 1995 | Products Do Not Meet Stability Specifications | (a) 1,900 bottles of batch 408 and 1,982 bottles of batch 409; (b) 2,586 jars of batch 641 and 2,529 jars of batch 642 were distributed; (a) firm estimated that 26 -4- bottles from batch 408 and 94 bottles from batch 409; (b) 48 jars from batch 641 and 281 jars from batch 642 remained on market at time of recall initiation | Nationwide
1996-01-10 | Devices | II | z-207/208-6 | Asm Mattress And Psm Mattress: (a) Alternating Pressure Mattress, Asm- 776/780; (b) Powered Low Air Loss Mattress, Psm-176/180 | Manufacturer | Gaymar Industries, Inc | July 18, 1995 | Some Of The Lower Cells Of The Mattress Systems Exhibited Internal Separation Of Button Welds Resulting In An Observable Bulging Effect Of The Mattress | (a) 1,880 units; (b) 38 units | Nationwide
1996-01-10 | Devices | II | z-261/263-6 | Sof | Manufacturer | Gaymar Industries, Inc | July 7, 1995 | The Bed Cushions May "balloon" Or "expand" Beyond The Cushions' Normal Dimensions If A Button Weld Failure Were To Occur And When Used With The Scn-or Cp-300 Continuous Inflation Pump | (a) 49,930 units; (b) 36,598 units; (c) 5,000 units were distributed | Nationwide And International
1996-01-10 | Devices | II | z-264/265-6 | 3m Avi Burette Iv Administration Sets, Models: 8c1230, 8c2650, 8c1630, Aj2304; (b) Iv Burette Administration Set With Interlink Injection Ports, Models Cj4304, Ck4303, Ck2304 | 3m Infusion Therapy | Korea Green Cross Corporation, Seoul, South Korea | December 3, 1995 | The Sets Leak At Connections Of The Lower Piece Of Tubing To The Y-piece And Connections Of The Lower Piece Of Tubing To The Luer Locks | (a) 500,000 sets; (b) 150,000 sets were distributed | Nationwide And International
1996-01-10 | Devices | II | z-267/319-6 | Horizon I | Mcgaw | Mcgaw Of Puerto Rico, Inc | September 26, 1995 | The Set Tubing Is Being Cut By The Pump Tube Pincher Mechanism, And As A Result, Leakage Of Medication Or Air Entering The I | 35,152 cases were distributed | Nationwide And International
1996-01-10 | Devices | II | z-320-6 | Olympus Brand Disinfectant Solution, For Use With Olympus Automated Blood Typing And Clinical Chemistry Analyzers | Olympus America | Decon Laboratories, Inc | November 27, 1995 | The Disinfecting Solution Bottles Are Ballooning And Distended, And Contain A Dark Colored Sediment | 86 units were distributed | Nationwide, Canada
1996-01-10 | Devices | III | z-260-6 | Trakstar 18 Ptca Dilatation Catheter With Slx Coating, Catalog #535-020, A Balloon Catheter Used To Widen Arteries During Cardiac Catherization Procedures | Manufacturer | Cordis Corporation, Miami Lakes, Florida | November 20, 1995 | The Incorrect Dimensions Are Listed On The Side Panel Of The Device's Shelf Packaging. The Front Box Label And The Inner Pouch Are Labeled With The Correct Balloon Dimensions Of 2. 0 Mm (inflated Balloon Width) X 20 Mm (inflated Balloon Length). The Mislabeled Side Panel Indicated Dimensions Of 3. 0 Mm X 30 Mm. -7-_______________ update Recall #z-255/256-6, Quinton Instrument Q-710 Stress Electrocardiograph And Q-4500 Stress Test System, Which Appeared In The January 3, 1996 Enforcement Report Should Read: Code: Serial Numbers: (a) 101 Through 181; (b) 2171 Through 2284 (serial Numbers Are Non- Consecutive). | 14 units were distributed | Indiana, Michigan, California, Arizona, Massachusetts, Florida, Illinois