## 1993-03-03
RECALL_CLASSIFICATION_DATE | PRODUCT_TYPE | CLASSIFICATION | RECALL_NUMBER | PRODUCT | RECALLING_FIRM | MANUFACTURER | RECALL_INITIALIZATION_DATE | REASON | VOLUME | DISTRIBUTION
1993-03-03 | Drugs | III | d-145/146-3 | Laxative Products Under The Rx Place Label: (a) Stool Softener Docusate Sodium 100 Mg 100 Softgels Otc; (b) Laxative With Stool Softener Casanthranol And Docusate Sodium 100 Softgels Otc. Recall #d-145/146-3. | G&w; Labs S. Plainfield New Jersey (responsible Firm) | Pharmacaps Inc. Elizabeth New Jersey. | October 27, 1992 | Main label panels correctly declare ingredients | (a) 1 512 Containers; (b) 1 440 Containers Were Distributed. | Massachusetts Rhode Island New Jersey New York Pennsylvania.
1993-03-03 | Drugs | III | d-160-3 | S. Product Nitrotan Germicidal Wound Spray In 4 Ounce And 8 Ounce Aerosol Cans An Otc For Topical Use. Recall #d-160-3. | Manufacturer | Cramer Products Inc. Gardner Kansas. | November 19, 1992 | Internal mechanism of the aerosol valve may fail, allowing product leakage | 52 980 Cans Were Distributed. | Nationwide.
1993-03-03 | Biologics | II | b-149-3 | Red Blood Cells. Recall #b-149-3. | Manufacturer | Hunter Blood Center Inc. Clearwater Florida. | March 18, 1992 | Blood product, which tested negative for all required and recommended viral markers, but was collected from a donor who experienced an occupational needle stick exposure approximately seven months prior to donation, was distributed | 1 Unit. | Florida.
1993-03-03 | Biologics | II | b-156-3 | Red Blood Cells. Recall #b-156-3. | Manufacturer | United Blood Services Blood Systems Inc. San Angelo Texas. | April 2, 1992 | Red blood cells, prepared more than eight hours after the collection of the whole blood units, were distributed | 9 Units. | Texas.
1993-03-03 | Devices | II | z-119/120-3 | (a) Stainless Steel Smoke Evacuation Vaginal Speculums Standard Size; (b) Stainless Steel Smoke Evacuation Vaginal Speculums Large Size. Recall #z-119/120-3. | Ellman International Inc. Hewlett New York | Macho International Pakistan. | April 21, 1992 | The underside of the arm of the speculum around the area of the evacuation tube may oxidize/rust depending on the sterilization technique used | 120 Units Were Distributed. | Nationwide.
1993-03-03 | Devices | II | z-231/244-3 | I.v. Catheter Extension And Y-type Extension Sets Single-use Devices Which Are Used On Both Adults And Children: Baxter Brand Catheters: (a) Catalog #2n1190; (b) Catalog #2n1192; (c) 2n1194; Medex Brand Catheters: (d) Mx451-fl (e) Mx453; (f) Mx453-l; (g) Mx453-2l (h) Mx453-sl (i) Mx456-l; (j) M458l; (k) Mx459; (l) Mx596-c; (m) Mx613; (n) Mx6251. Recall #z-231/244-3. | Medex Inc. Hillard Ohio | Baxter Healthcare Corporation Deerfield Illinois. | October 8, 1992 | The female luer is cracking and leaking | 1 808 783 Devices Were Distributed. | Nationwide And International. -3-
1993-03-03 | Devices | II | z-261-3 | Boston Envision Rigid Gas Permeable Daily Wear Contact Lenses (blue) Part #80168. Recall #z-261-3. | Manufacturer | Polymer Technology Wilmington Massachusetts. | November 2, 1992 | The lenses were cut as -3 | 51 Lenses Were Distributed. | Ohio Texas France Poland Yugoslavia.
1993-03-03 | Devices | II | z-262-3 | Jude Medical 9 Fr Rediguard Dl 40 Cc Intra-aortic Balloon Catheter Kit Catalog #060-0525 Containing Umi Cath-seal Percutaneous Catheter Introducer (pci). Recall #z-262-3. | St. Jude Medical Inc. Chelmsford Massachusetts. Salesforce Notified Hospitals On December 25 1992 Followed | Universal Medical Instruments Ballston Spa New York. | December 25, 1992 | The bond between the sheath and hub assembly may separate due to contact with aqueous solutions | 115 Kits Were Distributed. | Florida Texas Pennsylvania Missouri Michigan Illinois.
1993-03-03 | Devices | II | z-327-3 | Tens (transcutaneous Electrical Nerv Stimulator) A Rx Device Distributed Under The Dial Away Pain 400 And Synaptic 1000 Labels. Recall #z-327-3. | Relief From Pain Inc. Aurora Colorado | Halleck-willard Inc. Fredrick Colorado. | August 5, 1992 | Device may deliver a small shock to the patient | 16 Units Were Subject To Field Correction As Of 8/3/92. | Iowa California Michigan North Carolina Colorado New York Washington State Illinois Nebraska North Dakota.
1993-03-03 | Devices | II | z-330/331-3 | Sterile Mayfield Disposable Skull Pins Used To Keep A Patient's Head Stationary During Brain Surgery: (a) Stainless Steel; (b) Stainless Steel With A Polycarbonate Base Recall Z-330/331-3. | Manufacturer | Ohio Medical Instrument Company Cincinnati Ohio. -4- | January 22, 1993 | Both products were marketed without an approved 510(k) | Approximately (a) 58; (b) 439 Boxes (12 Units Per Box) Were Distributed. | Nationwide Germany Canada Japan Spain.
1993-03-03 | Devices | II | z-347-3 | Life Pulse High Frequency Ventilator Humidifier Cartridges Also Called Patient Breathing Circuits. Recall #z-347-3. | Manufacturer | Bunnell Inc. Salt Lake City Utah. | November 13, 1992 | The cartridges are susceptible to cracking at the hose ports, which could result in breakage and failure | Approximately 11 386 Units Were Distributed. | Nationwide Japan Korea.
1993-03-03 | Devices | III | z-283-3 | Nor-partigen Control Set I An In-vitro Diagnostic Kit For Use With The Monospecific Nor-partigen Plates For The Quantitative Determination Of Immunoglobulins In Human Serum. Recall #z-283-3. | Behring Diagnostics Somerville New Jersey | Behringwerke Ag Marburg Germany. | July 17, 1990 | Iqm constituent valve of these controls was reading greater than 15% of its high value | 188 Sets Were Distributed. Firm Estimates None Remains On The Market Product Is Expired. | Nationwide.
1993-03-03 | Devices | III | z-291-3 | Vacutainer Brand Sst Blood Collection Tubes. Recall #z-291-3. | Becton Dickinson Vacutainer Systems Rutherford New Jersey | Becton Dickinson Vacutainer Systems Sumter South Carolina. | July 20, 1992 | Product may break in the centrifuge due to firecrack defects in the bottom of the tube | 436 000 Units Were Distributed. | Nationwide Canada England Japan.
1993-03-03 | Devices | III | z-308-3 | Coumatrak Protime Controls Lyophilized Whole Blood Based Control Materials For Two Levels Of Prothrombin Time Normal And Abnormal Packaged In 1 Ml Vials For In-vitro Use Only As A Check On The Analytical Performance Of The Coumatrak Monitor And Protime Test System. Recall #z-308-3. | Biotrack Inc. Mountain View California | Consolidated Technologies Inc. Austin Texas. | November 23, 1992 | The controls take longer to clot than the clotting time ranges on the label | 11 748 Boxes (5 Vials Per Box) Were Distributed. | Delaware.
1993-03-03 | Devices | III | z-318/320-3 | Ventricular Drainage System Intended For Single Use To Facilitate The Collection Of Cerebrospinal Fluid (csf) Fluid Injection Csf Sampling And Intracranial Pressure Monitoring: (a) Ventricular Drainage System Model 041; (b) Ventricular Bolt Pressure Monitoring & Drainage Kit Model 110-4hc; (c) Microventricular Bolt Pressure Monitoring & Drainage Kit Model 110-hmc. Recall #z-318/320-3. | Camino Laboratories Inc. San Diego California | Plasco Inc. Gurnee Illinois (component). | September 1, 1992 | The bushings of the ventricular tubing do not fit properly which could lead to leakage at the bushing and tubing connection | 2 920 Units Were Distributed. | Nationwide Argentina Italy.
1993-03-03 | Devices | III | z-321/323-3 | Nova Stat Profile Control Level 2: (a) Nova Catalog #13425 (baxter Catalog #b6352-49) Stat Profile Control Level 2; (b) Nova Catalog #13427 (baxter Catalog #b6352-51) Stat Profile Control Multipak With Glucose; (c) Nova Catalog #13428 Stat Profile Control Qap Multipak; Recall #z-321/323-3. | Manufacturer | Nova Bio Medical Waltham Massachusetts. | August 26, 1992 | Some ampoules labeled as containing level 2 control actually contained level 1 control | (a) 23 Boxes Of 30; (b) 508 Boxes Of 30; (c) 20 Boxes Of 30 Ampoules Were Distributed. | Nationwide And International.
1993-03-03 | Devices | III | z-324-3 | Concise Strep A Kit An In-vitro Diagnostic Assay Used For The Qualitative Determination Of Group A Streptococcus Antigen From Throat Swabs. Recall #z-324-3. | Hybritech Inc. San Diego California | Pacific Biotech Inc. San Diego California. | December 2, 1992 | The product will produce a false positive result | 2 265 Kits Were Distributed. | Nationwide.
1993-03-03 | Devices | III | z-325-3 | Total Iron And Total Iron Binding Capacity (tibc) Test Kit. Recall #z-325-3. | Manufacturer | Bio-analytic Laboratories Inc. Palm City Florida. | May 22, 1991 | One of the reagents in the kit is precipitating | 10 Kits Were Distributed. | Michigan Mexico Egypt Turkey.
1993-03-03 | Devices | III | z-328-3 | Repliplate Cathra 256 Mcg/ml Sulfisoxazole (sul) Antibiotic Susceptibility Test Plate Used To Check The Susceptibility Of Disease-causing Bacteria To Antibiotics. Recall #z-328-3. | Manufacturer | Automed Inc. Arden Hills Minnesota. | December 11, 1992 | The plates were contaminated | 80 Sets Containing 1 Sulfisoxazole Plate Per Set Were Distributed. | Canada.
1993-03-03 | Devices | III | z-326-3 | Testosterone Test Standard A Component Of The Rsl 1251 Testosterone Test Kit Which Detects The Total Unconjugated Form Of This Steroid For In-vitro Diagnostic Use Only. Recall #z-326-3. -7- | Manufacturer | Icn Biomedicals Inc. Costa Mesa California. | October 12, 1992 | Lack of product stability when stored above the recommended -15c | 228 Sets Were Distributed. | California Illinois Indiana New Jersey Pennsylvania Texas Washington State Wisconsin.
1993-03-03 | Devices | III | z-329-3 | Abbott Testpack Plus Hcg-combo Test Kits List #3a60 A Self-performing Immunoassay In-vitro Diagnostic Designed For The Qualitative Determination Of Human Chorionic Gonadotropin (hcg) In Urine And Serum For Early Detection Of Pregnancy. Recall #z-329-3. | Abbott Laboratories Diagnostics Division Abbott Park Illinois | Abbott Laboratories North Chicago Illinois. | June 25, 1991 | Some clinical specimens could give false positive results | 1 338 20-test Kits And 1 868 40-test Kits Were Distributed; Firm Estimates None Remains On The Market. | Nationwide Canada Costa Rica Dominican Republic Venezuela Australia Germany Hong Kong.
1993-03-03 | Devices | III | z-334-3 | Concise Strep A In-vitro Diagnostic Kits Intended For The Qualitative Determination Of Group A Streptococcus Antigen From Throat Swabs. Recall #z-334-3. | Manufacturer | Hybritech Inc. San Diego California. | October 7, 1992 | The wrong revision directions for use were packaged with the kit | 773 Kits Were Distributed. | Nationwide Canada.