## 1994-06-29
RECALL_CLASSIFICATION_DATE | PRODUCT_TYPE | CLASSIFICATION | RECALL_NUMBER | PRODUCT | RECALLING_FIRM | MANUFACTURER | RECALL_INITIALIZATION_DATE | REASON | VOLUME | DISTRIBUTION
1994-06-29 | Drugs | II | d-305-4 | Pitressin (vasopressin Injection Usp) Steri-vial 10 Pressor Units In 0.5 Ml Synthetic 25/0.5 Ml Glass Vials. Recall #d-305-4. | Parke-davis Division Of Warner-lambert Company Morris Plains New Jersey | Warner-lambert Company Parke-davis Division Rochester Michigan. | January 20, 1994 | Portion of carton label bears incorrect route of administration | 791 Cartons Of 25 Vials Of Lot 022n2pn 1 024 Cartons Of 25 Vials Of Lot 03792ph Were Distributed. | Nationwide Canada.
1994-06-29 | Drugs | II | d-312/314-4 | Procan Sr Tablets (procaine Hydrochloride Extended Release Tablets) Indicated For The Treatment Of Documented Ventricular Arrhythmias: (a) 250 Mg; (b) 500 Mg; (c) 750 Mg. Recall #d-312/314-4. | Parke-davis Division Of Warner Lambert Morris Plain New Jersey | Warner Lambert Vega Baja Puerto Rico. | NOT-FOUND | Abbreviated new drug application deficiencies | Firm Estimates Little Product Remains On The Market.
1994-06-29 | Drugs | III | d-306-4 | Parke-davis Consumer Health Products Group Brand Benadryl Tablets 25 Mg Otc Antihistamine In Bottles Of 100. Recall #d-306-4. | Warner-wellcome Consumer Health Product Morris Plains New Jersey | Warner-lambert Company Vega Baja Puerto Rico. | April 25, 1994 | Subpotency. | 23 808 Bottles Were Distributed. | Nationwide.
1994-06-29 | Drugs | III | d-307-4 | Methenamine Mandelate 1.0 Tablets Oral Solid Dosage Form Packaged In 100 And 1 000 Tablet Bottles Rx Antiseptic For Urinary Tract Infections. Recall #d-307-4. | Manufacturer | Amide Pharmaceutical Inc. Little Falls New Jersey. | April 7, 1994 | Non-validated manufacturing process | 48 Bottles Of 1000 And 288 Bottles Of 100 Tablets Were Distributed. | New York Florida California Indiana.
1994-06-29 | Drugs | III | d-308/309-4 | Duragesic Fentanyl Transdermal System Controlled Substance Rx Drugs: (a) 25 Mcg/hr; (b) 50mcg/hr. Recall #d-308/309-4. | Janssen Pharmaceutica Titusville New Jersey | Alza Corporation Vacaville California. | February 25, 1994 | Release rate does not meet product specifications | 66 653 Cartons Of Lot 131305 19 061 Cartons Of Lot 131306 35 230 Cartons Of Lot 132306 And 15 159 Cartons Of Lot 132307 Were Distributed. | Nationwide.
1994-06-29 | Drugs | III | d-310-4 | Flurazepam Hydrochloride Capsules Usp 30 Mg Rx Schedule Iv Narcotic Oral Hypnotic For The Treatment Of Insomnia. Recall #d-310-4. | Udl Laboratories Rockford Illinois (repacker/responsible Firm) | Mylan Pharmaceuticals Morgantown West Virginia. | June 15, 1994 | Some unit cartons may be mislabeled as temazepam capsules 15 mg | 2 190 Unit Cartons Were Distributed. | Nationwide.
1994-06-29 | Drugs | III | d-311-4 | Tagamet (brand Of Cimetidine) 800 Mg In Bottles Of 30 Indicated For Short-term Treatment Of Active Duodenal Ulcer. Recall #d-311-4. | Manufacturer | Smith Kline Beecham Pharmaceuticals Cidra Puerto Rico. | May 23, 1994 | A tablet of relafen 750 mg was found in a bottle of tagamet 800 mg | 61 992 Units Were Distributed. | Nationwide.
1994-06-29 | Drugs | III | d-315-4 | Miochol (acetylcholine Chloride) 1:100 Intraocular 10 Mg/2ml In 2 Ml Vials Used To Obtain Complete Miosis Of The Iris In Seconds After Delivery Of The Lens In Cataract Surgery. Recall #d-315-4. | Iolab Claremont California | Iolab Inc. San German Puerto Rico. | May 20, 1994 | Product does not meet ph specifications | 12 225 Units Were Distributed; Firm Estimates None Remains On Market. | Nationwide.
1994-06-29 | Biologics | II | b-308/310-4 | (a) Red Blood Cells; (b) Platelets; (c) Recovered Plasma. Recall #b-308/310-4. | Manufacturer | The American National Red Cross Philadelphia Pennsylvania. | October 19, 1993 | Blood products collected from a donor taking proscar were distributed | 1 Unit Of Each Component. | Pennsylvania California.
1994-06-29 | Biologics | II | b-339-4 | Red Blood Cells. Recall #b-339-4. | Manufacturer | American Red Cross Atlanta Georgia. | August 8, 1993 | Red blood cells, collected from a donor who traveled to an endemic area for malaria and had taken prophylactic anti- malarial medication prior to donation, was distributed | 1 Unit. | Georgia.
1994-06-29 | Biologics | II | b-341/347-4 | (a) Red Blood Cells; (b) Platelets; (c) Cryoprecipitated Ahf; (d) Fresh Frozen Plasma; (e) Platelet Expired; (f) Recovered Plasma; (g) Recovered Human Serum (pool). Recall #b-341/347-4. | Manufacturer | American Red Cross Blood Services Buffalo New York. | September 15, 1993 | Blood products which were collected from a donor with a history of cancer; or tested negative but were collected from donors who previously tested repeatedly reactive for the antibody to the human immunodeficiency virus type 1 (anti-hiv-1), were distributed | (a) 11 Units; (b) 8 Units; (c) 3 Units; (d) 5 Units; (e) 1 Unit; (f) 5 Units; (g) 7 Units. | New York Pennsylvania California South Carolina Oklahoma Alabama Ohio Georgia Massachusetts Switzerland.
1994-06-29 | Biologics | II | b-350/352-4 | (a) Red Blood Cells; (b) Platelets Expired; (c) Recovered Plasma. Recall #b-350/352-4. | Manufacturer (a) | Northwest Florida Blood Center Inc. Pensacola Florida. -5- | May 25, 1993 | Blood products, which tested negative for all viral markers, but were collected from donors who participated in high risk behavior or experienced an occupational needle stick exposure less than 12 months prior to donation, were distributed | 2 Units Of Each Component. | Florida.
1994-06-29 | Biologics | III | b-338-4 | Influenza Virus Vaccine Trivalent Types A And B Subvirion Formula 1993-1994 Connaught Flu Vaccine In Prefilled 0.05 Ml Unit Dose Syringe. Recall #b-338-4. | Manufacturer | Connaught Laboratories Inc. Swiftwater Pennsylvania. | March 29, 1994 | In translating labeling from english to spanish, a portion of the indication section was omitted | 30 000 Units Were Distributed. | Argentina.
1994-06-29 | Biologics | III | b-348/349-4 | (a) Granulocytes Apheresis; (b) Recovered Plasma. Recall #b-348/349-4. | Manufacturer | American Red Cross Blood Services Buffalo New York. | September 15, 1993 | Blood products which were collected from a donor with a history of cancer or stored at an unacceptable temperature were distributed | (a) 1 Unit; (b) 2 Units. | New York California Massachusetts Switzerland.
1994-06-29 | Devices | II | z-740-4 | Prowess 2000 Radiation Treatment Planning Software Package Version 2.08. Recall #z-740-4. | Manufacturer | Small Systems Group Inc. Chico California. | January 13, 1993 | Calculations in the brachytherapy module produced dose errors of 48 percent | 96 Units. -6- | Nationwide And International.
1994-06-29 | Devices | II | z-793-4 | Ac Power Adaptor Or Transformer Labeled "component Power Supply Class 2 Part #394-4010-001 Intended For Use With Medical And Dental Equipment. Recall #z-793-4. | Cns Inc. Chanhassen Minnesota | Ault Inc. Minneapolis Minnesota. | November 8, 1993 | Some adaptors exceeded leakage current specifications, due to loose fuses or fuses in close proximity to the transformer lamination stack | 333 Units Were Distributed. | Nationwide And Canada.
1994-06-29 | Devices | II | z-931/939-4 | Sterile Ophthalmic Solutions For Use With Contact Lenses: (a) Sherman Pharmaceutical Stay-wet 3 Preserved Wetting And "in Eye" Lubricating Solution And Rewetting Drop For Use With Fluoro/silicone Crylate And Silicone Acrylate Rigid Gas Permeable (rpg) Contact Lenses .33 Ounces (10 Ml) Sample Size; (b) Sherman Pharmaceutical Stay-wet 3 Preserved Wetting And "in Eye" Lubricating Solution And Rewetting Drop For Use With Fluoro/silicone Acrylate And Silicone Acrylate Rpg Contact Lenses 1 Fluid Ounce (30 Ml) Size; (c) Wukvision Stay-wet 3 Preserved Wetting And "in Eye" Lubricating Solution And Rewetting Drop For Use With Fluoro/silicone Acrylate And Silicone Acrylate Rpg Contact Lenses 10 Ml Size; (d) Sherman Pharmaceutical De Stat 3 Cleaning Disinfecting And Storage Solution For Use With Fluoro/silicone Acrylate And Silicone Acrylate Rpg Contact Lenses 1 Fluid Ounce (30 Ml) Sample Size; (e) Sherman Pharmaceutical De Stat 3 Cleaning Disinfecting And Storage Solution For Use With Fluoro/silicone Acrylate And Silicone Acrylate Rpg Contact Lenses 4 Fluid Ounces (118 Ml) Size; (f) Wukvision De Stat 3 Cleaning Disinfecting And Storage Solution For Use With Fluoro/silicone Acrylate And Silicone Acrylate Rpg Contact Lenses 118 Ml Size; (g) Wukvision De Stat 3 Cleaning Disinfecting And Storage Solution For Use With Fluoro/silicone Acrylate And Silicone Acrylate Rpg Contact Lenses 30 Ml Size; (h) Sherman Pharmaceutical Sof/pro-clean Cleaning Solution For Use With All Soft (hydrophilic) Contact Lenses In Heat (thermal) And Chemical (not Heat) Lens Care Systems 1 Fluid Ounce (30 Ml) Size; -7- (i) Choongwae Pharma Corporation Sof/prp-clean Cleaning Solution For Use With All Soft (hydrophilic) Contact Lenses In Heat (thermal) And Chemical (not Heat) Lens Care Systems 30 Ml Size. Recall #z-931/939-4. | Manufacturer | Sherman Pharmaceuticals Inc. Mandeville Louisiana. | May 9, 1994 | Product sterility may have been compromised due to potential inadequate sterilization | Approximately 26 371 Units Of Stay-wet 3; Approximately 45 906 Units Of De-stat; And Approximately 63 786 Units Sof/pro-clean From All Lot Numbers Involved Were Distributed. | Nationwide And International.
1994-06-29 | Devices | II | z-963-4 | Ivac Core-check Tympanic Thermometer System Model 2090a A Hand-held Thermometer Which When Positioned In The External Auditory Canal Measures The Infrared Energy Emitted From The Tympanic Membrane. Recall #z-963-4. | Manufacturer | Ivac Corporation San Diego California. | February 19, 1992 | Inadequate shielding allowing high levels of electromagnetic or radio interference to possibly cause device malfunction | 11 518 Units Were Distributed From August 1991 To October 1992; Firm Estimated 10 255 Units Remained On Market At Time Of Recall. | Nationwide And International.
1994-06-29 | Devices | II | z-965-4 | Ipi Respiratory Therapy Fluid Transfer Sets With Luer Tip And Spring Loaded Clamp Catalog #1122 A Sterile Prescription Device Used To Transfer Respiratory Therapy Solutions To The Desired Respiratory Device. Recall #z-965-4. | Manufacturer | Ipi Medical Products Chicago Illinois. | March 10, 1994 | The clamp on the sets were glued in the open position | 37 Cases Were Distributed; Firm Estimated That 7 Cases Remained On Market At Time Of Recall. | Illinois Ohio North Carolina Massachusetts Arizona Georgia.
1994-06-29 | Devices | II | z-968/969-4 | Mde Defibrillator Model 19e And Mde Defibrillator With Pacer Model 20e Options Used With The Escort 300 Series Full Parameter Transport Monitor. Both Options Are Low Energy Dc Defibrillators Which Are Able To Deliver An Electrical Shock With A Range Of 2 To 360 Joules Of Energy. Recall #z-968/969-4. | Manufacturer | Medical Data Electronics Inc. Arleta California. | December 2, 1993 | The energy selector switch can be set between the detented increments of the switch causing higher than selected energy to be delivered | 110 Units Were Distributed. | Nationwide Canada Hong Kong Korea The Philippines.
1994-06-29 | Devices | II | z-978-4 | Hot Choice Microwave Oven Model 122000 Used In Heating Frozen Prepared Pizza Or Chicken Using Convection/microwave Energy. Recall #z-978-4. | Manufacturer. Fda Approved The Firm's Corrective Action Plan June 16 1994. Firm-initiated Field Correction Ongoing. | Kaiser Electroprecision Irvine California. | June 16, 1994 | Product failed to meet the certification labeling requirements | 30 Units. | California.
1994-06-29 | Devices | II | z-979/980-4 | First Medic Defibrillators (a) Model 90510 (fm 510); (b) Model 90526 (fm 610). Recall #z-979/980-4. | Manufacturer | Spacelabs Medical Inc. Redmond Washington. | April 26, 1994 | The material (tin) used for the plating of the isolation module shield is susceptible to metal fiber growth that, if oriented in a certain fashion, can make contact with the connector pins and result in a small electrical short circuit condition for the device | (a) 610 Units; (b) 483 Units Were Distributed. | Nationwide And International.
1994-06-29 | Devices | II | z-981-4 | Mevatron Linear Accelerator And Mevasim Therapy Simulator - Overhead Suspension Hand Control. Recall #z-981-4. | Manufacturer | Siemens Medical Systems Inc. Concord California. | October 29, 1994 | The devices were distributed with a defective weld that can fail and cause the overhead support plate (which weighs 45 pounds) to fall | 1 141 Units. | Nationwide And International.
1994-06-29 | Devices | II | z-982-4 | Mevatron Linear Accelerator Wedge Filter Assembly Used With The Mevatron Therapeutic X-ray Machine. Recall #z-982-4. | Manufacturer | Siemens Medical Systems Inc. Concord California. | October 29, 1993 | The devices were distributed with defective connection screws between the wedge filters and the aluminum trays which could allow the wedge to fall from the mounting tray | Approximately 4 000 Wedges Were Distributed 1976-1993. | Nationwide And International.
1994-06-29 | Devices | III | z-942-4 | Calibration Gas Mixture In Ud Size Cylinders Used As A Calibration Or Check Standard For Anesthesia Monitoring. Recall #z-942-4. | Manufacturer | Scott Medical Products Plumsteadville Pennsylvania. | April 22, 1994 | The product which is 30% oxygen is incorrectly labeled as containing 21% oxygen | 5 Cylinders Were Distributed. | Texas Indiana Iowa California.
1994-06-29 | Devices | III | z-950-4 | Cynomologus Monkey Kidney Tubes Catalog #70-177-d Used For The Isolation Of Virus In Clinical Specimens. Recall #z-950-4. | Manufacturer | Whittaker Bioproducts Inc. Walkersville Maryland. | January 20, 1992 | Product was contaminated with gram positive bacillus species | 125 Units Were Distributed. | Nationwide.
1994-06-29 | Devices | III | z-951-4 | Rhesus Monkey Kidney Tubes With Sv5 And Sv40 Antiserum Cell Cultures Which Is Used For Isolation Of Virus In Clinical Specimens. Recall #z-951-4. | Manufacturer | Whittaker Bioproducts Inc. Walkersville Maryland. | March 4, 1992 | Product was contaminated with micrococcus, bacillus, and coryneform bacteria | 1 474 Units Were Distributed; Firm Estimates None Remains On The Market. | Nationwide.
1994-06-29 | Devices | III | z-964-4 | Skull Mount Kit A-1330 Consisting Of Sterile Gloves Pvp Solutions Small Prep Tray Razor Sponge Gauze Towels Csr Wraps Needles Zylocaine Syringe Scalpel Drape Tray Skull Ring Bone Screws Screw Driver Drill Bit And Hand Drill. Recall #z-964-4. | Manufacturer | Ohio Medical Instruments Cincinnati Ohio. | December 8, 1993 | The product had no sterilization validation | 1 Kit Was Distributed. | Idaho.
1994-06-29 | Devices | III | z-973-4 | Human Neonatal Kidney (hnk) Cell Cultures In Tubes 50 Per Rack Living Cell Cultures For Viral Isolation Procedures In Clinical Virology Labs. Recall #z-973-4. | Manufacturer | Whittaker Bioproducts Inc. Walkersville Maryland. | December 18, 1991 | The device was contaminated with yeast | 3 380 Units Were Distributed; Firm Estimates None Remains On The Market. | Nationwide.
1994-06-29 | Devices | III | z-974-4 | Human Neonatal Kidney (hnk) Cell Cultures In Tubes 50 Per Rack Living Cell Cultures For Viral Isolation Procedures In Clinical Virology Labs. Recall #z-974-4. | Manufacturer | Whittaker Bioproducts Inc. Walkersville Maryland. | December 27, 1991 | Device was contaminated with bacteria | 1 995 Units Were Distributed; Firm Estimates None Remains On The Market. | Nationwide. -11-
1994-06-29 | Devices | III | z-977-4 | Seequence 2/medalist Trial Contact Lenses Not For Sale Intended For Use As Trial Lenses Only And Remain On The Premises Of The Eye Care Practitioner. Recall #z-977-4. | Manufacturer | Bausch & Lomb Inc. Contact Lens Division Rochester New York. | April 14, 1994 | While the blister packs are correctly labeled as -4 | 195 Cartons Were Distributed Between 12/20/93 And 3/25/94. | Nationwide.