## 1994-07-30
RECALL_CLASSIFICATION_DATE | PRODUCT_TYPE | CLASSIFICATION | RECALL_NUMBER | PRODUCT | RECALLING_FIRM | MANUFACTURER | RECALL_INITIALIZATION_DATE | REASON | VOLUME | DISTRIBUTION
1994-07-30 | Drugs | III | d-327-4 | Lindane Lotion Usp 1% Rx Topical Used For The Treatment Scabies In 1 Pint Bottles Sold Under The Following Labels: Pennex Pbi Goldline Gg Major Qualitest Rugby Schein. Recall #d-327-4. | Manufacturer | Pennex Pharmaceuticals Inc. Formerly Known As Pharmaceutical Basics Inc. Morton Grove Illinois. | June 29, 1994 | Lots 19490, 19516 and 19806 do not meet content uniformity specification | 28 856 Bottles Were Distributed; Firm Estimates 25% Of The Product Remains On The Market. | Nationwide.
1994-07-30 | Drugs | III | d-328/329-4 | Salsalate Tablets Usp: (a) 500 Mg In 100 And 1000 Count Bottles; (b) 750 Mg In 100 And 1000 Count Bottles. Recall #d-328/329-4. | Manufacturer | Caraco Pharmaceutical Laboratories Inc. Detroit Michigan. | July 20, 1993 | Products do not meet dissolution specifications | (a) 85 Bottles Of 100 Of Lots 3fe01b And 3ap02b; 26 Bottles Of 1 000 Of Lot 3fe01a; (b) 62 Bottles Of 100 Of Lot 3fe02b; And 316 Bottles Of 1 000 Of Lots 3ap01a And 3fe02a. | Colorado Missouri Arkansas Illinois Michigan Tennessee Louisiana Florida.
1994-07-30 | Biologics | II | b-316-4 | Blood Bank Information System (bbis) Computer Software. Recall #b-316-4. | Manufacturer | Collaborative Medical Systems Inc. Waltham Massachusetts. | June 25, 1993 | Software, that contained programming errors which could potentially result in the release of unsuitable blood products, was distributed | 5 Systems. -3- | Iowa Massachusetts Oklahoma.
1994-07-30 | Devices | II | z-954/955-4 | (a) Ef Synergy Quad Plated Media Used As A Predictor Of Synergy Of An Aminoglycoside And Penicillin When Testing Enterococci Isolates; (b) Chocolate Agar Mono Plated Media Used In The Culturing Of Pathogenic Neisseria And Haemophilus. Recall #z-954/955-4. | Manufacturer | Remel Limited Partnership Lenexa Kansas. | September 29, 1992 | The control quadrant for the ef synergy quad plated media failed to support the growth of the challenge organisms used as positive controls; and the quality control organism, haemophilus influenzae, for the chocolate agar mono plated media did not grow on this lot of product | (a) 660 Plates; (b) 1 700 Plates Were Distributed. Firm Estimates None Remains On The Market. | (a) Arkansas Maine Massachusetts New Hampshire New York Rhode Island California Ohio Michigan Utah Nebraska Missouri Florida Virginia Pennsylvania; (b) California Utah New Mexico.
1994-07-30 | Devices | II | z-983/984-4 | Pro-lok Beech Brand Shielded Needle Connector Needle Housing With Removable Needle 18ga X 5/8 (16 Mm) Sterile: (a) Catalog #01-9018; (b) Catalog #01-9020. Recall #z-983/984-4. | Beech Medical Products Newtown Pennsylvania | Vlv Associates Inc. East Hanover New Jersey. | August 28, 1992 | The female luer of the pro-lok housing cracks and may leak when used with certain iv solutions | Approximately 2 Million Plus Units Were Distributed; Firm Estimates None Remains On The Market. | Pennsylvania Maryland New York Massachusetts Illinois Texas Michigan California Ohio Colorado Missouri Washington State Minnesota Indiana.
1994-07-30 | Devices | II | z-1078-4 | Cardiopump Acd (active Compressed/decompression) Resuscitator Both Prototype And Production Models. Recall #z-1078-4. -4- | Ambu Inc. Linthicum Maryland | Ambu International A/s Glostrup Denmark. | April 19, 1994 | The firm was using data from an unapproved clinical investigation on human subjects to support a 510(k) application | 506 Devices. | California Minnesota New York.
1994-07-30 | Devices | II | z-1080-4 | Storz Intraocular Lens Model P010uv. Recall #z-1080-4. | Manufacturer | Storz Ophthalmics Inc. Clearwater Florida. | April 20, 1994 | The lenses were mislabeled with the incorrect refractive index | 21 Lenses Remain In Commerce That May Still Be Unimplanted. | Ohio Florida California Singapore Germany Brazil Guatemala.
1994-07-30 | Devices | II | z-1088-4 | Microvasive Ablaser Side-firing Laser Guide A Laser Delivery Guiding Catheter Indicated For Use For Soft Tissue Ablation And Hemostasis Treatment In Urological Conditions. Recall #z-1088-4. | Boston Scientific Corporation Watertown Massachusetts March 22 1994. Firm-initiated Recall Complete. | Boston Scientific Corporation Milford Massachusetts. | March 22, 1994 | The gold plated catheter tip may detach during surgery causing the tip to remain within the body or the laser reflecting mirror in the tip may char and burn tissue | 179 Units Were Distributed; Firm Estimates None Remains On The Market. | Michigan California Missouri Mississippi Illinois New York Kentucky Virginia Tennessee Alabama Indiana Massachusetts Pennsylvania Colorado Florida Georgia Washington D.c.
1994-07-30 | Devices | II | z-1089/1093-4 | Disposable Reservoir Bags For Infumed And Walkmed Infusion Pumps: (a) Ipr-86 Reservoir Bag 65 Ml; (b) Ipr-150 Reservoir Bag 150 Ml; (c) Ipr-250 Reservoir Bag 250 Ml; (d) Mf-65 Reservoir Bag 65 Ml; (e) Pr-70 Reservoir Bag 65 Ml Recall #z-1089/1093-4. -5- | Manufacturer | Medex Inc. Duluth Georgia. | June 11, 1993 | The female luer lock will crack when being used preventing administration of proper amount of prescribed fluids | 171 010 Bags. | Nationwide United Kingdom Germany.
1994-07-30 | Devices | II | z-1094/1095-4 | Baxter/travenol Fenwal Blood Warmer Used To Warm Blood Products And Intravenous Fluids Before Administration To Reduce Complications Associated With The Infusion Of Cold Solutions: (a) Product #4r4305 - 115 Volts (u.s. Distribution); (b) Product #4r4306 - 220 Volts (international Distribution). Recall #z-1094/1095-4. | Baxter Healthcare Corporation Deerfield Illinois | Panlmatic Company Elk Grove Village Illinois. | June 15 1994 | There is a potential for electrical shock if the wiring were to come loose and contact the electrical magnetic interference (emi) coating inside the housing | An Estimated 3 500 Units Are Affected And Remain On The Market. These Affected Units Were Distributed Between May 1985 And September 1987. | Nationwide Canada Colombia Brazil Venezuela United Kingdom.
1994-07-30 | Devices | II | z-1099/1102-4 | Ccs Maestro Ii Models 231 232 325 And 333 Cardiac Pacemakers. Recall #z-1099/1102-4. | Manufacturer | Cardiac Control Systems Inc. Palm Coast Florida. | March 25, 1994 | An intended future safety feature included in the hybrid circuitry of the present device design can be spontaneously activated in some units, causing the device to act as an asynchronous single-chamber pacer | 416 Pacers Were Distributed Domestically. Of Those 284 Were Implanted And 132 Were Recovered (per Final Report). A Total Of 574 Pacers Were Distributed Internationally And All Were Implanted. | Arkansas California Connecticut Florida Louisiana Massachusetts Maine Montana North Carolina New Jersey New York Ohio Pennsylvania Rhode Island South Carolina Tennessee Texas Wisconsin. -6-
1994-07-30 | Devices | III | z-956-4 | Beta Lysin Reagent 5 Ml A Liquid Microbial Reagent For The Identification Of Strep. Agalactiae. Recall #z-956-4. | Manufacturer | Remel Limited Partnership Lenexa Kansas. | January 20, 1994 | The expected positive qc reaction, an enhanced crescent of beta hemolysis with strep | 15 Units Were Distributed From 10/11/93 To 1/20/94. | Colorado Illinois Oklahoma.
1994-07-30 | Devices | III | z-957-4 | Mcfarland Equivalence Turbidity Standards 4.0 Used As Standards In Adjusting Densities Of Bacterial Suspensions. Recall #z-957-4. | Remel Limited Partnership Lenexa Kansas | Scientific Device Lab. Glenview Illinois. | October 20, 1993 | The individual tubes were incorrectly labeled | 37 Units Were Distributed. | Oklahoma Texas Florida New Jersey Washington State Michigan Kansas Missouri Illinois Colorado New Mexico New York Nebraska Guam Canada.
1994-07-30 | Devices | III | z-958-4 | Piperacilin 100 Mcg Susceptibility Test Disks For In-vitro Diagnostic Use. Recall #z-958-4. | Manufacturer | Remel Limited Partnership Lenexa Kansas. | May 7, 1993 | The individual cartridge labeled piperacillin was incorrectly labeled | Approximately 386 Units Were Distributed; Firm Estimates None Remains On The Market. | Nationwide.
1994-07-30 | Devices | III | z-959-4 | Virocult Ent In-vitro Diagnostic Collection And Transport System. Recall #z-959-4. | Remel Limited Partnership Lenexa Kansas | Medical Wiring & Equipment Sparta New Jersey. | January 20, 1994 | The product had an incorrect expiration date printed on the label of the outer container | 8 Units Of 125 Swabs Were Distributed. | New York Ohio Kansas.
1994-07-30 | Devices | III | z-960-4 | Remel Rpr Liquid Controls To Be Used With The Remel Rpr Card Test Kit. Recall #z-960-4. | Remel Limited Partnership Lenexa Kansas | Ampcor Camden New Jersey. | May 27, 1993 | The reactive control contained black specs and turbidity | 20 Units Were Distributed. | Illinois Missouri Iowa Tennessee Michigan Alabama.
1994-07-30 | Devices | III | z-961-4 | Contact I Sabouraud Dextrose Agar With Polysorbate 80 With Lecithin Used In The Detection And Enumeration Of Microorganisms Present On Surfaces Of Sanitary Importance. Recall #z-961-4. | Manufacturer | Remel Limited Partnership Lenexa Kansas. | October 19, 1993 | The product was incorrectly labeled as tsa instead of sab- dex | 500 Plates Were Distributed; Firm Estimates None Remains On The Market. | California.
1994-07-30 | Devices | III | z-1079-4 | Linvatec Cannuloc Threaded Cannula Used During Orthoscopic Surgery To Introduce Various Surgical Instruments Into The Surgical Site. Recall #z-1079-4. | Manufacturer | Linvatec Division Of Bristol-myers Squibb Largo Florida. | April 6, 1994 | 75 mm cannuloc threaded cannulas were labeled as 50 mm cannuloc threaded cannulas | 24 Boxes Of 10 Were Distributed. | Georgia Massachusetts Minnesota Maine Denmark Canada.
1994-07-30 | Devices | III | z-1083-4 | Wall Mount Brackets Abs Construction Part #008-0046-00 An Accessory Item For Propaq Patient Monitor Models 102 104 And 106. Recall #z-1083-4. | Manufacturer | Protocol Systems Inc. Beaverton Oregon. | January 19, 1993 | An environment of high vibration or severe force could cause fracture of the plastic portion of the bracket | 196 Units Were Distributed. | Nationwide And International.