## 1994-08-24
RECALL_CLASSIFICATION_DATE | PRODUCT_TYPE | CLASSIFICATION | RECALL_NUMBER | PRODUCT | RECALLING_FIRM | MANUFACTURER | RECALL_INITIALIZATION_DATE | REASON | VOLUME | DISTRIBUTION
1994-08-24 | Drugs | II | d-347-4 | Acephen Acetaminophen Suppositories Usp (pediatric) 120 Mg Foil Wrapped For Rectal Administration A Nonprescription Product. Recall #d-347-4. | Manufacturer | G & W Laboratories Inc. South Plainfield New Jersey. | April 8, 1994 | Product does not meet dissolution specifications | 23 370 Boxes Of 12 Were Distributed. | Nationwide.
1994-08-24 | Drugs | II | d-348-4 | Miacalcin (calcitonin-salmon) Injection Synthetic In 2 Ml Multi-dose Vials Rx Indicated For Hypercalcemia Post- Menopausal Osteoporosis And Paget's Disease. Recall #d-348-4. | Sandoz Pharmaceuticals Corporation East Hanover New Jersey | Schering-plough Corporation Manati Puerto Rico. | March 25, 1994 | Some vials were found to be unlabeled | 42 014 Vials Were Distributed. | Nationwide.
1994-08-24 | Drugs | II | d-349-4 | Benzonatate 100 Mg Usp Soft Gelatin Capsules In Bottles Of 100 Rx Indicated For The Symptomatic Relief Of Cough Under The Following Labels: Aligen Martec H.l. Moore Qualitest Sidmak U.r. L. Recall #d-349-4. -4- | Manufacturer | Banner Pharmacaps Inc. Elizabeth New Jersey. | July 13, 1994 | Product does not meet dissolution specifications | Undetermined. | New Jersey Wyoming Louisiana.
1994-08-24 | Drugs | II | d-350-4 | Clozaril Clozapine 25 Mg Tablets In Blister Packaging Of 100 (sandopaks) Rx Schizophrenia Treatment. Recall #d-350-4. | Manufacturer | Sandoz Pharmaceutical Corporation East Hanover New Jersey. | March 25, 1994 | Some of the unit dose blisters had no printing | 6 239 Sandopaks Were Distributed. | Nationwide Puerto Rico.
1994-08-24 | Drugs | III | d-351-4 | Triamterene And Hydrochlorothiazide Capsules 50 Mg In Bottles Of 100 And 1 000 An Antihypertensive. Recall #d-351-4. | Rugby Laboratories Norcross Georgia | Penn Labs Inc. Subsidiary Of Smithkline Beecham Cidra Puerto Rico. | July 5, 1994 | Some lots do not meet dissolution specification through the expiration date | Approximately 297 Million Capsules Were Manufactured. | Undetermined.
1994-08-24 | Drugs | III | d-346-4 | Sulfamethoxazole 400 Mg And Trimethoprim 80 Mg Tablets In Boxes Of 100 Rx Antibiotic. Recall #d-346-4. | Auro Pharmaceuticals Inc. Williamsville New York (repacker/responsible Firm) | Lemmon Company Sellersville Pennsylvania. | August 8, 1994 | Presence of a yellow powder in a blister strip | 361 Boxes Of 100 Tablets Each Were Distributed Between 3/4/93 And 6/28/93. Firm Estimated That 96 Boxes Remained On Market At Time Of Recall. | California Florida Indiana Michigan New York Pennsylvania. -5-
1994-08-24 | Biologics | II | b-431/433-4 | (a) Red Blood Cells; (b) Platelets; (c) Source Leukocytes. Recall #b-431/433-4. | Manufacturer | Civitan Regional Blood Center Inc. Gainesville Florida. | December 14, 1993 | Blood products that either tested repeatedly reactive for antibodies to human immunodeficiency virus types 1 and 2 (anti-hiv-1/2) or tested repeatedly reactive for the antibody to hepatitis b core antigen (anti-hbc) were distributed | (a) 1 Unit; (b) 3 Units; (c) 1 Unit. | Alabama Ohio Florida.
1994-08-24 | Devices | II | z-115/116-4 | Zoll Defibrillators Used For The Purposes Of Converting Ventricular Fibrillation (vf) A Cardiac Rhythm Incompatible With Life To Sinus Rhythm Or Other Cardiac Rhythms Capable Of Producing Hemodynamically Significant Heart Beats: (a) Zoll Pd 1200 Pacemaker/defibrillator; (b) Zoll D 900. Recall #z-115/116-4. | Manufacturer | Zoll Medical Corporation Burlington Massachusetts. | June 10, 1994 | Depleted or damaged batteries can compromise operation of the defibrillator | (a) 7 090 Units; (b) 1 635 Units Were Distributed. | Nationwide And International.
1994-08-24 | Devices | II | z-548-4 | Electrode Lead Wires 2.0 Mm Which Have Exposed Metal Pins. These Lead Wires Were Sold Under Nortech's And Empi's Brand Names And Are Used With Tens And Nmes (neuro-muscular Electrical Stimulation) Devices. The Product Numbers Of The Nortech Lead Wires Are 86602136 86604000 86604018 86604036 86604037 86604048 86604060 86604136 86604148 86604160. The Product Numbers Of The Empi Lead Wires Are 193012 193013 198060 198061 198086 198087 198136. Recall #z-548-4. -6- | Manufacturer | Empi Inc. St. Paul Minnesota. | July 7, 1994 | Device labeling fails to provide adequate directions for use and the devices pose an unwarranted risk of injury to patients, because the products are not labeled to contraindicate use with apnea monitors | 4 099 Empi Brand Lead Wires And 5 137 Nortech Brand Lead Wires Were Distributed. | Nationwide Mexico.
1994-08-24 | Devices | II | z-1151/1152-4 | Simplastin L Thromboplastin Reagent In-vitro Diagnostic Reagent Intended For Use In The Determination Of Prothrombin Time: (a) Product #59552 300 Tests; (b) Product #59553 1000 Tests. Recall #z-1151/1152-4. | Manufacturer | Organon Teknika Corporation Durham North Carolina. | November 19, 1993 | The reagent may give unusually prolonged prothrombin times in some coumadin patient specimens with inr's greater than 3 | 9 222 Kits Were Distributed. | Nationwide The Netherlands Canada.
1994-08-24 | Devices | II | z-1202-4 | Reaction Cuvettes Model 751060 Used To Hold Specimen Samples For Chemical Analyses Performed Using The Firm's Express 550 And Express Plus Spectrophotometrs. Recall #z-1202-4. | Ciba Corning Diagnostics Corporation Oberlin Ohio | Vermont Plastics Inc. Montpelier Vermont (contract Manufacturer). | February 24, 1993 | The device was found to have been manufactured with square corners instead of the customary rounded corners, which may cause reagent carry-over, potentially effecting test results | 593 Boxes Of 100 Magazines Were Distributed. Each Magazine Holds 13 Cuvettes Each Cuvette Consists Of 5 Wells. | Nationwide.
1994-08-24 | Devices | II | z-1228-4 | Sunbronze Cosmos 85/2 Tanning Beds. Recall #z-1228-4. | Manufacturer. Fda Approved The Firm's Corrective Action Plan July 1 1994. Firm-initiated Field Correction Ongoing. | European Sun Systems St. Louis Missouri. | July 1, 1994 | The remanufactured sunlamp products failed to incorporate timers and their quality control failed to indicate a check for installation of compliant timers | 10 Units. | Texas.
1994-08-24 | Devices | II | z-1234/1236-4 | Air-shields Transilluminators Used To Deliver Visible Light Through The Body Tissue Of Infants In The Placement Of I.v.'s And In The Diagnosis Of Pneumothorax: (a) Part #lt90700 (120); (b) Part #lt91500 (220-240v); (c) Part #lt96000 (100v). Recall #z-1234/1236-4. | Lumitrex Inc. Strongsville Ohio | Endura Plastics Kirtland Ohio (component Supplier). | August 19, 1993 | A plastic clip inside the transilluminator was manufactured utilizing an incomplete drying process | 100 Units Were Distributed; Firm Estimates That 29 Units Remained On Market At Time Of Recall. | Pennsylvania.
1994-08-24 | Devices | II | z-1239-4 | Sarns 8000 Line Occluder Part #16418 Indicated To Control The Flow Of Blood Through The Venous Line. Recall #z-1239-4. | Manufacturer | Sarns 3m Health Care Ann Arbor Michigan. | October 22, 1993 | The knob used to control gradual blood flow was incorrectly designed, resulting in inadequate control of blood flow in some units | 39 Units Were Distributed. | California Delaware Indiana Kentucky Massachusetts Michigan New Jersey New York North Carolina Ohio Pennsylvania South Carolina Tennessee Washington State Washington D.c. Singapore Hong Kong Canada Germany Japan Australia Scotland Italy.
1994-08-24 | Devices | II | z-1240/1241-4 | Sarns Air Sensors Air Bubble Detectors: (a) Part #13540 3/8"; (b) Part #15020 1/4". Recall #z-1240/1241-4. | Manufacturer | Sarns 3m Health Care Ann Arbor Michigan. | April 14, 1993 | An electrical short may occur in some units rendering these units nonfunctional | (a) 273 Units; (b) 46 Units Were Distributed. | Nationwide And International.
1994-08-24 | Devices | II | z-1258/1259-4 | Report Master Ii Version 2.0c Software Data Diskette For Gemstar Ii Chemistry System. A Software Program That Charts Flags Calculates And Stores Blood Test Chemistry Results: (a) Part #402275; (b) Part #402276. Recall #z-1258/1259-4. | Manufacturer | Schiapparelli Biosystems Inc. Fairfield New Jersey. | November 11, 1993 | A false bun/creatine ratio may be calculated when either a bun or creatine value has not been entered by the user | 87 Copies Were Distributed. | Nationwide.
1994-08-24 | Devices | II | z-1263-4 | Zoll Medical Corporation Internal Paddles Used For Defibrillation During Open Heart Surgery. Recall #z-1263-4. | Manufacturer | Zoll Medical Corporation Burlington Massachusetts. | February 27, 1992 | A situation could allow an improperly routed wire to shorten pins 9 and 10, causing the unit to fail | 10 Units Were Distributed. | Illinois Kentucky Tennessee Texas.
1994-08-24 | Devices | II | z-1271-4 | Model 1000 Ophthalmic Laser Workstation Used In Ophthalmic Surgery. Recall #z-1271-4. | Manufacturer. Fda Approved The Firm's Corrective Plan May 26 1994. Firm-initiated Recall Complete. | Phoenix Laser Systems Inc. Fremont California. | May 26, 1994 | The device was found to lack an emission indicator for the aiming beam, the warning logotype label carried inaccurate hene output information, and the operator's manual lacked calibration procedures | 2 Units. | Florida California.
1994-08-24 | Devices | II | z-1300-4 | Nd:yag Ophthalmic Laser System Used In Posterior Capsulotomies. Recall #z-1300-4. -9- | Manufacturer. Fda Approved The Firm's Corrective Action Plan June 27 1994. Firm-initiated Field Correction Ongoing. | Laserex Systems Adelaide S.a. Australia. | June 27, 1994 | The joystick aiming beam control in the slitlamp was poorly seated causing internal electrical wires to become prematurely worn and broken | 103 Units. | Nationwide And International.
1994-08-24 | Devices | III | z-1260-4 | Abbott Test Pac Rsv Enzyme Immunoassay For The Rapid Detection Of Respiratory Syncytial Virus From Nasopharyngeal Washes Aspirates And Swabs. Recall #z-1260-4. | Manufacturer | Abbott Laboratories Abbott Park Illinois. | November 17, 1993 | The rsv reaction discs in the kits could provide false positive results if the specimen is transported in saline | 285 Kits Were Distributed; Firm Estimates None Remains On The Market. | Nationwide Canada Chile Italy Japan Taiwan Germany.
1994-08-24 | Devices | III | z-1264/1266-4 | Immunocytochemical Assay In-vitro Diagnostic Kits Which All Contain 6 Ml Vials Of Pap Complex [peroxidase (horseradish): Anti-peroxidase (rat) Complex] From The Same Bulk Lot: (a) Abbott Er-ica Monoclonal Immunocytochemical Assay For The Detection Of Human Estrogen Receptor In Human Breast Tissue Used As An Aid In Predicting Patient Response To Hormonal Therapy And In The Prognosis And Management Of Breast Cancer Patients List #3087-18 30 Test Kit; (b) Abbott Pgr-ica Monoclonal Immunocytochemical Assay For The Detection Of Human Progesterone Receptor In Human Breast Tissue Used As An Aid In Predicting Patient Response To Hormonal Therapy And In The Management Of Breast Cancer Patients List #2a08-18 30 Test Kit; -10- (c) Abbott Pgr-ica Monoclonal For Investigational Use Only; Immunocytochemical Assay For The Detection Of Human Progesterone Receptor In Human Breast Tissue List #6a74-18 30 Test Kit. Recall #z-1264/1266-4. | Manufacturer | Abbott Laboratories Abbott Park Illinois. | February 23, 1994 | Certain er-ica test kits were found to have flocculent material in the pap complex and certain pgr-ica test kits were found to have microbial contamination in the pap complex | (a) 557 Kits; (b) 396 Kits; (c) 73 Kits Were Distributed. | Nationwide And International.
1994-08-24 | Devices | III | z-1267-4 | Kodak Ektachem Performance Verifier Ii And Diluent Set Ii Catalog #8231474 A Lyophilized Solution Used For Quality Control Of The Ektachem Analyzers Marketed By Kodak. Recall #z-1267-4. | Manufacturer | Eastman Kodak Company Rochester New York. | March 30, 1994 | Some bottles of performance verifier i are labeled as bottles of performance ii | 2 941 Boxes Were Distributed Between 2/14/94 And 3/28/94. | Nationwide And Canada.
1994-08-24 | Devices | III | z-1268/1270-4 | Hdl Cholesterol Kits Used In Conjunction With The Hdlc Slide To Quantitatively Measure High Density Lipoprotein Cholesterol: (a) Kodak Ektachem Hdl Cholesterol Cat 146 3256 Contains 100 Tubes Plus 1 Bottle Diluent And Is Used With The Ektachem Mainframe Analyzer; (b) Ektachem Dt Hdl Cholesterol Kit Cat 148 0664 Contains 25 Slides 27 Macro Tubes And One Bottle Of Diluent Used With The Desk Top Ektachem Analyzer; (c) Ektachem Dt Micro Hdl Cholesterol Kit Cat 133 5504 Contains 25 Slides 27 Micro Tubes And One Bottle Of Diluent Used With The Desk Top Ektachem Analyzer. Recall #z-1268/1270-4. | Eastman Kodak Company Clinical Diagnostics Division Rochester New York | Eastman Kodak Company Rochester New York. | October 19, 1993 | The tubes are filled with a small quantity of dextran sulfate which sticks to the top of the test tube cap | Approximately 14 500 Kits Used On The Mainframe Ektachem Analyzers And 14 500 Kits Used On The Desk Top Models Were Distributed Between August And November 1993. | Nationwide And Canada. -11-