## 1994-09-21
RECALL_CLASSIFICATION_DATE | PRODUCT_TYPE | CLASSIFICATION | RECALL_NUMBER | PRODUCT | RECALLING_FIRM | MANUFACTURER | RECALL_INITIALIZATION_DATE | REASON | VOLUME | DISTRIBUTION
1994-09-21 | Drugs | II | d-381/382-4 | Albuterol Sulfate Tablets Usp (a) 2 Mg; (b) 4 Mg In Bottles Of 100 For Asthma Relief. Recall #d-381/382-4. | Warner Chilcott Laboratories Morris Plains New Jersey | Warner Chilcott Laboratories Division Of Warner Lambert Company Lititz Pennsylvania. | July 15, 1994 | Possible discrepancies in anda data | Approximately 48 838 2-mg Bottles And 37 000 4-mg Bottles Were Distributed. | Nationwide.
1994-09-21 | Drugs | III | d-383/384-4 | Augmentin Chewable Tablets An Oral Antibacterial Combination: (a) Augmentin Chewable Tablets (125 Mg. Amoxicillin And 31.25 Mg Clavulanate Potassium) Foil Strips 3 Tablet Size (physician Samples) And 30 Tablet Size; (b) Augmentin Chewable Tablets 250 Mg Amoxicillin And 62.5 Mg Clavulanate Potassium Foil Strips 3 Tablet Size (physician Samples) And 30 Tablet Size. Recall #d-383/384-4. | Manufacturer | Smithkline Beecham Pharmaceuticals Bristol Tennessee. | August 31, 1994 | Marginal subpotency of the clavulanate potassium ingredient | 105 631 Of 30 Tablets Of The 125 Mg/31.25 Mg Product And 464 864 Packages Of 30 Tablets Of The 250 Mg/62.5 Product Were Distributed. The Quantity Of Physician Samples Distributed Is Undetermined. -4- | Nationwide.
1994-09-21 | Biologics | II | b-461-4 | Source Plasma. Recall #b-461-4. | Manufacturer | Am-rho Laboratories Inc. Jacksonville Florida. | October 29, 1993 | Blood products which tested repeatedly reactive for the hepatitis b surface antigen (hbsag); or tested negative, but were collected from a donor who previously tested repeatedly reactive for hbsag, were distributed | 17 Units. | New Jersey.
1994-09-21 | Biologics | III | b-449/450-4 | (a) Red Blood Cells; (b) Recovered Plasma. Recall #b-449/450-4. | Manufacturer (a) | The American National Red Cross Wichita Falls Texas. | March 26, 1994 | Blood products, collected in anticoagulant citrate phosphate dextrose adenine solution (cpda-1 bags) in which the outside was contaminated with mold, were distributed | (a) 91 Units; (b) 141 Units. | (a) Texas Oklahoma; (b) California.
1994-09-21 | Biologics | III | b-460-4 | Platelets. Recall #b-460-4. | Manufacturer | The American National Red Cross Toledo Ohio. | January 7, 1994 | Blood products, incorrectly tested for the antibody to the human immunodeficiency virus type 1/2 (anti-hiv-1/2), were distributed | 10 Units. | Maryland Massachusetts Ohio Washington D.c.
1994-09-21 | Devices | I | z-1154/1157-4 | Microscan Dried Gram Positive Panels Microtiter Plates With Dried Antimicrobial Dilutions And Supplemented Broth That Are Rehydrated With Suspensions Of Test Organisms: (a) Microscan Positive Combo Type 6; (b) Microscan Positive Mic Type 6; (c) Microscan Positive Mic Type 8; (d) Microscan Positive Breakpoint Combo Type 6. Recall #z-1154/1157-4. -6- | Manufacturer | Baxter Diagnostics Inc. Microscan West Sacramento California. | May 31, 1994 | The panels do not detect penicillin resistant strains of streptococcus pneumoniae which can result in serious health consequences, including death, if it is not detected | Undetermined. | Nationwide And International.
1994-09-21 | Devices | II | z-1272-4 | Howmedica P.c.a. Hip System Uhmwpe Acetabular Insert Size # 28 Mm I.d. 10o Hood Fits 64 Or 67 Mm Outer Shell Catalog #6283-6-647 Used To Provide The Articulating Surface For The Femoral Head Following Total Hip Replacement. Recall #z-1272-4. | Manufacturer | Howmedica Inc. Rutherford New Jersey. | April 13, 1994 | The device is labeled as a 28 mm i | 59 Units Were Distributed; Firm Estimated That 3 Units Remained On Market At Time Of Recall. | Nationwide Canada Ireland.
1994-09-21 | Devices | II | z-1282/1299-4 | Povidone Iodine (pi) Solution Packaged In Kendall Curity Anesthesia Trays: 1. Curity Spinal Anesthesia Tray Sterile Reorder #4705 2. Curity Spinal Anesthesia Tray Sterile Reorder #4716 3. Curity Spinal Anesthesia Tray Sterile Reorder #4720 4. Curity Spinal Anesthesia Tray Sterile Reorder #4722 5. Curity Spinal Anesthesia Tray Sterile Reorder #4724 6. Curity Spinal Anesthesia Tray Sterile Reorder #4727 7. Curity Spinal Anesthesia Tray Sterile Reorder #4730 8. Curity Saddle Block Tray Sterile Reorder #4732 9. Curity Saddle Block Tray Sterile Reorder #4740 10. Curity Saddle Block Tray Sterile Reorder #4750 11. Curity Saddle Block Spinal Anesthesia Tray Sterile Reorder #4763 12. Curity Saddle Block Spinal Anesthesia Tray Sterile Reorder #4787 13. Curity Saddle Block Spinal Anesthesia Tray Sterile Reorder #4795 14. Curity Spinal Anesthesia Tray Sterile Reorder #4920 15. Curity Single Shot Epidural Anesthesia Tray Sterile Reorder #5070 -7- 16. Curity Continuous Epidural Anesthesia Tray Sterile Reorder #5181 17. Curity Continuous Epidural Anesthesia Tray Sterile Reorder #5197 18. Curity Single Shot Epidural Anesthesia Tray Sterile Reorder #5554. Recall #z-1282/1299-4. | Kendall Healthcare Products Company Mansfield Massachusetts | Kendall Healthcare Products Augusta Georgia. | June 8, 1994 | Anesthesia trays contain packets of povidone iodine solution which may be subpotent | 2 053 Cases Were Distributed. | Nationwide.
1994-09-21 | Devices | II | z-1316/1317-4 | Feeding Tubes 5 Fr 15 Inches Long Sterile Single Patient Use Rx Device Used For Naso-gastric Enteral Feeding Of Infants: (a) Baxter Feeding Tube Catalog #k32; (b) Baxter Infant Feeding Tube With Radiopaque Stripe Catalog #r32. Recall #z-1316/1317-4. | Baxter Healthcare Corporation Round Lake Illinois | Baxter S.a. Cartago Costa Rica. | July 21, 1994 | Solvent bond failure between the tubing and the luer connector may cause leakage | 251 670 Tubes Were Distributed; Firm Estimated That 15 Percent Of The Product Remained On The Market At Time Of Recall. | Nationwide Canada Singapore Puerto Rico.
1994-09-21 | Devices | II | z-1375/1377-4 | Puritan-bennett Microprocessor Ventilator 7200ae Series Enables A Respiratory Care Practitioner To Provide A Patient With Ventilatory Assistance: (a) Model 7200ae Ventilator; (b) Model 7200spe Ventilator; (c) Model 7200e Ventilator. Recall #z-1375/1377-4. | Manufacturer | Puritan-bennett Corporation Carlsbad California. | June 29, 1994 | An intermittent wiring/pin connection in the pneumatic wire harness assembly could cause the ventilators to go into safety valve open (svo) position, resulting in bagging of the patient and a ventilator change | 29 Units Were Distributed; No Defective Units Remain In Commerce. | Arkansas Texas Florida Alabama Massachusetts Connecticut California Georgia Colorado New York.
1994-09-21 | Devices | II | z-1385/1386-4 | Endo-p Transducer A Diagnostic Ultrasound Transducer Compatible With The Sonoline Series Ultrasound System (si200/250 Si400/450/450d): (a) Part #874000; (b) Part #5725270. Recall #z-1385/1386-4. | Siemens Medical Systems Inc. Ultrasound Group Issaquah Washington | Siemens Ag Ub Med. Erlangen Germany. | May 20, 1994 | The transducers fail to properly change from axial to longitudinal scanning planes, resulting in misguided biopsy punctures | 283 Transducers Were Distributed. | Nationwide And International.
1994-09-21 | Devices | II | z-1379/1382-4 | (a) Endoflex Lumbar Discectomy Kit Catalog #2-6080 Intended For Lumbar Disc Herniations; (b) 3.5 Mm Endoprobe Catalog #2-6088 The Cutting Instrument For The Removal Of Nucleus Materials Sold Individually Or As Part Of A Kit; (c) 3 Mm Flexible Ring Curette Catalog #5-1020 Manually Used To Decompress A Herniated Nucleus Pulposus During Microdiscectomy And Laminectomy Surgery; (d) 3 Mm Flexible Rectangular Curette Catalog #5-1030 Manually Used To Decompress A Herniated Nucleus Pulposus During Microdiscectomy And Laminectomy Surgery. Recall #z-1379/1382-4. | Manufacturer | Surgical Dynamics Inc. Concord California. -9- | May 16, 1994 | The sterilization parameters used by the contract sterilizer had not been validated thereby compromising the sterility of the devices | 547 Units Were Distributed. | Nationwide Australia Italy Peoples Republic Of China Spain South Korea.
1994-09-21 | Devices | II | z-1394-4 | Cosgrove-edwards Annuloplasty System Model 4600 Indicated For Mitral Valvuloplasty. Recall #z-1394-4. | Manufacturer | Edwards Cvs Division Baxter Healthcare Corporation Irvine California. | June 29, 1994 | Sterility may be compromised due to defective packaging seals | 419 Units Were Distributed; Firm Estimates None Remains On The Market. | California Massachusetts Michigan Minnesota Missouri Ohio.
1994-09-21 | Devices | II | z-1405-4 | Baxter Univox Membrane Oxygenation System Used In Cardiopulmonary Bypass Surgery For 02/c02 Exchange. Recall #z-1405-4. | Baxter Healthcare Corporation Irvine California | Baxter Healthcare Corporation Anasco Puerto Rico. | June 3, 1994 | The uv adhesive bond may detach between the gas manifold and the oxygenator housing, allowing oxygen to be released into the environment instead of replacing the carbon dioxide in the blood returning to the patient | 11 955 Systems Were Distributed. | Nationwide And International.
1994-09-21 | Devices | III | z-1357-4 | Aldetect Formaldehyde Test Reagent In 2 Cc Glass Ampules An In-vitro Diagnostic Used To Detect Aldehydes In Dialysis Devices. Recall #z-1357-4. | Manufacturer Telephone February 2 And 3 1994. Firm- Initiated Recall Complete. | Pochemco Inc. Chicopee Massachusetts. | February 2, 1994 | The device failed finished product testing, and the expiration date was incorrectly labeled as 2 years instead of 1 year | Firm Estimates That Little If Any Product Remains On Market. -10- | Massachusetts New York New Jersey Ohio.
1994-09-21 | Devices | III | z-1363-4 | Microaire Surgical Oscillating Sawblade Part #mo-124 Power Instruments For Orthopedic Surgery. Recall #z-1363-4. | Manufacturer | Microaire Surgical Instruments Inc. Valencia California. | June 20, 1994 | Mislabeled. | 99 Blades Were Distributed. | Arizona California Florida Illinois Minnesota New Mexico.
1994-09-21 | Devices | III | z-1378-4 | Howmedica Centrax Biopolar Prosthesis The Acetabular Component Of Hip Implant Used In Orthopedic Surgery Implant Procedures. Recall #z-1378-4. | Manufacturer | Howmedica Inc. Rutherford New Jersey. | April 27, 1994 | Product was incorrectly packaged in a pre-printed carton identifying the device as a precision hip system sterile acetabular component | 6 Units Were Distributed; Firm Estimates None Remains On The Market. | California New York Florida Wisconsin.