## 1994-11-09
RECALL_CLASSIFICATION_DATE | PRODUCT_TYPE | CLASSIFICATION | RECALL_NUMBER | PRODUCT | RECALLING_FIRM | MANUFACTURER | RECALL_INITIALIZATION_DATE | REASON | VOLUME | DISTRIBUTION
1994-11-09 | Biologics | II | b-027-5 | Red Blood Cells. Recall #b-027-5. | Manufacturer | The American National Red Cross Detroit Michigan. | February 21, 1993 | Blood products, corresponding to units of platelets which were implicated in a case of post-transfusion serratia marcescens septicemia, were distributed | 3 Units. | Michigan.
1994-11-09 | Biologics | II | b-028/030-5 | (a) Red Blood Cells; (b) Recovered Plasma; (c) Recovered Plasma Pooled. Recall #b-028/030-5. | Manufacturer | Lane Memorial Blood Bank Eugene Oregon. | January 4, 1994 | Blood products, collected from an ineligible donor due to medication with human pituitary-derived growth hormone, were distributed | (a) 2 Units; (b) 1 Unit; (c) 1 Unit. | Oregon California Florida.
1994-11-09 | Biologics | II | b-031/032-5 | (a) Platelets For Further Manufacture; (b) Recovered Plasma. Recall #b-031/032-5. | Manufacturer | Lane Memorial Blood Bank Eugene Oregon. | April 18, 1994 | Blood products, that tested initially reactive for the antibody to hepatitis c virus encoded antigen (anti-hcv), were distributed | 1 Unit Of Each Component. | Florida.
1994-11-09 | Biologics | II | b-033/034-5 | (a) Platelets For Further Manufacture; (b) Recovered Plasma. Recall #b-033/034-5. | Manufacturer | Sarasota Community Blood Bank Inc. Sarasota Florida. | July 7, 1994 | Blood products that tested nonreactive for antibody to the human immunodeficiency virus type 1 (anti-hiv-1), but were collected from a donor who previously tested repeatedly reactive for anti-hiv-1, were distributed | 1 Unit Of Each Component. | Florida.
1994-11-09 | Devices | II | z-1429/1432-4 | Elastimide Silicone/polymide Posterior Chamber Intraocular Lenses Optic Diameter 6.30 Mm Indicated For The Visual Correction Of Aphakia: (a) Model Aq1016 J Loop 10 Degree Angle Length 13.5mm (b) Model Aq2003 C Loop Zero Degree Angle Length 12.5mm (c) Model Aq2013 C Loop 10 Degree Angle Length 12.5mm (d) Model Aq2010 C Loop 10 Degree Angle Length 13.5mm. Recall #z-1429/1432-4. | Manufacturer | Staar Surgical Company Inc. Monrovia California. | May 27, 1994 | Mislabeled. | 35 034 Lenses Were Distributed; Firm Estimates None Remains On Market. | Nationwide And International.
1994-11-09 | Devices | II | z-026-5 | Is Labeled With An Incorrect Optic Diameter 6.00 Mm Instead Of 6.30 Mm. _______________ Product 1000 Plane Sorter Circuit Board For The Ecat Scanner #953. Recall #z-026-5. | Manufacturer November 12 1993. Firm-initiated Recall Ongoing. | Cti Pet Systems Inc. Knoxville Tennessee. | November 12, 1993 | A component not permitted by the manufacturers specifications (115 meg instead of 185 meg) was used in the manufacture of the 1000 plane sorter boards | 2 Circuit Boards Were Distributed. | Sweden Germany.
1994-11-09 | Devices | II | z-041/044-5 | Bird Volume Ventilators With Flow Control Valve Cams Indicated For The Delivery Of Medical Gases To Adult And Pediatric Patients: (a) Model 09740; (b) Model 09611; (c) 15020; (d) 15336. Recall #z-041/044-5. -4- | Manufacturer | Bird Products Corporation Palm Springs California. | August 1, 1994 | The valve cam, which lacked chromium oxide plating, can wear and cause the ventilator to deliver higher than set volumes | Approximately 2 000 Valves Were Distributed. | Nationwide And International.
1994-11-09 | Devices | II | z-049-5 | Laser Pattern Generator Model Delta 200 Used To Create And Project Graphic Patterns Of Red Light. Recall #z-049-5. | Manufacturer. Fda Approved The Firm's Corrective Action Plan October 7 1994. Firm-initiated Field Correction Ongoing. | Laserworks Santa Ana California. | October 7, 1994 | Noncompliance with federal laser product performance standard in that product was improperly classified as class iiia; product was reported as a class iiia but could be used with a class iiib or class iv laser for which a variance is required; no variance had been issued | 4 Units. | Domestic And Foreign.
1994-11-09 | Devices | II | z-050-5 | Lacked Some Performance Features Required Of A Class Iiib Laser Product: Key Switch Remote Interlock Connector; Product Lacked Labeling Requirements; And User Instructions Lacked Adequate Instructions And Warnings For Proper Assembly And Safe Use Of The Product. _______________ Product Critikon Protective Iv Catheter Product #3055 18 Gauge X 1-1/4 Inches Sterile Used To Puncture The Skin And Vein For Iv Administration Of Fluids. Recall #z-050-5. | Critikon Inc. Tampa Florida | Critikon Inc. Southington Connecticut. | July 28, 1994 | The needle guard may not advance while in use, thereby putting the user at risk to be stuck by the needle | 3 187 Cases 4 Shelf Packs Of 50 Catheters Per Case Were Distributed. | Nationwide France England Thailand Ecuador Canada Australia. -5-
1994-11-09 | Devices | II | z-051/071-5 | Arrow Percutaneous Sheath Introducer Kits With Polyurethane Sheath: (a) Product No. Ako9801s; (b) Product No. Ak09810s; (c) Product No. Ak09810c; (d) Product No. Bd09880; (e) Product No. Ck01000; (f) Product No. Ds09850s; (g) Product No. Ic09880; (h) Product No. Mh09800; (i) Product No. H09810vks; (j) Product No. Nr09810s; (k) Product No. Ps09810s; (l) Product No. S109880se; (m) Product No. Sj09805; (n) Product No. Sv09804; (o) Product No. Ut09880; (p) Product No. Wb09810s; (q) Product No. Ck01410; (r) Product No. Ak09817s; (s) Product No. Ca09886; (t) Product No. Lj09807s; (u) Product No. Ph09806. Recall #z-051/071-5. | Arrow International Inc. Reading Pennsylvania | Arrow International Inc. Randleman North Carolina. | July 26, 1994 | A process change, file number rev-3958, was initiated to reduce the amount of scrap from shrinkage of the hub | 20 905 Units Were Distributed. | Nationwide Puerto Rico Canada Australia Holland Japan United Kingdom.
1994-11-09 | Devices | II | z-073/077-5 | Ophthalmic Surgical Kits Sterile And Non-sterile Single Use Devices Containing The Possibly Defective Component Kpe Male Connector P/n 126-2 Kits Are Identified With Lot Numbers Traveler Numbers: 8016 8017 608-88 608-93 608-96 And 609-04: (a) Phaco Pack Catalog #kpe-210 Fits All Cavitron Systems And Cam Receptacle Systems Contains 1 Irrigation Manifold 1 Aspiration Manifold 1 Drainage Bag And 1 Tray Support Cover Sterile Manufactured Under Hi-line Medical Label For Hi-line Medical; (b) Kpe 220 T Contains 1 Irrigation Manifold 1 Aspiration Manifold 1 Drainage Bag 1 Tray Support Cover 1 Irrigation Manifold And 1 Test Chamber Tip Wrench Sterile Manufactured Under Tri-star Ophthalmic Label For Tri-star Ophthalmic; (c) Part #504-c Part Of A Larger Kit Assembled By Sterile Concepts Richmond Va Contains 1 Aspiration Manifold 1 Irrigation Manifold 1 Drape 1 Drain Bag 1 Bio-wrap And Cohesive Paper; (d) Part #509-c Part Of A Larger Kit Assembled By Sterile Concepts Richmond Va Contains 1 Aspiration Manifold 1 Irrigation Manifold 1 Drain Bag 2 Test Chambers 2 Tip Sleeves 1 Wrench 2 "o" Rings 1 Pusher 1 Guide 1 Pouch And 1 Bio-wrap; (e) Part #510-c Part Of A Larger Kit Assembled By Sterile Concepts Richmond Va Contains 1 Aspiration Manifold 1 Irrigation Manifold 1 Drain Bag 2 Test Chambers 2 Tip Sleeves 1 Wrench 1 Pouch 1 Bio-wrap And Cohesive Paper. Recall #073/077-5. | Manufacturer | Ipax Inc. Englewood Colorado. | December 22, 1993 | The kpe male connectors may leak and compromise the sterility of the system and may also cause the system to fail to operate properly to adequately cool the handpiece, which could increase the risk of patient burns | 19 Cases Were Distributed; Firm Estimates None Remains On The Market. | Missouri California Virginia.
1994-11-09 | Devices | III | z-022-5 | Civco Spa (symmetrical Phased Array) Biopsy Bracket. Recall #z-022-5. | Manufacturer | Diasonics Ultrasound Milpitas California. -7- | July 13, 1994 | The original labeling on the biopsy bracket did not provide adequate directions for use | Firm Estimated That 45 Units Remained On Market At Time Of Recall. | Nationwide.
1994-11-09 | Devices | III | z-045-5 | Acuvue (etafilcon A) Disposable Contact Lenses. Recall #z-045-5. | Lens Express Inc. Deerfield Beach Florida | Vistakon Inc. A Division Of Johnson & Johnson Jacksonville Florida. | September 8, 1994 | The expiration dates on the outer carton and inner individually packed lenses do not correspond | Approximately 12 349 Lenses Were Distributed. | Nationwide And Canada.