## 1994-12-07
RECALL_CLASSIFICATION_DATE | PRODUCT_TYPE | CLASSIFICATION | RECALL_NUMBER | PRODUCT | RECALLING_FIRM | MANUFACTURER | RECALL_INITIALIZATION_DATE | REASON | VOLUME | DISTRIBUTION
1994-12-07 | Biologics | II | b-099-5 | Source Plasma. Recall #b-099-5. | Community Bio Resources Inc. Birmingham Alabama | Immuno-u.s. Inc. Columbus Georgia. | January 17, 1994 | Blood products, which tested non-reactive for hepatitis b surface antigen (hbsag), but were collected from a donor who previously tested repeatedly reactive for hbsag, were distributed | 1 Unit. | Austria.
1994-12-07 | Biologics | III | b-100/103-5 | (a) Whole Blood Cells; (b) Red Blood Cells; (c) Platelets; (d) Platelets Pheresis. Recall #b-100/103-5. | Manufacturer | Aurora Area Blood Bank Aurora Illinois. | September 25, 1994 | Blood products incorrectly tested for syphilis were distributed | (a) 2 Units; (b) 67 Units; (c) 43 Units; (d) 1 Unit. | Illinois Indiana.
1994-12-07 | Devices | II | z-151-5 | Brandtscan 900s Model 3151 High Speed Bar | Manufacturer. Fda Approved The Firm's Corrective Action Plan October 7 1994. Firm-initiated Field Correction Ongoing. | Brandt Inc. Bensalem Pennsylvania. | October 7, 1994 | Product failed to comply with the federal laser product performance standard in that it did not incorporate an interlock and failed to display all the required labeling | 11 Units Were Distributed | Nationwide.
1994-12-07 | Devices | II | z-152/159-5 | Blood Cardioplegia Delivery Sets Indicated For Mixing Cooling And Warming Of Oxygenated Blood/cardioplegic Solution And Delivering It To The Patient: (a) Model Bcd Adv1 Individually Packaged In Trays 6 Per Carton; (b) Model Bcd Adv4 Individually Packaged In Trays 6 Per Carton; (c) Model Bcd Adv C (with Shunt) Individually Packaged In Trays 6 Per Carton; (d) Model D721 (bcd Adv 1 Labeled For Dideco Italy) Individually Packaged In Trays 6 Per Carton; (e) Model D724 (bcd Adv 4 Labeled For Dideco Italy) Individually Packaged In Trays 6 Per Carton; (f) Model Mono 4116 (kit With Bcd Adv 1 And Other Devices); (g) Model Mono 3744 (kit With Bcd Adv 1 And Other Devices); (h) Model Scps 3898 (kit With Bcd Adv 1 And Other Devices). Recall #z-152/159-5. -4- | Sorin Biomedical Inc. Irvine California | Halkey Roberts Corporation St. Petersburg Florida. | October 11, 1994 | Due to a defective clamp on the crystalloid line, leakage may occur, resulting in a variation of infused potassium concentration in the crystalloid solution | 3 284 Units Were Distributed; Firm Estimated That 2 000 Units Remained On Market On 10/14/94. | Alabama California Florida Georgia Louisiana Michigan North Carolina Pennsylvania Tennessee Texas Washington State West Virginia Canada Italy Japan The Netherlands.
1994-12-07 | Devices | II | z-160-5 | Abbott Rv Pacing Lead Heparin Coated Transluminal Rv Pacing Lead For Temporary Ventricular Pacing Only. Recall #z-160-5. | Abbott Laboratories Abbott Park Illinois | Abbott Laboratories Mountain View California. | October 18, 1994 | There is a potential for the tip of the pacing lead to break off and remain in the patient | Approximately 5 000 To 6 000 Leads Were Distributed; Firm Estimated That 600 Leads Remained On The Market The Time The Recall Was Initiated. | Nationwide Taiwan Switzerland Argentina Spain The Netherlands Canada Japan Italy Puerto Rico.
1994-12-07 | Devices | III | z-101-5 | Lipase Test Reagent C 091-0192 Model P1 Used With The Model No. 91 Amylase-lipase Analysis System For In-vitro Diagnostic Use Only. Recall #z-101-5. | Bacharach Inc. Pittsburgh Pennsylvania | Electro-nucleonics Inc. Cranbury New Jersey. | July 9, 1990 | The reagents when used with the model 91 amylase-lipase analyzer were out of range after calibration | 1 089 Units Were Distributed; Firm Estimates None Remains On The Market. | Nationwide.
1994-12-07 | Devices | III | z-102-5 | Difco Durapak Media Chocolate Agar A Media Plate Used For Isolating And Cultivating Fastidious Microorganisms. Recall #z-102-5. -5- | Difco Laboratories Livonia Michigan | Difco Laboratories Sun Prairie Wisconsin. | March 5, 1993 | The plates were failing to grow hemophilus microorganisms when tested after storage for 2 to 4 months at room temperature | Firm Estimates None Remains On The Market. All Lots Of The Product Have Been Expired Since October 1993. | Nationwide Switzerland Italy Spain France Uruguay Portugal Finland Belgium.
1994-12-07 | Devices | III | z-149-5 | Epstein-barr Virus Igm Elisa Test. Recall #z-149-5. | Manufacturer | Gull Laboratories Inc. Salt Lake City Utah. | February 10, 1994 | The negative control component of the subject device did not meet the specification of s0 | 105 Units Were Distributed; Firm Estimates None Remains On The Market. | California Illinois Massachusetts Maryland Michigan Mississippi Virginia Austria Singapore South Africa Italy Switzerland Uruguay.