## 1995-03-22
RECALL_CLASSIFICATION_DATE | PRODUCT_TYPE | CLASSIFICATION | RECALL_NUMBER | PRODUCT | RECALLING_FIRM | MANUFACTURER | RECALL_INITIALIZATION_DATE | REASON | VOLUME | DISTRIBUTION
1995-03-22 | Drugs | II | d-086-5 | Liquid Oxygen For Respiratory Therapy In Size C D And E Cylinders | Manufacturer | Texas Air Supply Home Medical Equipment Inc | November 10 1994 | Current Good Manufacturing Practice Deficiencie | 500 Cylinders Were Distributed Firm Estimates None Remains On The Market | Texas
1995-03-22 | Drugs | II | d-087-5 | Medical Oxygen Cylinders Sizes A B D E H M | Manufacturer | Metrolina Welders Supply Inc | February 15 1995 | Current Good Manufacturing Practice Deficiencie | 12 Lots Were Distributed | North Carolina South Carolina
1995-03-22 | Drugs | II | d-088-5 | Oxygen Compressed Usp Compressed Medical Gas In Cylinders | Manufacturer | Valley Welding Supply Company Greensburg Pennsylvania | March 1 1995 | Current Good Manufacturing Practice Deficiencie | 131 Units Were Distributed | Pennsylvania
1995-03-22 | Drugs | III | d-085-5 | Indocyanine Green Usp A Prescription Bulk Pharmaceutical Chemical | Manufacturer | Eastman Kodak Company Rochester New York | October 18 1993 | Subpotenc | Approximately 7 000 Grams Were Distributed | Maryland New Mexico
1995-03-22 | Drugs | III | d-089-5 | Guaipax Pse Tablets Pseudophedrine Hcl 120 Mg And Guaifenesin 600 Mg Sustained Release 100 Tablets Rx Indicated For The Relief Of Nasal Congestion | Manufacturer | Eon Laboratories Manufacturing Inc | March 7 1995 | Product Does Not Meet Dissolution Specifications For The Guaifenesin Ingredien | 2 286 Units Of 100 Tablets Were Distributed Firm Estimates Very Little Of The Product Remains On The Market | Alabama California Connecticut Florida Georgia Hawaii Iowa Illinois Minnesota Missouri Mississippi North Carolina Nebraska New Jersey New York Oklahoma Ohio Oregon South Carolina Texas Puerto Rico
1995-03-22 | Drugs | III | d-090-5 | Triple Sulfa Vaginal Cream Usp Administered Intravaginally Used In The Treatment Of Vaginal Infection In 2 | Manufacturer | Clay Park Laboratories Inc | February 1 1995 | Product Does Not Meet Ph Specification | 698 338 Tubes Were Distributed | Nationwide The Virgin Islands West Indies The Bahamas Bulgaria
1995-03-22 | Drugs | III | d-091-5 | Nystatin Triamcinolone Acetonide Cream Usp 120 G 4 Ounce And 1 Pound Plastic Jars Used For The Treatment Of Topical Yeast Infection Under The Clay Park And Url Labels | Manufacturer | Clay Park Laboratories Inc | February 21 1995 | Superpotenc | Approximately 7 312 4 Ounce Jars And 901 1 Pound Jars Were Distributed | Nationwide
1995-03-22 | Drugs | III | d-092\093-5 | A Fluocinolone Acetonide Cream 0 | Manufacturer | Clay Park Laboratories Inc | March 8 1995 | Potency Not Assured Through Expiration Dat | A Approximately 11 666 Tubes B 21 800 Tubes Were Distributed | A Nationwide St
1995-03-22 | Biologics | III | b-197-5 | Red Blood Cells | Manufacturer | North Jersey Blood Center East Orange New Jersey | July 26 1994 | Blood Product Which Was Exposed To Unacceptable Shipping Temperatures Was Distributed | 1 Unit | New Jersey
1995-03-22 | Devices | II | z-397/402-5 | Baxter Pharmaseal Spinal And Saddle Block Anaesthesia Trays Containing 5 Cc Glass Syringes B D Glaspak 5293 With Metal Hub Spinal Needles Sterile Single Use Only 7 A Spinal Anesthesia Tray With Spinal Anesthetic And Spinal Needle 22 Ga Catalog 4500 C B Spinal Anesthesia Tray With Spinal Anesthetic And Spinal Needle 25 Ga Catalog 4510 C C Spinal Anesthesia Tray With Spinal Anesthetic Epinephrine And Spinal Needle 25 Ga Catalog 4511 C D Saddle Block Anesthesia Tray With Spinal Anesthetic Ephedrine Sulfate And Spinal Needle 25 Ga Catalog 4518 C E Custom Spinal Anesthesia Tray With Spinal Anesthetic And Spinal Needle 22 Ga Catalog 514504072 F Custom Anesthesia Tray With Spinal Anesthetic And Spinal Needle 25 Ga Catalog 6845141200 | Baxter Healthcare Corporation | Becton Dickinson Franklin Lakes New Jersey Syringe | October 3 1994 | There Is A Potential For Leaking To Occur At The Connection Between The 5 Cc Glass Syringe B D Glaspak 5293 And The Metal Hub Spinal Needle | 16 493 Trays Were Distributed Firm Estimated That 30 Of The Product Remained On Market At Time Of Recall Initiation | Nationwide And International
1995-03-22 | Devices | II | z-438/439-5 | Glutaraldehyde Disinfectant Sterilization Solution In 1 Gallon Containers A Baxter Omnicide 14 N | Manufacturer | Cottrell Ltd | August 16 1994 | A Quantity Of Glutaraldehyde Product Containing A Surfactant Was Inadvertently Pumped Into The Mixing Tank Dedicated To The Production Of The Subject Products Which Are Non Surfactant Glutaraldehyde Product | 5 876 Gallons Were Distributed Firm Estimated That 5 200 Remained On Market At Time Of Recall Initiation | Illinois Pennsylvania Virginia Nebraska Indiana
1995-03-22 | Devices | II | z-454-5 | Sechrist Model 2200 B Ventilator Indicated For The Delivery Of Medical Grade Air And Oxygen To Adult Pediatric Patients | Manufacturer | Sechrist Industries Inc | March 1 1994 | Due To A Defective Voltage Monitor The Product May Cease Ventilation Without Audible And Visual Alarm | 357 Units Were Distributed | Nationwide And International
1995-03-22 | Devices | II | z-496-5 | Sechrist Ventilator Model 2200 B A Microprocessor Controlled Ventilator Indicated For The Delivery Of Medical Grade Air And Oxygen To Adult Pediatric Patients | Manufacturer | Sechrist Industries Inc | February 23 1994 | Leakage Caused By A Defective Refill Valve Results In An Inaccurate Delivery Of Gas Volume And Pressur | 79 Units Were Distributed | Idaho And International
1995-03-22 | Devices | II | z-498/501-5 | Respironics Brand Bipap 30 Ventilatory Support Systems A Non Continuous Ventilator For Non Life Support Applications A Model 552006 Bipap S T D 30 B Model 552045 Bipap S T D 30 C Model 552049 Bipap S T D 30 D Model 552052 Bipap S T D 30 | Manufacturer | Respironics Inc | January 9 1995 | Under Certain Conditions The Interaction Between The Spontaneous Times Mode And The Breaths Per Minute Control Of The Affected Bipap S T D 30 And S T 30 Devices Results In The Devices Operating Outside The Product Specification | 189 Units Were Distributed | Nationwide Taiwan Spain Netherlands Kowloon France Canada The United Kingdom
1995-03-22 | Devices | II | z-497-5 | High Frequency Oscillatory Neonatal Ventilator Model 3100 A Indicated For Ventilatory Support And Treatment Of Respiratory Failure And Barotrauma In Neonates | Manufacturer | Sensormedics Corporation Yorba Linda California | November 22 1994 | A Wire Next To The Metal Amplifier Bracket May Abrade Due To Vibration Causing A Potential Electrical Shock Hazar | 625 Units Were Distributed | Nationwide Australia Canada Hong Kong Korea Mexico Netherlands Philippines Singapore Taiwan
1995-03-22 | Devices | II | z-510-5 | Sharpe Endo Assist Endoscopic Knot Pusher Product 87000 Product Size 10 Mm A Reusable Laparoscopic Device Used For The Extracorporeal Tying Of Suture Knots | Manufacturer | Sharpe Endosurgical Corporation St | February 6 1995 | Plastic Pieces Of The Sheath Plastic Tube That Covers The Laparoscopic Mechanism Of The Device Have Broken Off The Distal End Of The Tube Which Enters The Patient's Body Potentially Allowing Pieces To Fall Into A Patient's Body Cavit | 62 Units Were Distributed | Nationwide Canada England France Ireland Japan
1995-03-22 | Devices | II | z-511/515-5 | All Intervascular Inc | Manufacturer | Intervascular Inc | July 19 1994 | The Manufacturer Is Marketing The Device Without An Approved Premarket Notification Submission 510 K | 48 Units | California Illinois Pennsylvania Louisiana Tennessee Mississippi Florida And International
1995-03-22 | Devices | II | z-521/533-5 | Adult Expandable Breathing Circuits And All Products Including Various Anesthesia Kits As Components | Manufacturer | Vital Signs Totowa New Jersey | February 2 1994 | Exposure To Extremely Cold Temperatures May Cause The Expandable Tubing To Crack When Flexe | 2 225 Units Were Distributed | Alabama California Georgia Illinois Kansas Kentucky North Carolina Nebraska New York Texas Washington State Wyoming
1995-03-22 | Devices | III | z-520-5 | Beckman Synchron Cx 7 Clinical Analyzer Computer Enhanced Cx 4 Ce Cx 5 Ce Indicated For In Vitro Chemistry Testing | Manufacturer | Beckman Instruments Inc | July 14 1995 | Defective Software May Cause Patient Samples To Be Incorrectly Sampled Which May Cause Incorrect Patient Results To Be Reporte | 1 553 Units Were Distributed | Nationwide And International
1995-03-22 | Devices | III | z-534-5 | Stratus Prolactin Fluorometric Enzyme Immunoassay Kit With Code 35 Values | Manufacturer | Baxter Diagnostics Inc | September 24 1994 | The Technical Bulletin Which Is Enclosed With This Kit Has Incorrect Code 35 Values For The Stratus Ii Users Which Causes The Assay To Be Unusabl | 12 Kits Were Distributed | Texas West Virginia South Carolina Pennsylvania Ohio Nebraska Iowa Kansas Colorado Arizona Arkansas
1995-03-22 | Devices | III | z-535/538-5 | Dade Ck Mb Myoglobin Immunoassay Control Level 1 Level 2 And Level 3 A Catalog B 5200 B 5200 S Sample B Catalog B 5200 1 C Catalog B 5200 2 D Catalog B 5200 3 | Manufacturer | Baxter Diagnostics Inc | January 24 1995 | The Myoglobin Reconstituted Stability On Level Iii Maintains For Only Three Days After Reconstitutio | 3 888 Kits Were Distributed | Nationwide
1995-03-22 | Devices | III | z-539-5 | Insert States That The Reconstituted Product Is Stable For 14 Days When Stored At 2 To 8 Degrees Centigrade | Oxford Instruments Medical Systems | Oxford Instruments Ltd | November 10 1994 | A Faulty Capacitor In The Device Is Susceptible To Becoming Resistive Resulting In Inaccurate Blood Pressure Measurement | 4 Units Were Distributed | California Canada Argentina
1995-03-22 | Devices | III |  | Lusee Tlc Steerable Guidewires Designed To Provide Maximum Torque Control And Steerability Enabling The Guidewire To Transverse Tortuous Vessels And Tight Stenoses A Lusee Plc Steerable Guidewire Soft Super Soft Standard Tip Flexibilities B 0 |  | Cordis Corporation Miami Lakes Florida | N A | Tip Fractures Have Been Reported In Procedures Involving Total Occlusions Highly Tortuous Vasculature And Interferring Small Side Branche |  | Nationwide And International
1995-03-22 | Devices | III |  | Frozen Scallops 95 745 218 Et Al |  |  | N A | Of The Presence Therein Of Decomposed Scallop |  | N/A
1995-03-22 | 404 Not Found | III |  |  |  |  | N A |  |  | N/A
1995-03-22 | 404 Not Found | III |  |  |  |  | N A |  |  | N/A
1995-03-22 | 404 Not Found | III |  |  |  |  | N A |  |  | N/A
1995-03-22 | 404 Not Found | III |  |  |  |  | N A |  |  | N/A
1995-03-22 | 404 Not Found | III |  |  |  |  | N A |  |  | N/A
1995-03-22 | 404 Not Found | III |  |  |  |  | N A |  |  | N/A
1995-03-22 | 404 Not Found | III |  |  |  |  | N A |  |  | N/A
1995-03-22 | Devices | III |  | Bulk Spices |  |  | N A |  |  | N/A