## 1995-07-26
RECALL_CLASSIFICATION_DATE | PRODUCT_TYPE | CLASSIFICATION | RECALL_NUMBER | PRODUCT | RECALLING_FIRM | MANUFACTURER | RECALL_INITIALIZATION_DATE | REASON | VOLUME | DISTRIBUTION
1995-07-26 | Devices | II | z-853-5 | Bard Ptfe Vascular Grafts Thinwall Straight Intended For Use As A Bypass Conduit In The Reconstruction Of Occluded Or Diseased Blood Vessels | Bard Vascular Systems Division | Bard Vascular Systems Division Billerica Massachusetts | August 5 1994 | Longitudinal Crack In The Wal | 1 397 Units Were Distributed | Nationwide Ireland Italy
1995-07-26 | Devices | II | z-874/877-5 | Cardiac Pulse Generators Models 8220 E 8223 E 8224 8232 | Telectronics Pacing Systems | Telectronics Pacing Systems Miami Lakes Florida | April 19 1995 | Devices May Contain Defective Flex Circuits Which Can Cause An Intermittent Or Permanent No Output Conditio | 23 Units | Nationwide Canada Europe
1995-07-26 | Devices | II | z-879/881-5 | Port A Cath Dual Lumen Venous Access System With Silicone Catheter A Reorder Numbers 21 8010 And 21 8010 22 B Reorder Numbers 21 8050 And 21 8050 22 C Reorder Number 21 8052 | Manufacturer | Sims Deltec Inc | May 3 1995 | Possible Leak | 2 378 Units Were Distributed | Nationwide And International
1995-07-26 | Devices | II | z-944/945-5 | Safco Supergrade Autochuck Handpiece Headcaps Product Wphc A Component Of The 4 Hole Supergrade Plus Autochuck Dental Handpiece Product Wp 4 | Safco Dental Supply Company | Nakanishi Dental Manufacturing Company Ltd | June 16 1995 | The Headcap May Come Off The Handpiece During Us | 791 Handpieces And 383 Replacement Cartridges Were Distributed Affecting A Total Of 1 174 Headcaps | Nationwide
1995-07-26 | Devices | II | z-946-5 | Service Manuals Distributed With The Plumesafe 1202 Part Pss 1202 And Porta Plumesafe 602 Part Ppss 602 Smoke Evacuators And The Instruction Sheet Distributed With The Virosafe 12 Part Vs 12011 And Vs 12013 And Virosafe 6 Part Numbers Sing Vs 6 Vs 0606011 Vs 06013 Smoke Evacuator Filters | Manufacturer | Medetek Devices Inc | June 9 1995 | Service Manual And Instructions For These Devices Make Unsupported Claims Of Anti Microbial Propertie | 236 Lume Porta Plumesafe Devices And 88 Virosafe Filters | Connecticut Florida Georgia Illinois Iowa Michigan Minnesota Missouri Nebraska New Jersey New York North Carolina Oklahoma Pennsylvania Tennessee Texas Virginia Canada Austria Iceland Italy Malaysia
1995-07-26 | Devices | II | z-954/955-5 | Single Chamber Cardiac Pulse Generators Models 1206 And 1206 M | Telectronics Pacing Systems | Telectronics Pacing Systems Miami Lakes Florida | April 19 1995 | The Pulse Generators May Contain Defective Ceramic Substrate Pads Which Can Cause An Open Contact Condition In The Pacer Circuit And Result In Sudden Loss Of Outpu | 1 Unit Of Model 1206 M And 179 Units Of Model 1206 3 | Nationwide And International
1995-07-26 | Devices | III | z-889/890-5 | Software For The Abl 500 And 600 Series Analyzers Blood Gas Analyzers Used For Measuring A Patient's Ph Pc 02 And P 02 From Blood Samples | Radiometer America | Radiometer America Copenhagen Denmark | April 2 1994 | The Entry Of Patient Information Is Improperly Sequenced And The Analyzers May Ignore The Information And May Not Calculate Correctly The Blood Gas Results For The Patient's Actual Temperatur | 1 322 Units Were Distributed | Nationwide
1995-07-26 | Devices | III | z-891-5 | Ez Ch 50 Complement Reference Control Low Ivd Catalog 789 008 Intended To Be Used With Ez Complement Reagents To Measure Total Complement Activity In Human Serum By Immunochemical Techniques | Manufacturer | Diamedix Corporation Miami Florida | March 22 1995 | The Low Control Was Reported As Being Hig | 199 Units Were Distributed | Nationwide Spain Canada Germany The Netherlands Venezuela
1995-07-26 | Devices | III | z-892/900-5 | Anti Rnp Microassay Eia Ivd 96 Test Set Intended For The Detection And Quantitation Of Antibodies Against The Antigen Rnp In Serum As An Aid In The Diagnosis Prognosis And Monitoring Of Certain Rheumatic Autoimmune Diseases A Anti Rnp Microassay Catalog No | Manufacturer | Diamedix Corporation Miami Florida | October 24 1994 | The Calibrator Which Contains Both An Sm And Rnp Antigen Was Not Performing To Specification | 394 Sets Were Distributed Firm Estimates None Remains On The Market | Nationwide And International
1995-07-26 | Devices | III | z-901/909-5 | Anti Ssb Microassay Eia Ivd Catalog 783 260 96 Test Set Intended For The Detection And Quantitation Of Antibodies Against The Antigen Ssb La In Serum As An Aid In The Diagnosis And Prognosis Of Certain Rheumatic Autoimmune Diseases A Antigen Wells For Anti Ssb Microassay Catalog No | Manufacturer | Diamedix Corporation Miami Florida | August 26 1994 | The Calibrator Which Contains Ssb Antibodies Was Losing Activit | 214 Sets From Lot 71104 And 296 Sets From Lot 61204 Were Distributed Firm Estimates None Remains On The Market | Nationwide And International
1995-07-26 | Devices | III | z-910/918-5 | Anti Scl 70 Microassay Eia Test Set Ivd Intended For The Detection And Quantitation Of Antibodies Against The Antigen Scl 70 In Human Serum As An Aid In The Diagnosis And Prognosis Of Scleroderma A Antigen Wells For Anti Ssb Microassay Catalog No | Manufacturer | Diamedix Corporation Miami Florida | September 30 1994 | A High Incidence Of Non Specific Positive Result | 108 Sets Were Distributed Firm Estimates None Remains On The Market | Nationwide And International
1995-07-26 | Devices | III | z-953-5 | Software For The Electra 900 Automatic Coagulation Timer For The Electra 900 C Automatic Coagulation Timer For The Electra 1000 C Automatic Coagulation Timer And For The Electra 1000 C Automatic Coagulation With Bar Code Station Used To Test Different Chemical Clotting Factors Using Human Blood Plasma | Manufacturer | Medical Laboratories Automation Inc | September 19 1993 | Software Defec | 2 763 Units Were Distributed | Nationwide And International
1995-07-26 | Devices | III | z-990/1001-5: code type s1565, lot nos | Plastic Solution Containers For Eml Series Analyzers | Radiometer America | Radiometer America Copenhagen Denmark | May 1 1995 | The Gasket Behind The Foil Inlet Seal Of The Adapter That Connects The Bottle To The Analyzer May Not Be Seated Properl | 15 887 Containers Were Distributed | Nationwide
1995-07-26 | Devices | III | z-1004-5 | Axsym Reaction Vessels List 08 A 75 01 In Vitro Diagnostic Used To Hold The Patient's Sample And The Reagents For Miroparticle Enzyme Immunoassays Utilizing The Axsym Processing Center | Abbott Laboratories | Nypro Gurnee Illinois | June 20 1995 | The Use Of These Lots Of Reaction Vessels May Cause Imprecise Patient Result | 99 151 Cartons Were Distributed | Nationwide And International