## 1995-09-20
RECALL_CLASSIFICATION_DATE | PRODUCT_TYPE | CLASSIFICATION | RECALL_NUMBER | PRODUCT | RECALLING_FIRM | MANUFACTURER | RECALL_INITIALIZATION_DATE | REASON | VOLUME | DISTRIBUTION
1995-09-20 | Drugs | II | d-245-5 | Maximum Strength Theraflu Flu And Cold Medicine For Sore Throat Apple Cinnamon Flavor Packaged 6 Packets Per Box 24 Boxes Per Case Otc For Pain Relief Fever Reducing Nasal Decongestant And Antihistamine | Sandoz Consumer Pharmaceuticals | Sandoz Canada Inc | August 31 1995 | Product Contains Aspartame But Label And Labeling Do Not Declare Amount Of Phenylalanine Per Dosage Uni | 275 304 Boxes Were Distributed To Wholesale Or Retail Accounts 7 910 Boxes Were Distributed To Sandoz Representatives | Nationwide
1995-09-20 | Drugs | II | b-532-5 | Human Tissue From Cadaveric Donors Intended For Human Transplantation | Manufacturer | Washington State Tissue Services Seattle Washington | March 6 1995 | Human Tissue From Cadaveric Donors Was Distributedd Which Lacked Adequate Documentation About The Donor's Relevant Medical History Related To Risk Factors For Or Clinical Evidence Of Hepatitis B Hepatitis C Or Human Immunodeficiency Virus Hiv Infection Tested Repeatedly Reactive For The Antibody To The Hepatitis C Virus Encoded Antigen Anti Hcv And Lacked Adequate Documentation Indicating That Donors Received Less Than Four Units Of Blood Blood Components Colloids Or Crystalloids Within Forty Eight Hours Prior To Taking The Blood Sample Or Lacked An Adequate Algorithm To Ensure That There Was Not Hemodilution Sufficient To Alter Test Result | Approximately 402 Pieces | Texas Louisiana California
1995-09-20 | Biologics | II | b-375/376-5 | A Red Blood Cells B Fresh Frozen Plasma | Manufacturer | American Red Cross Blood Services Peoria Illinois | January 19 1995 | Blood Products Which Tested Initially Reactive For Antibody To The Human Immunodeficiency Virus Type 1 Anti Hiv 1 And Were Not Retested In Duplicate Were Distributed | 1 Unit Of Each Component | Illinois Iowa
1995-09-20 | Devices | I | z-1184/1202-5 | Blood Glucose Test Strips 1 | Manufacturer | Diagnostic Solutions Inc | January 26 1995 | Devices May Provide Inaccurate Test Results Which Would Give Falsely Elevated Reading Or Falsely Lower Reading | Approximately 1 000 000 Units Were Manufactured Of Which Most Of Which Were Distributed | Nationwide And International
1995-09-20 | Devices | II | z-888-5 | Sarns Turbo Membrane Oxygenator Catalog 9443 Part 98 0702 0545 9 | Manufacturer | Sarns 3 M Health Care Ann Arbor Michigan | May 2 1994 | There Is A Potential Packaging Defect In The Final Bag Seal Due To A Damaged Heater Band In The Vertrod Sealer Used To Seal The Polymylar To Polymylar Pouc | 61 Cases 3 Units Per Case Were Distributed 3 | California Massachusetts Missouri North Carolina Texas Sweden Japan
1995-09-20 | Devices | II | z-1041-5 | Columbia Sheep Blood Agar With 5 Prepared Culture Media Plates Product M 1013 | Manufacturer | Micro Diagnostics Inc | August 2 1995 | Some Of These Media Plates Labeled As Columbia Sheep Blood Agar Actually Contained Plates Of Tsa With 5 Bovine With Esculi | 240 Plates Were Distributed | Missouri Wisconsin Vermont
1995-09-20 | Devices | II | z-1055/1058-5 | Rusch Sterile Agt Preformed Endotracheal Tubes A Catalog 100180 Xxx B Catalog 100181 Xxx C Catalog 111780 Xxx D Catalog 111781 Xxx | Rusch | Willy Rusch Ag Kernen Germany | October 19 1994 | Pouch Seals Can Weaken And Cause The Packages To Ope | 10 050 Tubes Were Distributed Firm Estimates None Remains On The Market | Nationwide
1995-09-20 | Devices | II | z-1077/1085-5 | Centrysystem 3 Cartridge Blood Tubing Sets Catalog Numbers 003109 400 003109 410 003110 500 003111 500 003112 500 003113 500 003114 500 003210 500 003212 500 | Cobe Renal Care | Cobe Renal Care Inc | June 16 1995 | The Saline Administration Line Could Have An Occlusion Due To Plastic Molding Flashing Which May Not Be Detected During The Prime Or Recirculation Mode | 31 000 Units | Nationwide And International
1995-09-20 | Devices | II | z-1168-5 | Harmonic Scalpel 5 Mm Sterile Dissecting Hook Used With The Laparosonic Blade System Used During General And Gynecological Laparoscopic Surgery | Manufacturer | Ultracision Inc | July 27 1995 | The Packaging Tray May Be Punctured Compromising The Sterility Of The Devic | 144 Units Were Distributed | Georgia Colorado Massachusetts New York Maryland Kansas New Jersey Pennsylvania California Missouri
1995-09-20 | Devices | II | z-1173-5 | Star X Surgical Lasers Used In Podiatric Surgery | Manufacturer Fda Approved The Firm's Corrective Action Plan August 25 | Surg Assist Inc | August 25 1995 | Noncompliance With Performance Standards For Laser Products In That The Operator's Manuals Contained Inadequate Calibration Procedure | 30 Units | Nationwide And International
1995-09-20 | Devices | II |  | Gnomos Stereotactic System Used By Neurosurgeons To Determine Target Coordinates For Brain Biopsy | Nomos Corporation | Hyperion Inc | July 19 1995 | A Change In The Length Of The Rear Arc Bracket Component Can Result In Missed Target Coordinates Which Could Result In Removing Brain Tissue From An Unintended Area During A Brain Biopsy Procedur | 3 Units Were Distributed | Florida Massachusetts South Carolina
1995-09-20 | Devices | III | z-1037/1038-5 | Philips Gryoscan Magnetic Resonance Imaging Systems A Model T 5 With Vax Computer Revision 2 | Philips Medical Systems | Philips Medical Systems Nederland Best Netherlands | N A | Software Error Impeded Proper Image Analysi | 72 Units | Undetermined