## 1995-01-24
RECALL_CLASSIFICATION_DATE | PRODUCT_TYPE | CLASSIFICATION | RECALL_NUMBER | PRODUCT | RECALLING_FIRM | MANUFACTURER | RECALL_INITIALIZATION_DATE | REASON | VOLUME | DISTRIBUTION
1995-01-24 | Drugs | II | d-077-6 | Myochrysine Gold Sodium Thiomalate Injection 50 Mg Ml In 10 Ml Multi Dose Vials Used For Rheumatoid Arthritis Under The Merck And King Pharmaceuticals Label | Manufacturer | Merck And Company Inc | December 18 1995 | Product Does Not Meet Antimicrobial Effectiveness Tes | Approximately 90 000 Unit Packages Remained On Market At Time Of Recall Initiation | Nationwide Spain Peru
1995-01-24 | Drugs | III | d-078-6 | Sodium Iodide I 131 Solution For Oral Therapeutic Use In 100 Mci And 50 Mci Vials Used For The Treatment Of Hyperthyroidism And Selected Cases Of Carcinoma Of The Thyroid | Cis Us | Cis Bio International Subsidiary Of Compagnie Oris Industrie Sa Cedex France | December 15 1995 | Some Packages Contained A Package Insert For Sodium Iodide I 131 Capsule | 71 100 Mci Vials And 55 50 Mci Vials Were Distributed | Nationwide
1995-01-24 | Devices | II | z-321/325-6 | I V Extension Sets With T Connectors A Abbott Extension Set With T Connector | Abbott Laboratories | Abbott Laboratories Laurinburg North Carolina Hg Lot Suffix Abbott Health Products Barceloneta Puerto Rico H 1 Lot Suffix | December 8 1995 | Some Of The Extension Sets Leak And Disconnect Between The Male Fitment Of The Device And The Luer Fitment Of The | 2 120 180 Units Were Distributed Firm Estimated That 135 000 150 000 Units Remained On Market At Time Of Recall Initiation | Nationwide And International
1995-01-24 | Devices | II | z-332/334-6 | System 97 A Helium Charged Mobile Intra Aortic Balloon Pump Iabp Catalog 0998 00 0104 X For Use As A Patient Aid During Pre Intra Or Post Operative Open Heart Surgery Use Within Patients Demonstrating Unstable Angina Use Within Patients With Left Main Artery Occlusion Or Poor Left Ventricle Function A Power Supply B Front End Board C Solenoid Driver Board | Manufacturer | Datascope Corporation Paramus New Jersey | August 31 1995 | The System May Fail Because Of The Following Component Failures Power Supply Failure Of A Zener Diode Designated At Vr 9 In The System Power Supply Which Can Result In A Power On Condition Which Cannot Be Switched Off By Normal Mean | A 259 Units B 554 Units C 153 Units Were Distributed | Nationwide And International
1995-01-24 | Devices | II | z-335-6 | 3 M Sarns Brand Perfusion System 9000 Gas Flow System Used To Provide And Monitor Co 2 Flow To The Patient During Cardiopulmonary Bypass | Manufacturer | Sarns 3 M Health Care Ann Arbor Michigan | July 24 1995 | The Relay Controlling The Co 2 Flow Can Experience Random Failure And Latch In The On Position Causing Co 2 Gas To Continue To Flow After Completion Of The Pre Bypass Co 2 Flush Of The Perfusion Circui | Approximately 500 Units | Nationwide And International
1995-01-24 | Devices | II | z-336/337-6 | Duopulse And Unipulse Dental Laser System Used In Dental Surgery | Manufacturer Fda Approved The Firm's Corrective Action Plan December 28 | Excel Quantronix Corporation Hauppauge New York | December 28 1995 | Noncompliance With Performance Standards For Laser Products In That The Operator's Manuals Lacked Adequate Calibration Procedures And The Devices Had Several Labeling Noncompliance | 213 Units Were Distributed | Nationwide Germany Korea
1995-01-24 | Devices | II | z-346-6 | Tempo Tingle Timers Installed With Diagnostic X Ray Systems Used To Regulate The Duration And To Some Extent The Amount Of Exposure To X Radiation | Manufacturer Fda Approved The Firm's Corrective Action Plan January 3 | Tingle X Ray Products Inc | January 3 1996 | Devices Were Not Tested For Accuracy In Accordance With The Specifications And Did Not Bear Proper Certification And Identification Labels As Required By 21 Cfr 102 | 17 Timers | Nationwide
1995-01-24 | Devices | II | z-347-6 | Immunocard Helicobacter Pylori Test Kit Catalog 710030 A Rapid Enzyme Immunoassay For The Detection Of Igg Antibodies To Helicobacter Pylori In Human Serum And Plasma | Manufacturer | Meridian Diagnostics Inc | November 14 1995 | The Enzyme Conjugate In The Kit Was Prepared At The Wrong Dilution Resulting In An Improper Sensitivity Of The Test | 175 Test Kits Of Lot 710030 007 And 57 Test Kits Of Lot 710030 008 Were Distributed | Nationwide Italy Russia Puerto Rico
1995-01-24 | Devices | II | z-348-6 | Captia Rubella M Eia Test Kit Product 801 165 | Centocor | Centocor Uk Ltd | June 23 1995 | There Is A Higher Incidence Of Equivocal Positive Results Than Anticipate | 200 Kits Were Distributed | Nationwide And International
1995-01-24 | Devices | II | z-349/357-6 | Critikon Brand Protectiv And Protectiv Plus I | Johnson And Johnson | Johnson And Johnson Company Formerly Known As Critikon Inc | April 10 1995 | An Opening May Be Present In The Clear Plastic Portion Of The Blister Package Therefore Compromising The Sterile Barrier Of The Packag | 778 191 Catheters Were Distributed | Nationwide And International
1995-01-24 | Devices | III | z-358-6 | Cellfree Interleukin 2 Receptor Il 2 R Bead Assay Kit Catalog Ak 3120 For The Quantitative Measurement Of Interleukin 2 Receptor Level In Human Serum | Manufacturer | T Cell Diagnostics Inc | January 28 1994 | Controls Provided With The Product Were Not Meeting The Assigned Ranges And Fell Below Specified Limit | 20 Kits Were Distributed | Massachusetts Texas Connecticut Louisiana Greece
1995-01-24 | Devices | III | z-360-6 | Glucometer Encore Test Strips Used To Measure Glucose In Whole Blood | Manufacturer | Bayer Corporation Elkhart Indiana | July 7 1995 | Subject Lots May Fail To Give Blood Glucose Reading And The Glucometer Encore Meter Will Display An E 3 Error Cod | 45 451 Cartons | Nationwide