## 1996-04-17
RECALL_CLASSIFICATION_DATE | PRODUCT_TYPE | CLASSIFICATION | RECALL_NUMBER | PRODUCT | RECALLING_FIRM | MANUFACTURER | RECALL_INITIALIZATION_DATE | REASON | VOLUME | DISTRIBUTION
1996-04-17 | Drugs | II | d-117-6 | Children's Cherry Flavored Acetaminophen Suspension Liquid In 4 Ounce Bottles Under The Following Brand Names American Fare Brooks Chateau Cvs Food Lion Goldline Good Sense Health Life Homebest Kroger Marquee The Medicine Shoppe Meijer Nex Osco Pathmark The Pharxm Phar Mor Revco Smiths Thrifty Topcare Walgreens | Manufacturer | Perrigo Company Allegan Michigan | February 21 1996 | Precipitate Of Active Ingredient Which May Result In Non Uniform Dosag | 3 979 Dozen Bottles Were Distributed | Nationwide
1996-04-17 | Drugs | III | d-116-6 | Gelusil Liquid Antacid Anti Gas Peppermint Flavor 355 Ml In 12 Fluid Ounce Bottles | Warner Lambert Company | Warner Lambert Company Brockville Ontario Canada | December 11 1995 | Some Units Exhibited Off Odor Off Taste And Presence Of A Bacillus Specie | 293 046 Bottles Were Distributed Firm Estimated That 290 000 Bottles Remained On Market At Time Of Recall Initiation | Nationwide
1996-04-17 | Drugs | III | d-118-6 | Sterile Triamcinolone Diacetate Suspension Usp 40 Mg Ml Injectable Rx Drug For Intramuscular Intra Articular Or Soft Tissue Use For Endocrine And Rheumatic Disorders | Steris Laboratories | The Upjohn Company Kalamazoo Michigan Raw Material Steris Laboratories Inc | November 17 1995 | Product Contains Trace Quantities Of Decabromdiphenylether Resulting From Synthesis Reagent Used To Manufacture Raw Materia | Approximately 14 500 Units Were Distributed Firm Estimated That Little If Any Product Remained On Market At Time Of Recall Initiation | Nationwide
1996-04-17 | Biologics | III | b-363/365-6 | A Platelets B Fresh Frozen Plasma C Cryoprecipitated Ahf | Manufacturer | Flushing Hospital Medical Center Flushing New York | January 26 1996 | Blood Products Which Were Misbranded Were Distributedd For Further Manufactur | 66 Units B 459 Units C 10 Units | Pennsylvania
1996-04-17 | Devices | I | z-562-6 | Siemens 3 Gas Model 962 Anesthesia Mixer Part 906588 Designed For Use With The Siemens Servo Ventilator And May Be Used For Mixing 02 Oxygen With Either N 20 Nitrous Oxide Or Air | Siemens Medical Systems | Servo Ventilator Siemens Elma Germany Mixer | March 11 1996 | A Manufacturing Process Error In Which The Oxygen Hose Was Connected To The Nitrous Oxide Outlet And Vice Vers | 15 Units Were Distributed | California Washington D
1996-04-17 | Devices | II | z-560-6 | Hunter High Performance Air Purifiers Model 30300 For Home And Business Use | Manufacturer | Hunter Fan Company Memphis Tennessee | January 30 1996 | There Is A Potential Fire Hazard Due To High Voltage Arcin | 42 283 Units Were Distributed Firm Estimated That 39 146 Units Remained On Market At Time Of Recall Initiation | Nationwide And Puerto Rico
1996-04-17 | Devices | II | z-561-6 | Ivac I | Ivac Corporation | Ensembles Medicos S | October 17 1995 | The Device Is Labeled As Containing 60 Drops Ml Sets Whereas It Contains 20 Drops Ml Set | 475 Cases Were Distributed Firm Estimates None Remains On The Market | Nationwide And Canada
1996-04-17 | Devices | II | z-568/569-6 | Nuclear Imaging Systems Used In Organ And Whole Body Scans A Biad Two Head Detector Spect System B Monad Single Head Detector Spect System | Manufacturer | Trionix Research Laboratory Twinsburg Ohio | April 3 1995 | The Detector Head May Drop On A Patient Or Healthcare Provider Because Of Metal Fatigue At The Device's Ball Screw Assembl | 80 Biad Units Were Distributed Six 6 Biad Units And Five 5 Monad Units Are Still In Stock | Nationwide Korea Australia Egypt Switzerland
1996-04-17 | Devices | II | z-582-6 | P 300 Power Wheelchair With Airless Solid Casters And P 300 Power Wheelchair With Pneumatic Casters Using Labac Specialized Seating System | Manufacturer | Quickie Designs Inc | December 6 1995 | The Forks Holding The Front Caster Assemblies In Place May Be Subject To Cracking Which Could Cause The Caster Assembly To Collapse And The User To Potentially Fall Out Of The Wheelchai | 3 762 Units | Nationwide
1996-04-17 | Devices | II | z-588/591-6 | Medline Custom Procedural Trays Which Contain Johnson Johnson Barrier Vu Thru Cardiovascular Split Sheets A Cardiovascular Packs 1 Catalog No | Medline Industries | Medline Industrues Waukegan Illinois Kit Johnson Johnson Medical Inc | January 4 1996 | The Fabric Of The Clipping Tabs Are Subject To Tearing Presenting Potential Harm To Patien | Firm Estimates None Remains On The Market | Florida Georgia Kansas Kentucky Pennsylvania South Carolina Texas Utah Virginia
1996-04-17 | Devices | II | z-592/593-6 | Nezhat Dorsey Smokevac Trumpet Valve For Pump Ii A Laparoscopic Suction Irrigation System With Metered Smoke Evacuation Mechanism A Nezhat Dorsey Smokevac Trumpet Valve And Pump Ii Cartridge Catalog No | Davol | Prent Thermoforming Janesville Wisconsin Component | June 13 1995 | The Sterility Of The Device May Be Compromised Due To Loss Of Package Integrity Caused By Crack | A 7 112 Units B 1 134 Units Were Distributed | Nationwide And International
1996-04-17 | Devices | II | z-594-6 | Television Projectors Incorporated In The Moving Target Simulators Mts Used To Project Target Images On A Screen | Manufacturer Fda Approved The Firm's Corrective Action Plan April 1 | Aai Corporation Cockeysville Maryland | April 1 1996 | The Television Projectors Incorporated In The Moving Target Simulators Failed To Comply With The X Radiation Emission Limit Of The Federal Performance Standard For Television Receivers 21 Cfr 102 | 16 Units | Nationwide And Germany
1996-04-17 | Devices | II | z-595-6 | Model Ws 2200 Laser System Used In Metal Cutting | Manufacturer Fda Approved The Firm's Corrective Action Plan April 1 | Convergent Energy Sturbridge Massachusetts | April 1 1996 | The Product Was Found To Be Noncompliance With The Federal Performance Standard For Laser Products 21 Cfr 104 | 1 Unit | Mississippi
1996-04-17 | Devices | II | z-596-6 | Cordis Trakstar 18 Long Balloon Catheters Catalog 535 303 | Manufacturer | Cordis Corporation Miami Florida | January 15 1996 | Small Axially Oriented Abrasions Present On Some Balloons Could Cause A Burst Pressure Below That Of The Rated On | 50 Units Were Distributed | Nationwide
1996-04-17 | Devices | III | z-583-6 | Stratus Cardiac Troponin I Fluorometric Enzyme Immunoassay Kit Catalog B 5700 64 Used In An Automated Procedure For The Quantitative Determination Of Troponin I Levels In Serum And Heparinized Plasma | Manufacturer | Dade International Inc | January 10 1996 | Calibrator Lot Xtn 215 Was Incorrectly Shipped To Be Used With Kit Lot N | 118 Kits Were Distributed | Texas California Florida Georgia Kentucky Louisiana Michigan Minnesota Missouri New York Ohio Pennsylvania South Carolina Wisconsin And International
1996-04-17 | Devices | III | z-584-6 | S P Boiling Water Bath Used For All Laboratory Procedures Requiring 100 Degree C Incubation | Manufacturer | Dade International Inc | February 19 1996 | The Water Bath Has Water Leakage In The Rear Of The Unit Which Could Cause A Fire Or Electrical Shock Hazar | 109 Units | Illinois
1996-04-17 | Devices | III |  | Medtronic Cardiorhythm Atakr Ablation System Model 0610 Intended Primarily For Use In Intracardiac Radiofrequency Ablation Procedures |  | Medtronic Cardiorhythm San Jose California | N A | Device Does Not Turn Off When The Foot Is Released From The Foot Peda | 430 Units Were Distributed | Nationwide And International