## 1996-05-15
RECALL_CLASSIFICATION_DATE | PRODUCT_TYPE | CLASSIFICATION | RECALL_NUMBER | PRODUCT | RECALLING_FIRM | MANUFACTURER | RECALL_INITIALIZATION_DATE | REASON | VOLUME | DISTRIBUTION
1996-05-15 | Drugs | II | d-136-6 | Medline Epi Clenz Instant Antiseptic Hand Cleanser An Otc Liquid For Handwashing To Decrease Bacteria On The Skin | Medline Industries | Polychem Corporation New Haven Connecticut | April 20 1996 | Some Bottles Were Mislabeled As Tlc Mouth Rins | 297 Cases Were Distributed Firm Estimated That 25 Percent Of Product Remained On Market At Time Of Recall Initiation | Nationwide
1996-05-15 | Drugs | II | d-137-6 | Co 2 Compressed Gas | Manufacturer | Mg Industries Inc | April 23 1996 | Presence Of Trace Quantities Of Hcn Hydrogen Cyanide | Undetermined | Pennsylvania Maryland West Virginia Virginia Tennessee Kentucky Georgia North Carolina New Jersey Bermuda
1996-05-15 | Drugs | II | d-138-6 | Carbon Dioxide Usp 50 Pound And 7 Pound Cylinder Capacity | Repacker | Sunox Inc | May 2 1996 | Presence Of Trace Quantities Of Hcn Hydrogen Cyanide | Approximately 1 200 Cylinders 50 Pound Capacity And 159 Cylinders 7 Pound Capacity Were Produced | North Carolina Virginia
1996-05-15 | Drugs | III | d-133-6 | Cytarabine Injection 20 Mg Ml 1000 Mg 50 Ml Cytarabine Injection 20 Mg Ml 100 Mg 5 Ml Rx Drug For Iv Or Subcutaneous Use | Faulding Pharmaceutical Company Purepac | Faulding Pharmaceutical Company David Bull Laboratories Mulgrave Victoria 3170 Australia | March 8 1996 | Particulate Matter Formation Calcium And Silicate | 32 400 Bottles Were Distributed Firm Estimated That 50 60 Percent Of Product Remained On Market At Time Of Recall Initiation | Nationwide
1996-05-15 | Drugs | III | WARNING | Barre National A Bromanate Elixir In 4 Ounce Bottles Otc Decongestant Antihistamine Under The Albertsons And Janet Lee Labels 6 B Tussex Cough Syrup In 4 Ounce Bottles Otc Cough Syrup Cough Suppressant Nasal Decongestant Expectorant Under The Moore And Major Labels | Manufacturer | Barre National Inc | April 26 1996 | Product Labeling Not In Accord With Otc Monograph Missing Warning Statement | Firm Estimates None Remains On The Market | California Connecticut Florida Michigan Texas
1996-05-15 | Biologics | II | b-393-6 | Platelets | Manufacturer | United Blood Services Chicago Illinois | February 12 1996 | Blood Product Prepared From A Unit Of Whole Blood Designated As A Difficult Collection Was Distributed | 1 Unit | Illinois
1996-05-15 | Biologics | III | b-366-6 | Red Blood Cells | Manufacturer | United Blood Services Chicago Illinois | February 8 1996 | Red Blood Cells Stored At Room Temperature For An Unknown Amount Of Time Were Distributed | 2 Units 7 | Illinois
1996-05-15 | Biologics | III | b-395-6 | Recovered Plasma | Manufacturer | American Red Cross Blood Services Peoria Illinois | October 30 1995 | Blood Product Corresponding To A Unit Of Red Blood Cells Implicated In A Febrile Transfusion Reaction Was Distributed | 1 Unit | California
1996-05-15 | Devices | II | z-640/646-6 | Sterile Intravenous Administration Iv Sets With 1 | Mcgaw | Mcgaw Of Puerto Rico Inc | January 19 1996 | An Inadequate Bond Design May Cause Tube Filter Separation Resulting In Leakag | 21 145 Cases 895 324 Units Were Distributed | Nationwide Puerto Rico Canada
1996-05-15 | Devices | II | z-724-6 | Tanfit 24 Suntanning Bed | Manufacturer | Puretan International Asterisk Inc | November 14 1995 | Eighteen 18 Gauge Wire Was Used In The Timers In Lieu Of 12 Gauge Which Caused The Timers To Shut The Lamps Off Prematurel | 38 Units | Nationwide
1996-05-15 | Devices | II | z-727-6 | Lumin Laparoscopic Uterine Manipulator Injector Catalog Mis 100 Used For Manipulation Of The Uterus During Diagnostic And Surgical Procedures | Manufacturer | Utah Medical Products Inc | March 25 1996 | The Sterility Of The Device May Be Compromised Due To Loss Of Package Integrit | 305 Units Were Distributed | Arizona California Indiana Kansas Maine Michigan Tennessee Texas Virginia And International
1996-05-15 | Devices | II | z-742/746-6 | Puritan Bennett 7200 Series Ventilators Indicated For Delivery Of Gas To Patients Dependent On Artificial Respiration A Model No | Manufacturer | Nellcor Puritan Bennett Npb Carlsbad California | October 2 1995 | Device Contains A Defective Pressure Transducer Causing It To Stop Ventilation At Set Parameter | 2 987 Units Were Distributed | Nationwide And International
1996-05-15 | Devices | II | z-748/749-6 | Biostar Strep A Oia Test Intended For In Vitro Diagnostic Use As An Aid In The Diagnosis Of Group A Streptococcal Pharyngitis A Catalog No | Manufacturer | Biostar Inc | April 18 1996 | The Test Kits May Produce False Positive Results When Performed At The Maximum Assay Times Stated In The Package Inser | Firm Estimates 125 000 Test Devices Or Approximately 1 250 100 Test Kits Were Distributed | Nationwide And International
1996-05-15 | Devices | II | z-750/751-6 | Action Cat Brand Power Scooter Three Wheeled Wheel Chair A Cat Traditional B Cat Aviation | Manufacturer | Invacare Corporation Elyria Ohio | January 19 1996 | The Brake Key Pin Which Connects The Motor Shaft To The Scooter Brake May Slide Out Of Position Causing The Brake To Become Non Functiona | 671 Scooters Were Distributed | Nationwide
1996-05-15 | Devices | II | z-752/753-6 | Uroview Diagnostic Imaging Systems | Manufacturer Fda Approved The Firm's Corrective Action Plan November 24 | Oec Medical Systems Salt Lake City Utah | November 24 1993 | The Devices Were Found To Be In Noncompliance With 21 Cfr 1032 A And 21 Cfr 102 | 61 Units | Nationwide And International
1996-05-15 | Devices | II | z-754-6 | Model 9400 Mobile Fluoroscopy Imaging System | Manufacturer Fda Approved The Firm's Corrective Action Plan November 17 | Oec Medical Systems Salt Lake City Utah | November 17 1993 | The Device Was Found To Be In Noncompliance With 21 Cfr 1032 C In That Wires In The X Ray Control May Cause An Accidental Radiation Occurrence Aro Resulting In The X Ray Uni | 767 Units | Nationwide And International
1996-05-15 | Devices | II | z-755-6 | Aruba Suntanning Bed | Manufacturer | Puretan International Asterisk Inc | May 16 1994 | The Shaft On Timer Was Too Short And This Caused The Knob Of The Timer To Jam When The Temperature Increased And Expanded The Plastic Canopy Of The Devic | 368 Units Were Distributed | Nationwide And Canada
1996-05-15 | Devices | II | z-756-6 | Ophthalas 532 Laser System Used In Ophthalmic Surgery | Manufacturer Fda Approved The Firm's Corrective Action Plan May 2 | Alcon Laboratories Inc | May 2 1996 | The Device Was Defective Within 21 Cfr 100 | 848 Units 11 | Nationwide And International
1996-05-15 | Devices | II | z-757-6 | Tri Rolls Brand Power Scooter A Three Wheeled Wheelchair | Manufacturer | Invacare Corporation Elyria Ohio | February 21 1994 | Under Severe Use Or Heavy Abuse Tiller Spring That Turns The Tiller Throttle Can Break And Or Slip Off Due To A Metal Burr Causing The Scooter To Remain In Drive Unless The Drive Control Is Manually Moved Into Neutral Positio | 5 631 Scooters Were Distributed | Nationwide
1996-05-15 | Devices | II | z-764/765-6 | Hydromat Hydrovalve And Tubing Set For Laparoscopic Suction And Irrigation A Hydromat Hydrovalve Tubing Set Catalog 263101 61 Series 1 Set B Hydromat Hydrovalve Tubing Set Catalog 263104 61 Series 4 Set | Karl Storz Endoscopy | Infinity Extrusion And Engineering Inc | January 30 1996 | Device Sterility May Be Compromised Due To A Defective Packaging Sea | 720 Sets Were Distributed | California Colorado Florida Georgia Kentucky Massachusetts Maryland New York Ohio Oklahoma Pennsylvania South Carolina Texas Virginia Washington State West Virginia
1996-05-15 | Devices | III | z-581-6 | Sentinel Seal Autotransfusion Chest Drainage System Cdu Catalog 8888 571489 A Sterile Single Use Prescription Device | Manufacturer | Sherwood Medical Company Commerce Texas | October 12 1994 | Device Is Subject To Air Leak Which May Occur Between The Patient Connection Site And The Chest Drainage System Collection Chambe | 2 680 Units Were Distributed | Belgium France Greece Italy Spain
1996-05-15 | Devices | III | z-726-6 | Baxter Myelogram Tray Containing An 18 Gauge X 3 1 2 Spinal Needle With Stylet And Nine Holed Aspiration Needle Used To Obtain Spinal Fluid Samples And Removal Of Oil Based Contrast Media During Myelograms | Baxter Healthcare Corporation | Baxter Healthcare Corporation Mannford Oklahoma | April 19 1996 | The Aspiration Cannula May Not Fit Inside The Spinal Needle As Intende | 470 Cases Were Distributed Firm Estimated That 50 Percent Of Product Remained On Market At Time Of Recall Initiation | Nationwide