## 1996-05-22
RECALL_CLASSIFICATION_DATE | PRODUCT_TYPE | CLASSIFICATION | RECALL_NUMBER | PRODUCT | RECALLING_FIRM | MANUFACTURER | RECALL_INITIALIZATION_DATE | REASON | VOLUME | DISTRIBUTION
1996-05-22 | Drugs | III | d-139/140-6 | Doxycycline Hyclate Rx Broad Spectrum Antibiotic In Bottles Of 50 And 500 Under The Lederle Label A Doxycycline Capsules Usp 100 Mg B Doxycycline Tablets 100 Mg | Danbury Pharmacal | Danbury Pharmacal Inc | April 11 1996 | Products Distributedd With Non Current Package Inser | A 500 S 3 214 Bottles To Lederle 5 20 94 12 16 94 50 S 5 000 Bottles To Lederle 12 13 94 B 500 S 6 849 Bottles To Lederle 12 23 94 3 13 96 50 S 10 000 Bottles To Lederle 12 23 94 And 3 13 96 | Pennsylvania
1996-05-22 | Devices | II | z-716/721-6 | Calibration Gases Under The Scott And Or Kontron Labels A 4 Isoflurane Balance Argon B 1 Isoflurane 5 Co 2 65 N 2 O 21 O 2 Balance N 2 C 2 Isoflurane 21 O 2 Balance N 2 D 2 Isoflurane Balance N 2 E 2 Enflurane Balance N 2 | Manufacturer | Scott Specialty Gases Inc | February 5 1996 | Devices Do Not Meet Their Concentration Specifications As Stated On The Labeling Because Of A Defect In A Certain Sealing Gaske | 633 Units Were Distributed | Colorado New Jersey Virginia Sweden England
1996-05-22 | Devices | II | z-747-6 | Multitask Operating System Version 2 | Clintec Nutrition Company | Compdisk Morton Grove Illinois | April 12 1996 | A Software Bug In The Automatic Additive Selector Option May Result In Dispensing An Inaccurate Measure Of Electrolyte | 658 Software Units Were Distributed | Nationwide
1996-05-22 | Devices | II | z-758-6 | Image Amplification Surgical Table Model 2080 I | Amsco American Sterilizer Company | Amsco American Sterilizer Company Montgomery Alabama | January 6 1989 | There Is A Tipping Hazard When Positioning Obese Patients In The Reverse Orientatio | 1 395 Tables Were Distributed | Nationwide And International
1996-05-22 | Devices | II | z-759/760-6 | A Plv 100 Ventilators B Plv 102 Ventilators | Manufacturer | Lifecare International Inc | March 27 1996 | There Is A Potential For Alarm Malfunction And Or Unit Shutdown With No Alar | A 72 Units B 80 Units Were Distributed | Nationwide And International
1996-05-22 | Devices | II | z-766-6 | Sperti Brand Sunlamp Model P 164 Rt Riviera Table Model Used For Skin Tanning | Kbd | Time Teck Whitewater Wisconsin Component Supplier Responsible Firm | April 9 1996 | The Nut Which Secures The Timer To The Shell Of The Sunlamp Can Break The Plastic Timer Casing When It Is Tightene | 38 Units Were Distributed | Illinois Arizona Oregon Colorado Florida Nevada Maryland
1996-05-22 | Devices | II | z-788-6 | Ubm Model 840 Scanhead Joint Assembly Ophthalmic Ultrasound System | Carl Zeiss | Humphery Instruments Inc | January 17 1996 | The System's Scanhead May Drop On A Patient Because Of Metal Fatigue At The Device's Scanhead Joint Assembl | 34 Units Were Distributed | Nationwide And Canada
1996-05-22 | Devices | II | z-789-6 | Equipment Manuals For Amsco 3080 Rl Sp Surgical Table | Amsco American Sterilizer Company | Amsco American Sterilizer Company Montgomery Alabama Tables | August 10 1995 | The Manuals Specify An Erroneous Patient Weight Limitation Of 500 Pounds Instead Of 400 Pound | 1 901 Tables Were Distributed | Nationwide And International
1996-05-22 | Devices | II | z-790-6 | Sterile E Z Clean Cautery Tip Catalog 0012 An Electrode Used In Electrosurgical Pencils | Manufacturer | Megadyne Medical Products Inc | April 3 1996 | The Sterility Of The Device Has Been Compromised Due To Loss Of Package Integrit | Approximately 190 000 Units Were Distributed Firm Estimated That A Majority Of The Product Had Already Been Used At Time Of Recall Initiation | Nationwide Japan Italy
1996-05-22 | Devices | III | z-728-6 | Stratus Brand Ck Mb Fluorometric Enzyme Immunoassay A Rapid And Sensitive Automated Procedure For The Quantitative Determination Of Ck Mb In Human Serum | Manufacturer | Dade International Inc | January 26 1996 | The Above Lot Of Material May Have Units Which Do Not Contain Dulbecco's Phosphate Buffered Saline But Contain Wate | 116 Kits Were Distributed | Nationwide
1996-05-22 | Devices | III | z-767-6 | Inhibitory Mold Agar Media Plates Catalog 01 505 An In Vitro Diagnostic Packaged 10 Plates Per Box | Manufacturer | Remel Limited Partnership Lenexa Kansas | August 2 1995 | The Plates Inside The Box Were Labeled As Bhi Blood Heart Infusion Agar But The Box Was Labeled As Inhibitory Mold Agar Ima | 2 710 Units Were Distributed Firm Estimates None Remains On The Market | Nationwide
1996-05-22 | Devices | III | z-768-6 | Bactidrop Spot Indole Catalog 21 550 A Reagent Used To Determine The Ability Of An Organism To Split Indole From The Tryptophan Molecule Packaged 50 | Manufacturer | Remel Limited Partnership Lenexa Kansas | June 5 1995 | The Device Did Not Give The Expected Blue Color Reaction Because Ampule Contained Kovac's Indole Reagent Instead Of Spot Indol | 203 Boxes Were Distributed Firm Estimates None Remains On The Market | Georgia Florida Virginia Connecticut Wisconsin Oklahoma Kansas California Alaska New Jersey Pennsylvania Missouri
1996-05-22 | Devices | III | z-771-6 | Mrl Yellow Agar Double Pour Media Plates Catalog 01 543 An In Vitro Diagnostic Packaged 10 Units Per Box | Manufacturer | Remel Limited Partnership Lenexa Kansas | December 28 1995 | Several Packages Labeled Mrl Yellow Agar Contained D Nase Test Aga | 40 Units Were Distributed Firm Estimates None Remains On The Market | Missouri
1996-05-22 | Devices | III | z-772-6 | Mrl Red Double Pour Plates Catalog 01 541 An In Vitro Diagnostic Packaged 10 Plates Per Box | Manufacturer | Remel Limited Partnership Lenexa Kansas | November 13 1995 | The Packages Labeled As Mrl Red Contained Mrl Gree | 140 Units Were Distributed Firm Estimates None Remains On The Market | Missouri
1996-05-22 | Devices | III | z-778-6 | Brain Heart Infusion Bhi Agar With Blood With Chloramphenicol With Gentamicin Media Plates Catalog 10 248 An In Vitro Diagnostic | Manufacturer | Remel Limited Partnership Lenexa Kansas | August 17 1995 | The Individual Plates Were Labeled As Bhi With Blood Instead Of Bhi With Blood With Chloramphenicol With Gentamici | 140 Units Were Distributed Firm Estimates None Remains On The Market | Illinois California Utah