## 1996-05-29
RECALL_CLASSIFICATION_DATE | PRODUCT_TYPE | CLASSIFICATION | RECALL_NUMBER | PRODUCT | RECALLING_FIRM | MANUFACTURER | RECALL_INITIALIZATION_DATE | REASON | VOLUME | DISTRIBUTION
1996-05-29 | Drugs | II | d-142-6 | Dilantin Extended Phenytoin Sodium Usp 30 Mg In Bottles Of 100 Anticonvulsant | Manufacturer | Warner Lambert Company Lititz Pennsylvania | May 6 1996 | Product Does Not Meet Dissolution Specification | 28 397 Bottles Were Distributed | Nationwide
1996-05-29 | Drugs | II | d-143/144-6 | Puritan Bennett A Carbon Dioxide Usp 99 In The Following Cylinder Sizes B D E P F M G T And Xl 45 Refrigerated Liquid Container B Carbon Dioxide 5 Oxygen 95 Mixture In The Following Cylinder Sizes E G And H | Puritan Bennett Corporation | S Puritan Bennett Corporation A Subsidiary Of Nellcor Puritan Bennett Inc | April 19 1996 | Presence Of Trace Quantities Of Hcn Hydrogen Cyanide | The Baltimore Branch Distributed 1 253 Cylinders Or Xl 45 S Co 2 Between Philadelphia 1 29 96 4 19 96 Branch Distributed 5 Cylinders Co 2 Between 4 11 15 96 The Baltimore Branch Distributed 106 Cylinders Of Co 2 O 2 Mixture Between 2 9 96 4 18 96 | Pennsylvania Maryland New Jersey Washington D
1996-05-29 | Drugs | III | d-141-6 | Pharmics Brand Polytinic Rx Diet Supplement In Bottles Of 100 | Pharmics | Gull Laboratories Inc | March 25 1996 | Subpotency Of The Folic Acid Ingredien | 646 Bottles Were Distributed Firm Estimated That Less Than 100 Bottles Remained On Market At Time Of Recall Initiation | Missouri Utah Idaho Nevada Colorado Iowa Ohio Florida Wyoming
1996-05-29 | Drugs | III | d-145/146-6 | A Caffeine Tablets Manufactured For Specialty Products Co | Manufacturer | Ion Labs Inc | March 18 1996 | Caffeine Product Does Not Meet Content Uniformity Specification | A 2 300 Bottles B 468 Bottles Were Distributed | New Jersey
1996-05-29 | Drugs | III | d-147-6 | Is Subpotent | Whitehall Robins | Wyeth Ayerst Lederle Laboratories Inc | May 8 1996 | Potency Not Assured Through Expiration Dat | 29 093 Bottles Were Distributed | Nationwide And Australia
1996-05-29 | Drugs | III | d-148-6 | Tricolan Antimicrobial Soap In 32 Ounce Bottles | Manufacturer | Perfecto Products Manufacturing Inc | April 8 1996 | Subpotenc | 1 365 Bottles Were Distributed Firm Estimated That 30 40 Of The Lot Remained On Market At Time Of Recall Initiation | Nationwide
1996-05-29 | Drugs | III | b-410-6 | Human Tissues For Transplant Human Eye Tissues | Manufacturer | Mid South Eye Bank Memphis Tennessee | January 23 1996 | Human Tissues Not Tested For Hiv 2 Were Distributed | 112 Shipments | Tennessee Arkansas
1996-05-29 | Biologics | II | b-402-6 | Platelets Pheresis | Manufacturer | Irwin Memorial Blood Center San Francisco California | January 11 1994 | Blood Product Was Collected From A Donor Who Previously Tested Repeat Reactive For Hbsag And Was Not Properly Reentere | 1 Unit | California
1996-05-29 | Biologics | II | b-411/412-6 | A Red Blood Cells B Fresh Frozen Plasma | American Red Cross Blood Services | American Red Cross Blood Services Rock Falls Illinois | October 17 1995 | Blood Products Collected From A Donor Who Travelled To An Area Designated As Endemic To Malaria Were Distributed | A 4 Units B 1 Unit | Illinois Missouri California Tennessee
1996-05-29 | Biologics | II | b-414-6 | Red Blood Cells | Manufacturer | American Red Cross Detroit Michigan | October 24 1995 | Blood Product Which Was Collected From A Donor With A Medical History Of Hepatitis Was Distributed | 1 Unit | Michigan
1996-05-29 | Biologics | II | b-415/417-6 | A Red Blood Cells B Platelets C Fresh Frozen Plasma | Manufacturer | American Red Cross Fort Wayne Indiana | November 21 1995 | Blood Products Which Tested Negative For All Infectious Diseases But Were Collected From A Donor Who Previously Tested Repeat Reactive For Anti Hiv 1 And Was Never Properly Reentered Were Distributed | A 2 Units B 1 Unit C 1 Unit | Indiana
1996-05-29 | Biologics | II | b-418/420-6 | A Red Blood Cells B Platelets C Fresh Frozen Plasma | Manufacturer | American Red Cross Detroit Michigan | October 26 1995 | Blood Products Collected From A Donor Taking The Drug Proscar Were Distributed | 1 Unit Of Each Component | Michigan
1996-05-29 | Biologics | II | b-421-6 | Source Plasma | Manufacturer | Community Bio Resources Inc | October 16 1995 | Blood Products Which Were Collected From A Donor Who Reported His Lips Pierced Were Distributed | 11 Units | Austria
1996-05-29 | Biologics | II | b-422/424-6 | A Red Blood Cells B Platelets C Recovered Plasma | Manufacturer | Carlisle Hospital Blood Bank Carlisle Pennsylvania | August 21 1995 | Blood Products Collected From Donors Taking The Drug Proscar Were Distributed | A 2 Units B 2 Units C 3 Units | Pennsylvania
1996-05-29 | Biologics | III | b-403-6 | As 1 Red Blood Cells | Manufacturer | Irwin Memorial Blood Center San Francisco California | January 11 1994 | Blood Product Was Shipped With An Extended Expiratio | 1 Unit | California
1996-05-29 | Biologics | III | b-404/405-6 | A As 1 Red Blood Cells B Platelets | Manufacturer | Lifesource Glenview Illinois | September 29 1995 | Blood Products Collected From A Donor Who Had Received A Mmr Vaccination Twenty Five Days Prior To The Donation Were Distributed | 1 Unit Of Each Component | Illinois
1996-05-29 | Biologics | III | b-426-6 | Red Blood Cells | Blood Services | United Blood Service Mcallen Texas | February 20 1996 | Red Blood Cells Remained At Room Temperature For An Unknown Amount Of Time Were Distributed | 1 Unit | Texas
1996-05-29 | Devices | I | z-709-6 | Synthetic Serum Substitute Catalog 99193 In 12 Ml Vials Designed For Use As A Tissue Culture Protein Supplement | Manufacturer | Irvine Scientific Sales Company Inc | December 14 1995 | The Device Is Contaminated With Penicillium S | 1 669 Vials Were Distributed | Alabama California Connecticut Florida Hawaii Illinois Massachusetts Missouri New Jersey Nevada New York Ohio Pennsylvania Tennessee Texas Washington State Japan
1996-05-29 | Devices | II | z-725-6 | Harvey Vapo Steril Solution A Sterilizing Compound For Use With Harvey Chemiclave Sterilizers For Sterilizing Dental And Surgical Instruments | Manufacturer | Mdt Biologics Company Rancho Dominguez California | February 1 1996 | The Product May Not Adequately Sterilize Since An Active Ingredient Formaldehyde Was Found To Be Less Than Declared In Some Lots And Some Bottles In One Lot Contained No Formaldehyde At Al | 82 606 Bottles Were Distributed | Nationwide New Zealand Taiwan
1996-05-29 | Devices | II | z-761/763-6 | Transducer Protector In Custom Sterile Trays A Component Of A Sterile Custom Tray Identified With The User's Name And A Unique Work Order Number A Dualex Transducer Protector B Dualex Plus Transducer Protector C Dualex Ultra Transducer Protector | Sterile Design | Millipore Corporation Bedford Massachusetts Transducers | February 21 1995 | The Transducer Protectors May Exhibit Air Leakage At The Female Luer Lock Connection To The Male Luer Of The Hemodialysis Machine Under Normal Operating Condition | 490 578 Custom Kits Were Distributed | New Jersey
1996-05-29 | Devices | II | z-779-6 | Bacticard Neisseria A Test Kit For Use In The Rapid Presumptive Identification Of Pathogenic Neisseria Species Isolated From Primary Culture Of Clinical Specimens On Selective Media | Manufacturer | Remel Limited Partnership Lenexa Kansas | July 28 1995 | The Tests Resulted In A Positive Glut Reaction With Neisseria Gonorrhoea | 84 Units Were Distributed | California Oklahoma Michigan Virginia Georgia Wisconsin Vermont Tennessee Minnesota South Carolina Ohio Missouri Maryland Florida New Mexico Louisiana New York Illinois Rhode Island
1996-05-29 | Devices | II | z-791/808-6 | Orthopedic Bone Screws With Cancellous Or Cortical Thread Designs A Cortical Locking Screw All Lengths Part No | Manufacturer | Ace Medical Company El Segundo California | June 20 1995 | The Device Labels Did Not Contain The Following Required Warning Statement Warning This Device Is Not Approved For Screw Attachment Or Fixation To The Posterior Elements Pedicles Of The Cervical Thoracic Or Lumbar Spin | Approximately 105 195 Bone Screws Were Distributed Firm Estimates None Of The Misbranded Product Remains On The Market | Nationwide
1996-05-29 | Devices | II | z-809/810-6 | Stryker Rugged Ambulance Cot Models 6060 And 6070 For Ambulance Loading Unloading | Manufacturer | Stryker Medical Division Of Stryker Corporation Kalamazoo Michigan | April 2 1996 | The Siderail Mounting Holes On The Outer Support Rail Were Drilled Incorrectl | 99 Cots | Nationwide And Canada
1996-05-29 | Devices | II | z-815-6 | Serachem Plus Clinical Chemistry Control Serum Bovine Unassayed Level Ii Intended For Monitoring The Precision Of Assays For Digoxin Theophylline Phenobarbital Phenytoin Gentamicin And Tobramycin | Instrumentation Laboratory Company Ilc | Instrumentation Laboratory Company Orangeburg New York | September 25 1996 | The Reconstituted Control Serum Had Turbidity And A Drop In Calcium Creatine And Other Analyte | Approximately 1 051 Units Were Distributed | Pennsylvania Kentucky Tennessee Florida Washington State Oregon Montana Texas California Idaho Michigan Missouri Connecticut Wisconsin
1996-05-29 | Devices | II | z-816/818-6 | Hip Prosthesis Stem A Apollo Hip Non Porous Stem Size 2 Catalog Number 7370 00 02 B Natural Hip Non Porous Cocr Stem Size 1 Catalog Number 7352 35 001 C Natural Hip Porous Collarless Stem Size 2 Left Catalog Number 7354 01 102 | Manufacturer | Intermedics Orthopedics Inc | April 17 1996 | The Devices Were Mislabele | 19 Units Were Distributed Firm Estimates None Remains On The Market | Missouri Ohio Pennsylvania South Dakota Tennessee
1996-05-29 | Devices | II | z-819/830-6 | Disposable Steri Shield Togas Used During Surgical Procedures In Order To Help Protect The Wearer Against Contamination And Or Exposure To Body Fluids And Microorganisms A Part No | Stryker Instruments | Biomedical Devices Irvine California | April 16 1996 | The Sterility Of The Devices Have Been Compromised Due To Loss Of Package Integrit | 79 530 Togas Were Distributed | Nationwide And International
1996-05-29 | Devices | II | z-831-6 | Ecg Pre Amp Pc Board Assembly For The Ecg Option Of The Siemens Gamma Camera System | Siemens Medical Systems | Solectron Technology Inc | April 9 1996 | The Circuit Board Exceeds The Maximum Leakage Current Safety Requirement For Patient Connected Device | 67 Units Were Distributed | Nationwide And International
1996-05-29 | Devices | II | z-832/834-6 | 7200 Series Microprocessor Ventilators Indicated For Delivery Of Gas To Patients Dependent On Artificial Respiration A Model No | Manufacturer | Nellcor Puritan Bennett Npb Carlsbad California | October 2 1995 | Ventilators Can Enter The Svo State Stopped Ventilating When They Are First Turned On Due To Defective Softwar | 8 402 Ventilators Were Distributed | Nationwide And International
1996-05-29 | Devices | II | z-835/840-6 | Power Wheelchairs Marketed Under The Following Brand Names A Ranger X Power Wheelchair 14 B Ranger Ii Power Wheelchair C Arrow Power Wheelchair D Arrow Xt Power Wheelchair E Jaguar Power Wheelchair F Power 9000 Action Cat | Invacare Corporation | Invacare Corporation Elyria Ohio Wheelchairs Gould Electronics Mexico Component | November 22 1994 | A Wire Harness Connecting The Battery To The Controller Contains A Fuse Holder That May Cause The Wheelchair To Lose Power Due To A Faulty Soldering Wel | 4 761 Chairs Were Distributed | Nationwide
1996-05-29 | Devices | II | z-850/851-6 | Laser Systems Used As Optical Oscillators For Highly Specialized Laboratory Applications A Model Kerbbqs Eafb Laser System B Model Cylf Ktp 1 Optical Parametric Oscillator Opo | Manufacturer Fda Approved The Firm's Corrective Action Plan On May 3 | Schwartz Electro Optics Inc | May 3 1996 | The Laser Systems Were Found To Be In Noncompliance With 21 Cfr 101 | 1 Unit Each | Southeast And Southwest Region Of The United States
1996-05-29 | Devices | III | z-769-6 | Urea Broth For Afb 1 | Manufacturer | Remel Limited Partnership Lexena Kansas | November 7 1995 | The Broth Was Pink In Color While Being Stored In Refrigeration At 2 8 Centigrade Compared To Another Lot That Was Yello | 540 Units Were Distributed Firm Estimates None Remains On The Market | California Virginia Connecticut Wisconsin Michigan Pennsylvania Arkansas Kansas New York Florida Ohio Indiana Puerto Rico
1996-05-29 | Devices | III | z-770-6 | Phenol Red Broth With 1 Sucrose Tubed Media For In Vitro Diagnostic Use | Manufacturer | Remel Limited Partnership Lexena Kansas | February 7 1996 | The Product Appeared As A Semi Solid That Was Light Pink In Colo | 260 Units Were Distributed Firm Estimates None Remains On The Market | Ohio California Nebraska Illinois New York Missouri
1996-05-29 | Devices | III | z-773-6 | Is As A Broth That Is Orange Red In Color | Manufacturer | Remel Limited Partnership Lexena Kansas | October 4 1995 | The Device Does Not Support Growth Including Growth Of The Positive Control Organism | Approximately 2 030 Units Were Distributed Firm Estimates None Remains On The Market | Texas Ohio Georgia Florida Colorado California Virginia Connecticut Wisconsin Kansas
1996-05-29 | Devices | III | z-774-6 | Cystine Heart Agar Without Blood 20 Ml Tubed Media For In Vitro Diagnostic Use | Manufacturer | Remel Limited Partnership Lexena Kansas | November 7 1995 | The Device Was Made With Heart Infusion Broth Instead Of Heart Infusion Agar Causing It To Be Liquid Instead Of A Soli | 40 Units Were Distributed | Kansas Oklahoma
1996-05-29 | Devices | III | z-775/776-6 | Glycerol Buffered Saline For In Vitro Diagnostic Use In Qualitative Procedures For The Preservation Of Enteric Pathogens And The Transportation Of Fecal Specimens To The Laboratory A Glycerol Buffered Saline Fecal Transport System 15 Ml Vials Packaged 12 Vials Per Box Product Catalog No | Manufacturer | Remel Limited Partnership Lenexa Kansas | January 23 1996 | The Devices Contained Black Particles Within The Saline Solutio | A 39 Boxes B 2 Boxes Were Distributed | Nebraska Indiana Oklahoma
1996-05-29 | Devices | III | z-777-6 | Buffered Cye With Ccvc Charcoal Yeast Extract Agar With Cycloheximide Cephalothin Vancomycin And Colistin For In Vitro Diagnostic Use | Manufacturer | Remel Limited Partnership Lenexa Kansas | October 9 1995 | The Bcye With Ccvc Was Inadvertently Released Into Distribution With Poor Growth Of Legionella Pheumophil | 70 Units Were Distributed | Illinois New York
1996-05-29 | Devices | III | z-780-6 | Surecaltm 1 Calibrator For Use In The Calibration Of The Il Phoenixtm Chemistry Eletrolyte Analyzer For In Vitro Diagnostic Use | Instrumentation Laboratory Company Ilc | Instrumentation Laboratory Company Orangeburg New York | October 23 1995 | Chloride Channel For The Il Phoenix Analyzer Would Not Calibrate Or The Chloride Recoveries Were High For Patient Value | Approximately 103 Cartons 12 5 Ml Bottles Were Distributed | New York New Jersey North Carolina South Carolina Mississippi Florida Minnesota Missouri
1996-05-29 | Devices | III | z-787-6 | Cutter Perfourm Light C | Heraeus Kulzer | Heraeus Kulzer Gmbh Company Kg Dormagen Germany | April 3 1996 | The Device Fails To Set Within The Specified Time Fram | 1 327 Kits Were Distributed | Nationwide And Canada
1996-05-29 | Devices | III | z-811/812-6 | Referril B Calibrator And Il Test Tco 2 Reagent A Referril B Calibrators Catalog No | Instrumentation Laboratory Co | A Instrumentation Laboratory Company Orangeburg New York B Medical Analysis Systems Inc | March 11 1996 | The Tco 2 Recoveries Were Lower Than Labeled On The Monarch System Using A New Liquid Form Of Reagent | Approximately 2 046 Units Of Referril B And 502 Units Of Tc 02 Reagent Monarch Were Distributed 18 | Nationwide And International
1996-05-29 | Devices | III | z-813/814-6 | Serachem Assayed Control Sera For Monitoring The Ilab 500 Ilab 900 Ilab 1800 Monarch And Phoenix Chemistry Systems And The Il 943 Flame Photometer A Level 1 Catalog No | Instrumentation Laboratory Company | Instrumentation Laboratory Company Orangeburg New York | November 22 1995 | The Label Inserts Had The Incorrect Amylase And Triglyceride Value | Approximately 21 12 5 Ml Vials Packages Were Distributed | New York New Jersey Massachusetts Florida Wisconsin California Maryland Virginia Arkansas Oklahoma
1996-05-29 | Devices | III | z-846-6 | Dade Protein C Clotting Assay Kit In Vitro Diagnostic Product Used For The Quantitative Determination Of The Anticoagulant Activity Of Protein C In Plasma | Manufacturer | Dade International Inc | February 19 1996 | The Kit Was Packaged Using Actin Activated Cephaloplastin Reagent Which Has Been Found To Perform Outside Its Specification | 82 Kits Were Distributed | Texas Maryland Wisconsin Illinois Kentucky Michigan Ohio California Australia
1996-05-29 | Devices | III | z-847/849-6 | Isolab's Hemocard S And A S Kits In Vitro Diagnostic Test Kits For The Detection Of Hemoglobin Variants Hb A S Or Hb S | Manufacturer | Isolab Inc | April 15 1996 | The Signal For Hb A And Hb S Has Weakened And May Produce A Negative Signal On A Positive Control Or Patient Sample Prior To The Kits Expiration Dat | 186 Kits Were Distributed | Nationwide Puerto Rico Belgium Israel Portugal Canada England