## 1996-06-05
RECALL_CLASSIFICATION_DATE | PRODUCT_TYPE | CLASSIFICATION | RECALL_NUMBER | PRODUCT | RECALLING_FIRM | MANUFACTURER | RECALL_INITIALIZATION_DATE | REASON | VOLUME | DISTRIBUTION
1996-06-05 | Drugs | III | d-150-6 | J B Williams Cepacol Maximum Strength Sore Throat Spray Cool Menthol | J B Williams Company | Les Emballages Knowlton Knowlton Quebec Canada | September 5 1995 | Subpotenc | 58 212 Units Were Distributed | Nationwide
1996-06-05 | Devices | I | z-729/741-6 | Avian Transport Adult And Pediatric Ventilator A Avian Ventilator Part No | Manufacturer | Bird Products Corporation Palm Springs California | April 1 1996 | All Device Models Have Oxygen Bleeding Or Leaking Internally Which Is Not Being Adequately Vented To The Outside Of The Cas | Approximately 3 200 Units Were Distributed | Nationwide And International
1996-06-05 | Devices | II | z-781/786-6 | Receptal Suction Liner System Used With The Receptal Hard Canister For Vacuum Suction Waste Body Fluid Management In Hospital Operating Rooms A List 43024 01 2000 Ml Liner Lid 1 4 Port V 3024 01 Vha Plus Label B List 43027 01 2000 Ml Liner Lid With Pour Spout Arthroscopy V 3027 01 Vha Plus Label C List 43042 01 2000 Ml Liner Lid With Vac Gard Internal Non Mechanical Shut Off Valve Filter V 3042 01 Vha Plus Label D List 43044 05 2000 Ml Liner Lid With Vac Gard Internal Non Mechanical Shut Off Valve Filter Pour Spout V 3044 05 Vha Plus Label E List 43043 01 1500 Ml Liner Lid With Vac Gard Internal Non Mechanical Shut Off Valve Filter V 3043 01 Vha Plus Label 2 F List 43044 01 1500 Ml Liner Lid With Vac Gard Internal Non Mechanical Shut Off Valve Filter Pour Spout | Abbott Laboratories | Abbott Laboratories Salt Lake City Utah | November 8 1995 | The Receptal Liner Lids May Pop Off The Suction Canisters Compromising The Function Of The Devic | Approximately 8 000 000 Units Are Distributed Annually Firm Estimates None Remains On The Market | Nationwide
1996-06-05 | Devices | II | z-841-6 | Calibration Gas Cylinders 10 02 Balance N 2 Used To Calibrate And Check An Oxygen Carbon Dioxide Analyzer | Manufacturer | Scott Specialty Gases Inc | April 13 1995 | The Cylinders Labeled As 10 Oxygen Balance Nitrogen Actually Contain 10 Carbon Dioxide Balance Nitroge | 36 Cylinders Were Distributed | Washington State Indiana New York Iowa Kansas Wisconsin Korea Canada
1996-06-05 | Devices | II | z-842/845-6 | Bird Ventilators Indicated For Delivery Of Medical Gases To Patients Dependent On Artificial Respiration A Model No | Bird Products Corporation | Vernay Sfernice Div | December 21 1995 | The Devices Contain A Defective Potentiometer Which Can Cause The Ventilators To Cease Operating While Alarms Remain Functiona | 799 Ventilators And Upgrade Kits Are Subject To Recall | Nationwide And International