## 2002-7-31
RECALL_CLASSIFICATION_DATE | PRODUCT_TYPE | CLASSIFICATION | RECALL_NUMBER | PRODUCT | RECALLING_FIRM | MANUFACTURER | RECALL_INITIALIZATION_DATE | REASON | VOLUME | DISTRIBUTION
2002-7-31 | Drugs | III | d-353-2 | Acyclovir Tablets Usp 400 Mg Rx Only 100-count Bottles Manufactured By: Ivax Pharmaceuticals Miami Fl 33137 Ndc 0172-4267-60. Recall # D-353-2. | Ivax Pharmaceuticals Inc. Miami Fl | Ivax Pharmaceuticals Caribe Inc. Cidra Puerto Rico. | May 21, 2002 | Tablet crumbling. | 11 483 bottles | Nationwide
2002-7-31 | Drugs | III | d-354-2 | A) luvox (fluvoxamine Maleate) Tablets (scored) 25 Mg 100 Count Bottles And Unit Dose Packages Of 100 Rx Only Recall # D-354-2; B) Luvox (fluvoxamine Maleate) Tablets (scored) 50 Mg 100 Count Bottles And Unit Dose Packages Of 100 Rx Only. Recall # D-355-2; C) Luvox (fluvoxamine Maleate) Tablets (scored) 100 Mg 100 Count Bottles And Unit Dose Packages Of 100 Rx Only. Recall # D-356-2. | Solvay Pharmaceuticals Inc. Marietta Ga | Solvay Pharmaceuticals Inc. Baudette Mn. | May 29, 2002 | Inaccuracies in data submitted to the new drug application by solvay (stability). | 1 503 114 units | Nationwide
2002-7-31 | Biologics | II | b-1657-2 | Source Plasma Recall # B-1657-2. | Nabi Chattanooga Tn | Nabi Chattanooga Tn | December 7, 2000 | Blood product, collected in a manner that compromises the sterility of the unit, was distributed. | 1 unit | NC
2002-7-31 | Biologics | II | b-1658-2 | Source Plasma Recall # B-1658-2 | Nabi Johnson City Tn | Nabi Johnson City Tn | May 13, 2000 | Blood product, collected in a manner that compromises the sterility of the unit, was distributed. | 1 unit | CA
2002-7-31 | Biologics | II | b-1659-2 | Source Plasma Recall # B-1659-2 | Nabi Johnson City Tn | Nabi Johnson City Tn | July 5, 2000 | Blood product, collected from a donor who had not completed the cjd/nvcjd increased risk questions, was distributed. | 1 unit | CA
2002-7-31 | Biologics | II | b-1696-2 | A) Platelets Pheresis Leukocytes Reduced Recall # B-1696-2; B) Platelets Pheresis Leukocytes Reduced Irradiated Recall # B-1697-2 | Blood Bank Of San Bernadino And Riverside Counties San Bernadino Ca | Blood Bank Of San Bernadino And Riverside Counties San Bernadino Ca | February 5, 2002 | Platelets with a decreased platelet count were distributed. | a 1 130 units b 136 units | CA
2002-7-31 | Biologics | II | b-1736-2 | Source Plasma Recall # B-1736-2 | Community Bio-resources Inc. Hoover Al | community Bio-resources Inc. St. Cloud Mn. | January 20, 1998 | Blood product, that was not quarantined after testing elevated for the alanine aminotransferase (alt) assay, was distributed. | 1 unit | Austria
2002-7-31 | Biologics | III | b-1735-2 | Source Plasma Recall # B-1735-2. | Seracare Inc. Topeka Ks | Seracare Inc. Topeka Ks | July 27, 2001 | Blood products, that tested negative for the antibody to hepatitis c virus (anti-hcv), were collected from a donor that subsequently tested repeatedly reactive to anti-hcv, were not properly quarantined, and were distributed. | 2 units | CA
2002-7-31 | Devices | II | z-1123-2 | Cryovalve Allograft - Heart Valve. Recall # Z-1123-2. | Cryolife Inc. Kennesaw Ga | Cryolife Inc. Kennesaw Ga | April 5, 2002 | Aortic valve donor was associated with alleged case of endocarditis. | one valve | CA
2002-7-31 | Devices | II | z-1136-2 | Ltv Series Ventilators Recall # Z-1136-2. | Pulmonetic Systems Inc. Colton Ca | Pulmonetic Systems Inc. Colton Ca | November 30, 2001 | Emi shield short circuits units, potential for fire. | 64 | PA OH FL NC GA PA WA CA and TO Canada Japan and Taiwan
2002-7-31 | Devices | II | z-1145-2 | Ct Tri Pak Sterile Disposable Syringe Kits Recall # Z-1145-2. | Medrad Inc. Indianola Pa | Medrad Inc. Indianola Pa | May 7, 2002 | May exceed endotoxin specification limits | 41650 kits | Nationwide and Sultanate OF Oman Korea Taiwan and BR
2002-7-31 | Devices | II | z-1147-2 | Intera 10t And 1.5t Mr Systems Magnetic Resonance Systems Are Indicated For Use As Diagnostic Devices That Display The Internal Structure And/or Function Of The Head Body Or Extremeties. The Coils Are Used To Obtain Optimal Imaging. Recall # Z-1147-2/z-1148-2. | Philips Medical Systems North America Inc. Bothell Wa | Philips Medical Systems Nederland B.v. Netherlands. | October 26, 2001 | Higher than expected current in the insulation of the coil cable. | 3 | FL and WA
2002-7-31 | Devices | II | z-1151-2 | Staar Surgical Uv Absorbing Posterior Chamber Intraocular Lens Model - Aq5010v Recall # Z-1151-2. | Staar Surgical Co. Monrovia Ca | Staar Surgical Co. Monrovia Ca | May 15, 2002 | Mislabeled as to diopter. | 156 | Nationwide and Puerto Rico and the Dominican Republic
2002-7-31 | Devices | II | z-1152-2 | Ela/angeion Lyra Model 2020 2021 And 2022 Implantable Cardioverter Defibrillators Recall # Z-1152-2/z-1154-2. | Angeion Corporation St. Paul Mn | Ela-angeion Plymouth Mn | June 27, 2002 | The devices could stop providing therapy due to premature battery depletion. | 494 | Nationwide and Internationally
2002-7-31 | Devices | II | z-1155-2 | Dyonics Direct View And Video Arthroscopes 40 Mm As Follows: A) Part Number: 3622r Video Arthroscope 4mm X 30 Degree Repair. Recall # Z-1155-2; B) Part Number: 3622s Video Arthroscope 4mm X 30 Degree Service. Recall # Z-1156-2; C) Part Number: 3626r Arthroscope 4mm X 30 Degree Repair. Recall # Z-1157-2. D) Part Number: 3626s Arthroscope 4mm X 30 Degree Service. Replacement. Recall # Z-1158-2. E) Part Number: 4125s Arthroscope 4mm X 30 Degree Service Replacement. Recall # Z-1159-2. | Smith & Nephew Inc. Andover Ma | Smith & Nephew Inc. Andover Ma | April 29, 2002 | Arthroscopes may be missing weld and contamination may penetrate resulting in inadequate sterilization. | 33 | Nationwide and Australia Mexico and the United Kingdom
2002-7-31 | Devices | II | z-1160-2 | Stryker Brand Tps U2 Drill For Use With The Stryker Total Performance System Recall # Z-1160-2. | Stryker Corporation Stryker Instruments Division Kalamazoo Mi | Stryker Corporation Stryker Instruments Division Kalamazoo Mi | April 29, 2002 | Possible shock hazard. | 84 drills | Nationwide and Internationally
2002-7-31 | Devices | II | z-1161-2 | Surgeons Vest Recall # Z-1161-2. | Precept Medical Products Inc. Arden Nc | Precept Medical Products Inc. Arden Nc | April 3, 2002 | Incomplete seal on packaging. | 187 cases 30 per case or 5 610 units | Nationwide
2002-7-31 | Devices | II | z-1162-2 | Sonoline Antares Diagnostic Ultrasound System The Sonoline Antares Is A High-resolution Digital Broadband Diagnostic Ultrasound Imaging System. It Supports The Following Diagnostic Applications: Abdominal Ob/gyn Small Parts Musculoskeletal Vascular And Urology. Recall # Z-1162-2. | Semens Medical Systems Inc. Issaquah Wa | Semens Medical Systems Inc. Issaquah Wa | January 20, 2002 | Software anomalies that affect measurement and calculation data. | 260 units | Nationwide
2002-7-31 | Devices | II | z-1163-2 | Kimberly-clark Multifunction Electrodes R2 Pads Self-adhesive Pre-gelled Low Impedance Electrodes Manufactured By Ballard Medical Products. Recall # Z-1163-2. | Ballard Medical Products Draper Ut | Ballard Medical Products Draper Ut | March 25, 2002 | Unit label incorrectly identifies the pads as radiolucent. | 88 518 units | Nationwide and Internationally
2002-7-31 | Devices | II | z-1165-2 | Ispan Sulfur Hexafluoride (sf6) And Perfluoropane (c3f8) Liquified Gas Under Pressure For Intraocular Use Cylinder Series Ld2 Ld1 And Ld. Recall # Z--1164-2/z-1165-2. | Scott Medical Products Plumsteadville Pa | Scott Medical Products Plumsteadville Pa | April 3, 2001 | Post-operative warning against use of nitrous oxide gas before intraocular gas has dissipated from the eye. | 2900 cylinders 2nd notification 4205 cylinders | Nationwide
2002-7-31 | Devices | II | z-1166-2 | Baxter Ipump Pain Management System Infusion Pump Product | Baxter Healthcare Corp. Round Lake IL | Baxter Healthcare Pte. Ltd. Singapore Sg. | April 10, 2002 | Software anomaly allows access to change prescription or program bolus without security. | 9 927 pumps | Nationwide
2002-7-31 | Devices | II | z-1167-2 | Boston Scientific/medi-tech Dual Lumen Vaxcel Picc (peripherally Inserted Central Catheter) As Follows: Catalog Number product Description A) 45-413 vaxcel 5f Dual Lumen Picc Recall # Z-1167-2; B) 45-426 Vaxcel 6f Dual Lumen Picc. Recall # Z-1168-2. C) 45-445 Picc Vaxcel 5f Dl/clampable/60cm Wire/ob. Recall Z-1169-2; D) 45-446 Picc Vaxcel 6f Dl/clampable/60cm Wire/ob. Recall # Z-1170-2. | Boston Scientific Corp./medi-tech Watertown Ma | Medsource Technologies Llc Navajoa Sonora. | April 22, 2002 | Catheter extension tube detached from hub upon being flexed. | 5 214 | Nationwide
2002-7-31 | Devices | II | z-1173-2 | Vaxcel Single Lumen Peripherally Inserted Central Catheter (picc) As Follows: Catalogue Number Description A) 45-430 4f Sl/clamp/130 Cm Picc Recall # Z-1173-2; B) 45-432 5f Sl/clamp/130 Cm Picc. Recall # Z-1174-2. | Boston Scientific Corp.medi-tech Natick Ma | Medsource Technologies Llc Navajoa Sonora. | March 25, 2002 | Guidewire is incorrectly coiled in protective hoop with luer fitting at wrong end. | 1112 units | Nationwide and Canada
2002-7-31 | Devices | II | z-1175-2 | A) Model Im2 Licox Bolt System Double Lumen Recall # Z-1175-2; B) Model Im3 Licox Bolt System Triple Lumen. Recall # Z-1176-2; C) Model Im2.s Complete Licox Probe Kit Double Lumen. Recall # Z-1177-2; D) Model Im3.s Complete Licox Probe Kit Triple Lumen. Recall # Z-1178-2; E) Model Im3.st Complete Licox Probe Kit With Temperature Probe Triple Lumen. Recall # Z-1179-2. | Integra Neuro Care Llc San Diego Ca | Integra Neuro Care Llc San Diego Ca | April 30, 2002 | Catheter/ bolt incompatibilty may result in erroneous intracranial pressure measurements. | 388 bolts | Nationwide
2002-7-31 | Devices | II | z-1180-2 | Camino Intracranial Pressure Monitoring Catheter With Licox Imc Bolt Fitting Model 110-4l. Recall # Z-1180-2. | Integra Neurocare Llc San Diego Ca | Integra Neurocare Llc San Diego Ca | April 26, 2002 | Catheter collapse may result in erroneously high pressure readings. | 44 estimated in | Distribution Nationwide
2002-7-31 | Devices | II | z-1181-2 | Ge Signa 5t 1.0t And 1.5t Mr/i Extremity Coils. Recall # Z-1181-2. | General Electric Medical Systems Waukesha Wi | General Electric Medical Systems Waukesha Wi | March 8, 2002 | Extremity coil pads could catch on fire and burn patient. | 2573 devices | Nationwide and Internationally
2002-7-31 | Devices | II | z-1182-2 | Ti Cann Troch Fixation Nail (trochanteric Fixation Nail) The Responsible Firm On The Label Is Synthes Usa Monument Co. Recall # Z-1182-2/z-1183-2. | Synthes (usa) Paoli Pa | Synthes (usa) Monument Co. | May 16, 2002 | Manufactured outside design specification. | 16 units | AZ CO FL MO MS OR PA and TX
2002-7-31 | Devices | III | z-1146-2 | Choice A.b. (aberation Blocking) Daily Wear Soft (hydrophilic) Contact Lens. The Product Is Packaged In Trial And Multipacks. Recall # Z-1146-2. | Ciba Vision Corp. Duluth Ga | Specialty Ultravision St. Hubert Quebec Canada. | May 24, 2002 | The polymer was undercured, lens may appear cloudy. | 781 units | Nationwide
2002-7-31 | Devices | III | z-1149-2 | A)cardiac/ Peritoneal Catheter Standard Barium Impregnated 90cm Catalog No 43103. Recall # Z-1149-2; B) Peritoneal Catheter Open End With Wall Slits Standard Barium Impregnated 90 Cm Catalog No. 43522. Recall # Z-1150-2. | Medtronic Neurosurgery Goleta Ca | Medtronic Neurosurgery Goleta Ca | May 14, 2002 | Product mix-up/ substitution/ mislabeling. | 219 | Nationwide and Canada Australia and Korea
2002-7-31 | Devices | III | z-1171-2 | Bone Stack Calcium Sulfate Recall # Z-1171-2. | Howmedica Osteonics Corp. Allendale Nj | Stryker Howmedica Osteonics Limerick Ireland. | January 10, 2001 | Volume of pellets is less than listed on the labeling | 193 units | Nationwide
2002-7-31 | Devices | III | z-1172-2 | Hemashield Gold Woven Double Velour 4-branch Vascular Graft In 28 And 30mm Sizes Catalog Number 175928 And 175930. Recall # Z-1172-2. | Boston Scientific Corp./medi-tech Wayne Nj | Boston Scientific Corp./medi-tech Wayne Nj | March 25, 2002 | 28 and 30mm sizes were mislabeled. | 17 28 mm size 17 30 mm size | Nationwide