## 1999-5-12
RECALL_CLASSIFICATION_DATE | PRODUCT_TYPE | CLASSIFICATION | RECALL_NUMBER | PRODUCT | RECALLING_FIRM | MANUFACTURER | RECALL_INITIALIZATION_DATE | REASON | VOLUME | DISTRIBUTION
1999-5-12 | Devices | II | z-817-9 | 25 G Whitacre Needle Spinal Tray Intended To Be Used In The Administration Of A Regional Or Local Anesthesia To A Patient | Manufacturer | Sims Portex Inc | March 26 1999 | Product Labeled As Sterile Was Not Sterilized | 3 Cases 10 Units Per Case Were Distributed | Colorado And Indiana
1999-5-12 | Devices | II | z-831/832-9 | Hyperinflation Sets And Resuscitator | Manufacturer | Owens Bigam Medical Company Newland North Carolina | November 25 1998 | Damage To The Oxygen Tubing Could Cause Leaks In The Tubing That Would Result In Inconsistent Flow Rates | 128 Cases And 63 Samples Were Distributed | Arizona California Colorado Georgia Illinois Kansas Louisiana Maryland Minnesota Missouri New Jersey North Carolina Ohio Oregon Pennsylvania Texas
1999-5-12 | Devices | II | z-852-9 | Dual Purpose Rocker Instrument Part 4009 6 | Interpore Cross International | P A K Manufacturing | February 18 1999 | The Device Will Not Allow The Proper Assembly Of A Slider Into A Hook When Used In Accordance With The Product Instructions For Use | 27 Units Were Distributed | California Kentucky Illinois Massachusetts Michigan Minnesota Missouri North Carolina Nebraska Tennessee Texas Virginia
1999-5-12 | Devices | II | z-853/854-9 | A 5 Fr Angiographic Catheters Model Nos | Manufacturer | Millar Instruments Houston Texas | March 19 1999 | A Coagulated Clot Was Found Within The Catheter Which Were Made Of Nylon Instead Of Polyurethane As Specified In The Firm�s 510 K | 290 Units Were Distributed | Nationwide Canada Japan
1999-5-12 | Devices | II | z-855/859-9 | Quinton Q Cath Hemodynamic Analysis System A Hemodynamic Recording System Which Acquires Displays Stores And May Print Out Waveform Data Acquired From A Patient Undergoing A Cardiac Catheterization Procedure A Q Cath B Q Cath Ii C Q Cath Ii Nt D Q Cath Ds E Q Cath Remote Station | Manufacturer | Quinton Instrument Company Bothell Washington | October 1 1998 | Device Has Potential To Display Previous Patient's Blood Pressure Product Gem Implantable Cardioverter Defibrillator A Model No | Approximately 81 Units Were Distributed | Nationwide Canada The Netherlands Thailand New Zealand Egypt
1999-5-12 | Devices | II | z-862-9 | Fleximount Ceiling Track Mount System With Trolly Model 0102180 Part Of The Mounting System Designed To Connect The Burton Outpatient Plus Outpatient Ii Coolspot And Coolspot Lights To The Ceiling | Manufacturer | Burton Medical Products Corporation Chatsworth California | April 13 1999 | The Device Has The Potential To Fall Because Setscrews Used To Secure The Down Tube On The Trolley Were Omitted From The Installation Pack | 122 Units Were Distributed | Nationwide And Canada
1999-5-12 | Devices | II | z-863-9 | Pressure Reduction Valve Part 20132 065 A Component Of The Argon Plasma Coagulation Apc Unit | Erbe Usa | Erbe Tubingen Germany Recalled By Erbe Usa Inc | April 5 1999 | The Valve Has The Potential To Separate From The Argon Tank Because It Does Not Fit Securely In The Tank | 50 Units With The Affected Serial Numbers Have Been Distributed | Nationwide And Canada
1999-5-12 | Devices | II | z-864-9 | Nc 1 Floor To Ceiling Mount An Accessory To A Microscope Used For Various Surgical Procedures | Carl Zeiss | Oberlikon Contraves A | April 1 1999 | The Device Has The Potential To Fall During Use | 52 Units Were Distributed | Florida Kentucky Texas North Carolina Illinois Alabama Missouri Ohio Louisiana Arizona Georgia California Hawaii Delaware Colorado Utah Pennsylvania Arkansas
1999-5-12 | Devices | II | z-867/868-9 | Storz Silicone Tire Implant For Retinal Detachment Surgery A Catalog No | Bausch And Lomb Surgical | Vesta Inc | February 24 1999 | The Package Labeling May Not Correctly Identify The Size And Style Of The Contents | 473 Pouches Were Distributed | Nationwide And International
1999-5-12 | Devices | II | z-869/870-9 | Bayer Advia 120 Hematology System A Quantitative Automated Hematology Analyzer That Provides A Leukocyte Differential Count And Reticulocyte Analysis For In Vitro Diagnostic Use In Clinical Laboratories A Advia 120 Hematology System Catalog No | Bayer Corporation | Bayer Diagnostic Manufacturing Ltd | February 9 1999 | The Device May Give Erroneous Results Because Of The Valve Gasket Deterioration In The Baso Channel Check Valve And The Software Designed Will Not Flag Significant Discrepancies In The Results Between The Two Methodologies At Extremely Low Counts | Approximately 920 Units Were Distributed | Nationwide And International
1999-5-12 | Devices | II | z-871-9 | Ocufilcon D 55 Visibility Tinted Contact Lenses | Manufacturer | Ocular Sciences Inc | April 26 1999 | Incomplete Seal Could Compromise Product Sterility | 1 283 472 Lenses Were Distributed | Nationwide And International
1999-5-12 | Devices | II | z-872-9 | Colleague 3 Volumetric Infusion Pump Triple Channel For Continuous Or Intermittent Delivery Through Clinically Acceptable Routes Of Administration Such As Intravenous Intra Arterial Subcutaneous Epidural Or Irrigation Of Fluid Spaces Applications | Baxter Healthcare Corporation | Baxter Healthcare Pte | April 16 1999 | A Software Communication Error Occurs Predominately When All Three Channels Are In Use Causing An Alarm Condition 533 320 717 000 That Will Stop The Function Of All Of The Channels In Use | 6 992 Pumps Were Distributed | Nationwide
1999-5-12 | Devices | II | z-874/907-9 | Dmr 2 Disposable Manual Resuscitators 1 | Mallinckrodt | Nellcor Puritan Bennett De Mexico S | March 30 1999 | The Non Rebreathing Valve Nrv Housing Assembly May Crack Resulting In The Swivel Elbow Fitting Becoming Loose And Disengaging From The Nrv During Use | Approximately 100 716 Units Were Distributed | Nationwide And International