## 1995-02-15
RECALL_CLASSIFICATION_DATE | PRODUCT_TYPE | CLASSIFICATION | RECALL_NUMBER | PRODUCT | RECALLING_FIRM | MANUFACTURER | RECALL_INITIALIZATION_DATE | REASON | VOLUME | DISTRIBUTION
1995-02-15 | Drugs | II | d-059/060-5 | A Lortab 5 15 Tablets Hydrocodone Bitartrate 5 Mg And Acetaminophen 500 Mg B Propoxyphene With Apap 15 Tablets Propoxyphene Napsylate 100 Mg Acetaminophen 650 Mg Rx Analgesic | Repacker | Sentara Pharmacy Services Virginia Beach Virginia Repacker | January 4 1995 | A Bottle Containing Lortab 5 Tablets Was Mislabeled As Propoxyphene Acetaminophen Tablet | 36 Bottles Were Distributed | Virginia
1995-02-15 | Drugs | II | d-062/063-5 | Zestoretic Tablets Lisinopril Hydrochlorothiazide 20 12 | Zeneca Pharmaceuticals | Zeneca Pharmaceuticals Newark Delaware | December 9 1994 | Lack Of Assurance That The Products Will Meet Dissolution Specifications For The Hydrochlorothiazide Ingredien | 4 272 Cases 96 Blisters Per Case And 8 826 Cases 12 Bottles Were Distributed Firm Estimates That No More Than 5 Percent Of Amount Distributed Remained On Market At Time Of Recall Initiation | Nationwide
1995-02-15 | Drugs | II | d-064-5 | Isordil Tembids Isosorbide Dinitrate 4 Mg Extended Release Tablets In 100 Tablet Bottles A Vasodilator And Is Indicated For The Treatment And Prevention Of Angina Pectoris | Wyeth Ayerst Laboratories | Wyeth Ayerst Laboratories Inc | December 30 1994 | Use Of An Unapproved Grade Of Aluminum Stearat | 27 972 Bottles Were Distributed | Nationwide
1995-02-15 | Drugs | III | d-061-5 | Estratab Esterified Estrogens Tablets Usp 1 | Solvay Pharmaceuticals | Solvay Pharmaceuticals Baudette Minnesota | November 18 1994 | Product Below Potency Specification Prior To Expiration Dat | 12 160 100 Tablet Bottles And 735 1000 Tablet Bottles Were Distributed | Nationwide
1995-02-15 | Drugs | III | d-065-5 | Doctor's Best Brand Dry Skin Creme In 2 Ounce And 4 Ounce Plastic Jars Containing 20 Percent Urea Cream Designed For Treatment And Relief Of Eczematous Dry Skin Otc For Topical Use | Sherwood Pharmaceutical | Vita Erb Limited Springfield Missouri | January 24 1995 | Presence Of Mol | 300 4 Ounce Jars And 1 678 2 Ounce Jars Were Distributed | Missouri Illinois Arkansas Mississippi Florida Michigan Kansas Alabama Kentucky Texas Tennessee
1995-02-15 | Biologics | II | b-158-5 | Lifeline Blood Bank Management System Software | Manufacturer | Informedics Inc | Not-found | Computer Software That Contained Programming Errors Which Could Result In The Release Of Unsuitable Blood Products Was Distributed | Approximately 200 Copies Of The Various Releases Have Been Distributed | Nationwide And International
1995-02-15 | Biologics | III | b-159-5 | Lifeline Blood Bank Management System Software | Manufacturer | Informedics Inc | Not-found | Computer Software That Contained Programming Errors Which Could Result In Deficiencies In Record Keeping Required By Good Manufacturing Practice Regulations Was Distributed | Approximately 200 Copies Of The Various Releases Have Been Distributed | Nationwide And International
1995-02-15 | Devices | II | z-255/257-5 | Clinac Medical Linear Accelerators A Model 2100 C D B Model 2300 C D C Model 600 C | Manufacturer | Varian Associates Inc | February 18 1992 | A Software Problem Causes The Systems To Not Properly Identify An Incorrect Accessory In Positio | 113 Units | Nationwide
1995-02-15 | Devices | II | z-931/934-4 which appeared in the june 29, 1994 enforcement report: (a) sherman stay-wet 4 - wetting, lubricating, rewetting drops for fluoro/silicone acrylate and silicone acrylate (rgp) contact lenses, 30/30 ml; -9- (b) sherman de-stat 4 - cleaning, disinfecting and storage solution for fluoro/silicone and silicone acrylate (rgp) contact lenses, 118/120 ml; (c) sherman sof/pro clean (thimerosal), daily cleaner for soft (hydrophilic) contact lenses, 30 ml | Sterile Ophthalmic Solutions For Use With Contact Lenses Extension Of Recall Z 931 934 4 Which Appeared In The June 29 1994 Enforcement Report A Sherman Stay Wet 4 Wetting Lubricating Rewetting Drops For Fluoro Silicone Acrylate And Silicone Acrylate Rgp Contact Lenses 30 30 Ml 9 B Sherman De Stat 4 Cleaning Disinfecting And Storage Solution For Fluoro Silicone And Silicone Acrylate Rgp Contact Lenses 118 120 Ml C Sherman Sof Pro Clean Thimerosal Daily Cleaner For Soft Hydrophilic Contact Lenses 30 Ml | Sherman Pharmaceuticals | Sherman Pharmaceuticals Abita Springs Louisiana | July 6 1994 | Sterility May Have Been Compromised Due To Potential Inadequate Sterilizatio | Approximately 6 587 Units Of De Stat 4 7 088 Units Of Stay Wet 4 And 8 936 Units Of Sof Pro Clean Thimerosal Were Distributed | Nationwide
1995-02-15 | Devices | II | z-376/381-5 | Children's Sunglasses Included In Beachbag And Tote Bag Sets A Style 3342 Sunglasses Included In Barney Logo Bag With Plastic Strap And Snap Closing Or Cord Strap And Gathered Closing B Style 3343 Sunglasses Included In Case Attached To Barney Logo Bag With Plastic Strap And Snap Closing C Style 3370 Sunglasses Included In Case Attached To Barney Logo Bag With Plastic Strap And Velcro Closing D Style 3312 Sunglasses In Plastic Pouch Which Is Either Attached Directly To Bag Or In Plastic Pouch Attached With Plastic Strap Bag Has Beach Scene With Cats With Words Along The Edge Of The Bag With A Plastic Strap With Velcro Closing E Style 3315 Sunglasses In Plastic Pouch Which Is Either Attached Directly To Bag Or In Plastic Pouch Attached With Plastic Strap Bag Has A Beach Scene With People No Wording Snap Closing F Style 3371 Sunglasses In Plastic Pouch Which Is Either Attached Directly To Bag Or In Plastic Pouch Attached With Plastic Strap Bag Has Beach Scene With Cats With Words Along The Edge Of The Bag With A Plastic Strap With Velcro Closing Smaller Version Of Style 3312 | Jaclyln | Hoi Tak Handbag Factory Kowloon Hong Kong | March 11 1993 | The Sunglasses May Fail To Meet Fda's Requirements For Impact Resistant Lense | Approximately 210 624 Beach Bags Containing Sunglasses Were Distributed | Nationwide
1995-02-15 | Devices | II | z-382-5 | Hanks Balanced Salt Solution For In Vitro Diagnostic Use | Manufacturer | Life Technologies Inc | January 12 1995 | Sterility Of The Device Is Compromised Due To Bacterial Contamination With Gram Negative Rod | 1 660 Units Were Distributed | Nationwide Canada Taiwan Japan Hong Kong
1995-02-15 | Devices | II | z-386-5 | Millar Model Tc 510 Control Unit A Control Unit Which Is An Interface Between The Catheter Pressure Transducer And The Strain Gauge Amplifiers | Millar Instruments | Concord Technologies Houston Texas | July 25 1994 | Unauthorized Modifications Were Made By The Contract Manufacturer Which Resulted In The Potential For Inaccurate Erratic Or No Blood Pressure Reading | 9 Units Were Distributed And Recovered From The Market | Nationwide
1995-02-15 | Devices | II | z-346-4 which appeared in the march 16, 1994, designed for long term access to the vascular system for the delivery of medications, intravenous fluids, parenteral nutrition solutions, blood products and imaging solutions, and for the withdrawal of blood samples | Cv Catheter With Groshong Valve Long Term Single Lumen 8 French Product 7711800 And Cv Catheter 5 French With Groshong Valves Long Term Dual Lumen Percutaneous Placement Procedure Tray | Manufacturer | Bard Access Systems Inc | November 22 1993 | The Guidewires Exhibited A Greater Than Acceptable Level Of Pyrogen | A Approximately 500 Units B 31 Units C 10 Units D 2 Units Were Distributed | Nationwide
1995-02-15 | Devices | II | z-396-5 | Medtronic Capsure Sp Model 5524 M 53 Cm Steroid Eluting Bipolar Implantable Tined Atrial Transvenous Leads | Medtronic | Medtronic Puerto Rico Inc | December 21 1994 | The Connector Sleeve At The Proximal End Of The Leads Are Labeled Ven Instead Of Atr | 18 Leads Were Distributed | Nationwide
1995-02-15 | Devices | III | z-366/375-5 | Scalpex Reusable Ophthalmic Microsurgical Scalpels Used In Surgical Procedures Primarily Ophthalmic For Cutting Tissues A Model 9036 Angled Bevel Down Slit 2 | Eagle Vision | D S Manufacturing | August 1 1994 | The Device Is Labeled For Reuse However The Polymer Sleeve Melts During Sterilization In An Autoclav | 744 Scalpels Were Distributed | Nationwide And International
1995-02-15 | Devices | III | z-391/392-5 | Nrs 2 And Nrs 4 Electroencephalograph Biofeedback Systems | Manufacturer | Lexicor Medical Technology Inc | March 8 1994 | Some Units May Have Been Manufactured With The Ground And Reference Inputs Connected Togethe | Approximately 130 Units | Nationwide And International
1995-02-15 | Devices | III | z-393-5 | Model 1089 Mk Ii Checktrode Electrode Tester | Lexicor Medical Technology | Ufi Morro Bay California | February 18 1994 | Some Units May Have Been Manufactured With Grossly Shorter Battery Life Electrode Impedance Readings Which Are Erratic Or Which Bounce Between Several Numbers And Test Reading Below Normal Value 49 51 K | 36 Units Were Distributed | Nationwide Italy Germany Norway Japan Canada Mexico Israel