## 1993-02-24
RECALL_CLASSIFICATION_DATE | PRODUCT_TYPE | CLASSIFICATION | RECALL_NUMBER | PRODUCT | RECALLING_FIRM | MANUFACTURER | RECALL_INITIALIZATION_DATE | REASON | VOLUME | DISTRIBUTION
1993-02-24 | Drugs | II | d-138-3 | Oridol-dm Liquid Rx Liquid Oral Antitussive/mucolytic-expectorant 10 Mg Packaged In 16 Ounce Brown Plastic Bottles With White Plastic Caps Under The Following Labels: Luchem Qualitest And Major. Recall #d-138-3. | Manufacturer | H.n. Norton & Company (formerly Luchem Pharmaceuticals Inc.) Shreveport Louisiana. | December 28, 1992 | One or more bottles of oridol-c was mislabeled as oridol-d | 2 794 Units Under The Luchem Label 10 000 Units Under The Qualitest Label And 744 Units Under The Major Label Were Distributed. | Nationwide.
1993-02-24 | Drugs | II | d-139-3 | Redi-script Cephalexin Capsules Usp A Rx Antibiotic In Bottles Labeled To Contain 100 500-mg Capsules. Recall #d-139-3. | Repacker | Prepackage Specialists Division Of Pdrx Pharmaceuticals Inc. Oklahoma City Oklahoma (repacker/own Label Distributor). | February 1, 1993 | Some bottles labeled as 500 mg strength contain 250 mg capsules | 9 504 Bottles Were Distributed; Firm Estimates That Little If Any Of The Lot Remains On The Market. | New Jersey California Illinois.
1993-02-24 | Drugs | III | d-137-3 | Cotatate Dh Syrup Under The Major Label A Liquid Oral Antitussive/expectorant Rx Drug Consisting Of Hydrocodone Bitartrate 5 Mg And Guaifenesin 100 Mg Packaged In 16 Ounce Brown Plastic Bottles. Recall #d-137-3. | Manufacturer | H.n. Norton & Company (formerly Luchem Pharmaceuticals Inc.) Shreveport Louisiana. | December 28, 1992 | Product does not meet stability specifications | 3 428 Units Were Distributed. | Nationwide.
1993-02-24 | Drugs | III | d-140/143-3 | Rx Drugs For Injection In Partial Additive Plastic I.v. Bag Called Pab Bag: (a) 5% Dextrose Usp 100 Ml 50 Ml And 25 Ml; (b) 0.9% Sodium Chloride Usp 100 Ml 50 Ml; (c) 0.4% Theophylline And 5% Dextrose 100 Ml 50 Ml; (d) Metronidazole Usp 100 Ml. Recall #d-140/143-3. -2- | Manufacturer | Mcgaw Inc. Irvine California. | December 21, 1993 | Some units may contain rubber particulate matter | (a) 239 552 Units (100 Ml) 71 568 Units (50 Ml) 19 952 Units (25 Ml); (b) 316 736 Units (100 Ml) 341 040 Units (50 Ml); (c) 5 424 Units (100 Ml) 2 184 Units (50 Ml); (d) 74 616 Units Were Distributed. | Nationwide Canada.
1993-02-24 | Drugs | III | d-144-3 | Chem-tuss Elixir A Liquid Antihistaminic Nasal Vasoconstrictor Rx Drug Packaged In 16 Ounce Brown Plastic Bottles. Recall #d-144-3. | Manufacturer | H.n. Norton & Company (formerly Luchem Pharmaceuticals Inc.) Shreveport Louisiana. | November 3, 1992 | Product contains excess fd&c; blue #1 color | 104 Units Were Distributed. | Nationwide.
1993-02-24 | Devices | II | z-712-2 | Kendall Curity Disposable Non-sterile Latex Exam Gloves Large Ambidextrous A Disposable Device Intended For A Variety Of Uses In Hospitals And Other Health Areas. Recall #z-712-2. | Kendall Healthcare Products Company Mansfield Massachusetts | Lovytex Bdn Bhd Panglima Garang Malaysia. | May 21, 1992 | Cartons had visible mold contamination and culture tests were positive for mold and aerobic bacteria | Undetermined. | Nationwide.
1993-02-24 | Devices | II | z-823/828-2 | Porter: Mxr Flowmeter/nitrous Oxide - Oxygen Flowmeter Used For Nitrous Sedation During Dental Procedures: Mxr Model Numbers: (a) 2000; (b) 2000-os; (c) 2050; (d) 2055; (e) 2056; (f) 2060. Recall #z-823/828-2. | Manufacturer | Porter Instrument Company Inc. Hatfield Pennsylvania. | NOT-FOUND | Malfunctions may allow for 100% nitrous oxide with 0% 0xygen flow | Approximately 6 500 Units Were Distributed Between 1/86 And 3/13/90. | Nationwide Canada Denmark.
1993-02-24 | Devices | II | z-162-3 | Safedge Oscillating Blade (17.0 Mm Cut Edge) Part #1117 A Cutting Accessory For Orthopedic Bone Saw. Recall #z-162-3. | Manufacturer | Stryker Corporation Instruments Division Kalamazoo Michigan. | September 28, 1992 | Wrong product, part #1100, section blade, may be found in package, instead of part #1117, safedge oscillating blade | 12 Units Were Distributed. | Florida Michigan Tennessee Texas Vermont Virginia California Turkey.
1993-02-24 | Devices | II | z-265/266-3 | (a) White Irrigating Sleeve Single-use For Cataract Surgery Catalog #opo-ws; (b) Collared Phaco I/a Tubing Set Catalog #opo-30wc. The Sleeve Is Contained In The Amo Opo-30wc Collared Phaco I/a Tubing Set With White Irrigation Set And Is Also Sold Separately As Opo-ws White Irrigation Sleeve. Recall #z-265/266-3. | Manufacturer | Allergan Medical Optics (amo) Irvine California. | November 24, 1992 | Sleeve may tear or become deformed due to heat being generated during use | Approximately 28 800 Units Were Distributed. | Nationwide Argentina South Africa Japan Canada Hong Kong Spain.
1993-02-24 | Devices | II | z-267-3 | Therac Model 20 (t20) Cobalt 60 Linear Accelerator Radiation Therapy Devices Used To Treat Cancer Patients. Recall #z-267-3. | Theratronics International Ltd. (formerly Atomic Energy Of Canada Ltd.) Kanata Ontario Canada | Theratronics International Ltd. Carrollton Texas. | December 92, 1992 | Improper dose distribution may occur in electron mode if hand control pushbuttons are operated quickly in a particular sequence | 15 Units Were Distributed. | Alaska Alabama California Florida Illinois Kansas Kentucky Michigan New York Pennsylvania Texas.
1993-02-24 | Devices | II | z-268-3 | Therac Model 6(t6) Cobalt 60 Linear Accelerator Radiation Therapy Devices Used To Treat Cancer Patients. Recall #z-268-3. | Theratronics International Ltd. (formerly Atomic Energy Of Canada Ltd.) Kanata Ontario Canada | Theratronics International Ltd. Carrollton Texas. | December 92, 1992 | Unwanted gentry motion may occur with the treatment room door closed | 14 Units Were Distributed. | Alabama California Illinois Kentucky Oklahoma Minnesota Michigan Montana New Jersey New York Pennsylvania Texas Tennessee.
1993-02-24 | Devices | II | z-269/275-3 | Eldorado Theratron And Phoenix Model Cobalt 60 Radiation Therapy Devices: (a) Eldorado 76 (e76); (b) Eldorado 78 (e78); (c) Theratron 765 (t765); (d) Theratron 780 (t780); (e) Theratron 780c (t780c); (f) Theratron 1000 (t000); (g) Phoenix. Recall #z-269/275-3. | Theratronics International Ltd. (formerly Atomic Energy Of Canada Ltd) Kanata Ontario Canada | Theratronics International Ltd. Carrollton Texas. | December 92, 1992 | Potential problems involving control system pneumatics, mode push button, electrons, emergency stop switch, hand control cable, hand control support, and dual timer on these devices can affect treatment timing and gentry operation | 187 Units. | Nationwide.
1993-02-24 | Devices | II | z-281-3 | Ultraviolet (uv) Lamps. Recall #z-281-3. | Manufacturer. Fda Approved The Firm's Corrective Action Plan October 19 1992. Firm-initiated Field Correction Ongoing. | Quality Lamp Inc. Canton Ohio. | October 19, 1992 | Noncompliance with performance standards for light emitting products since the lamps are relabeled and are unreported to cdrh, and since they were manufactured in an unregistered medical establishment without a quality control testing and certification program in effect, they remain adulterated and misbranded | 7 721 Units Were Distributed. | Nationwide.
1993-02-24 | Devices | II | z-282-3 | Bud Ultraviolet Device Intended For Surface Cleaning And Surface Energizing Of Endosseous Implants. Recall #z-282-3. | Manufacturer | Bud Industries Inc. East Aurora New York. | 1/26/93 | Product was distributed without an approved 510(k), pma or ide | 5 Units Were Distributed. All Units Has Been Returned As Of 1/26/93.
1993-02-24 | Devices | II | z-290-3 | Allergan Mini Contact Lens. Recall #z-290-3. | Allergan Pharmaceuticals Irvine California | Allergan Hydron Farnborough Hampshire United Kingdom. | December 23, 1992 | Product may be contaminated with pseudomonas cepacia | 10 Units Were Distributed. | Nationwide.
1993-02-24 | Devices | II | z-292-3 | Burkhart Roentgen Brand Ceiling-mounted Support Column Which Can Hold Two Maneuverable Arms Capable Of Holding A Surgical Light An X-ray Shield Or Ct Injector. Recall #z-292-3. | Manufacturer | Burkhart Roentgen Inc. Pinellas Park Florida. | January 19, 1993 | The support column may fail causing the arm to fall | 27 Units Were Distributed. | Pennsylvania Tennessee Texas Florida Illinois Minnesota Missouri.
1993-02-24 | Devices | II | z-293-3 | Hyperformer Pump Dispensing Set Catalog #v9900-01 Used To Prepare Total Parenteral Nutrition Solutions In Hospital -6- Pharmacies. Recall #z-293-3. | Mcgaw Inc. Irvine California | Mcgaw Of Puerto Rico Inc. Sabana Grande Puerto Rico. | November 6, 1992 | Some tubing sets are improperly color-coded, which can result in incorrect compounding of solution | Approximately 32 720 Units Were Distributed; Firm Estimates 2 000 Units Remain On The Market. | Nationwide And Puerto Rico.
1993-02-24 | Devices | II | z-294/296-3 | Endopath Disposable Surgical Trocar Used To Establish A Path Of Entry For Endoscopic Instruments During Gynecologic Laparoscopy And Other Abdominal Procedures: (a) Model Lr005; (b) Model R1011; (c) Model B1011. Recall #z-294/296-3. | Ethicon Inc. Cincinnati Ohio | Ethicon Inc. Juarez Mexico. | July 25, 1992 | The devices were leaking air due to excessive sealant which interferes with the closure on the flapper valve on the sleeve component | 78 428 Units Were Distributed. | Nationwide And International.
1993-02-24 | Devices | II | z-336/337-3 | (a) Tracheal Tubes With Flexibend Adapters Catalog #5-10817 8.5 Mm Catalog #5-10818 9 Mm And Catalog #5-10820 10 Mm And Catalog #5-19112 6 Mm Size Sterile Devices Sold In Units Of 10; (b) Flexibend Adapters Catalog #5-14018 9 Mm And Catalog #5-14020 10 Mm Size Non-sterile Sold In Bulk In Units Of 100. The Adapter Is An Interim Connection Between The Tracheal Tube And The 15 Mm Adapter That Hooks To The Respirator Circuit. Recall #z-336/337-3. | Manufacturer | Sheridan Catheter Corporation Argyle New York. -7- | September 14, 1992 | Devices may crack which may cause them to be non-functional | (a) 4 120 Units; (b) 400 Units Were Distributed. | (a) Louisiana Virginia Pennsylvania Oklahoma Hawaii Finland Japan Brazil Italy Holland Australia; (b) United Kingdom And Norway.
1993-02-24 | Devices | II | z-306-3 | Ecat Scanner Positron Emission Tomography System Model #953. Recall #z-306-3. | Siemens Gammasonics Knoxville Tennessee | Cti Pet Systems Knoxville Tennessee. | NOT-FOUND | A problem in the software on the sorter printed circuit board could cause the sorter to cease functioning | 2 Units Were Distributed. | Germany.
1993-02-24 | Devices | II | z-307-3 | Ecat Scanner Positron Emission Tomography System Model #951. Recall #z-307-3. | Siemens Gammasonics Knoxville Tennessee | Cti Pet Systems Knoxville Tennessee. | August 17, 1991 | The ring source functions to normalize the device prior to scanning a patient | 2 Units Were Distributed. | Illinois Arizona.
1993-02-24 | Devices | II | z-309/313-3 | High Pressure Extension Tubes Used To Connect The Angiographic Injector Syringe To A Catheter Which Is Inserted Into The Patient's Vascular System: (a) 10" Tube With Fixed Adapter; (b) 10" Tube With Rotating Adapter; (c) 20" Tube With Rotating Adapter; (d) 30" Tube With Rotating Adapter; (e) 30" Tube With Rotating Adapter. Recall #z-309/313-3. | Liebel-flarsheim Company Cincinnati Ohio | Critical Specialties Inc. West Chester Pennsylvania. | November 10, 1992 | The 10" and 30" tubes have a defect in the seal between the -8- luer lock end fitting and the tube which could lead to leakage or air entering the tubing system | (a) 375 Units; (b) 25 Units; (c) 575 Units; (d) 1 000 Units; (e) 850 Units Were Distributed. | Nationwide And International.
1993-02-24 | Devices | III | z-298/305-3 | Ecat Scanner Positron Emission Tomography System Model Nos. 713 831 931 933 951 951r 953 And 953b. Not All Units Of These Model Numbers Are Subject To Recall. (a) Ecat Scanner Model No. 713; (b) Ecat Scanner Model No. 831; (c) Ecat Scanner Model No. 931; (d) Ecat Scanner Model No. 951; (e) Ecat Scanner Model No. 951r; (f) Ecat Scanner Model No. 953; (g) Ecat Scanner Model No. 953b; (h) Ecat Scanner Model No. 933. Recall #z-298/305-3. | Siemens Gammasonics Inc. Knoxville Tennessee | Two Different Models Were Assigned Identical Serial Numbers. Manufacturer Cti Pet Systems Knoxville Tennessee. | April 25, 1992 | Bed failed to return to zero position | 40 Units Were Distributed. | Arizona California Illinois Michigan Minnesota Missouri -9- Nebraska New York North Carolina Tennessee Australia Belgium Canada England Germany France The Netherlands Italy.
1993-02-24 | Devices | III | z-288-3 | Dupont Aca(r) Digoxin Antibody Conjugate Reagent (dgn-abc Reagent) An Accessory Of The Dupont Aca(r) Discrete Clinical Analyzer System Used To Treat A Patient Specimen Of Blood Serum In Preparation For The Measurement Of Its Digoxin Concentration. Recall #z-288-3. | Manufacturer | E.i. Dupont De Nemours & Company Wilmington Delaware. | August 16, 1991 | Inaccurate digoxin (dgn) concentration levels and the device's inability to calibrate the aca(r) discrete clinical analyzer for the dgn test method at the 0 | 10 851 Kits Of Lot 1bd694 And 3 549 Kits Of Lot 1dd635 Were Distributed; Firm Estimates None Remains On The Market. | Nationwide And International.